12070 | 1220648-2025-25751 | 21224325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224325&pc=OZD | 01/08/2025 | 02/10/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that cold extremity noted post impella cp procedure, with lack of distal pulses. imaging confirmed limb ischemia. decision was made to explant the pump and achieve hemostasis with manual compression. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25751 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 reporting contact fax number was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25751.
12071 | 1220648-2025-25752 | 21224363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224363&pc=OZD | 01/08/2025 | 02/10/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528844 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. numerous attempts made to pass impella 5.5 through left subclavian artery as there was narrowing near ostium of vessel. the medical doctor requested new device be opened due to kink distal to motor housing from manipulation on insertion. a second 5.5 device way successfully passed by lesion in subclavian artery. there was no patient harm. | [5 empty] | The investigation for the product damage and delivery issues has been completed. product was not returned for review. based on clinical description, the root cause of cannula kink was as a result of the delivery issue, which was most likely due to patient condition. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12072 | 1220648-2025-25753 | 21224438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224438&pc=OZD | 01/08/2025 | 03/03/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596829 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the purge pressure gradually increased. the motor current also increased, and the pump stopped. the purge system including cassette and fluid was changed. the examination was completed regularly and the impella was properly removed with no harm tot he patient. | [5 empty] | The investigation of pump stop has been completed. the returned pump was inspected and little biomaterial was observed in the purge gap. the pump stops were reproduced in the lab, with motor current high alarms and high purge pressure. the purge line was inspected and blood ingress was observed with an obstruction between the 25cm and 35cm mark on the catheter. the data logs from the case were not provided. the root cause of the pump stop was most likely high purge pressure from biomaterial interaction. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25753 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12073 | 1220648-2025-25754 | 21224491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224491&pc=OZD | 01/08/2025 | 03/06/2025 |  | 01/23/2025 |  |  | 01/08/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570672 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being delivered for care escalation when the patient's deep vascular anatomy prevented the advancement of the impella. the right axillary was examined and was also too small. the 5.5 escalation was aborted and there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12074 | 1220648-2025-25755 | 21224531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224531&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, aortic waveform was not present. there was also a placement signal unreliable alarm. the pump started on auto and the flow tile stated impella flow calculation disabled then read 0/0 with a motor current peaking at 962/626 (628). the pump was explanted and replaced. there was no patient harm. | [5 empty] | The investigation for the optical signal issue and low pump flow has been completed. the pump was not returned for investigation. the data log shows that the placement signal was an unreliable alarm, with a value of 3000 from the pump plugged into the console. additionally, the pump was starred, and the impella flow showed 0 l/min. all 5s optical signal values were within specifications. the imc log confirms the suspicious calibration values during startup. optical signal issue might be related to an aic software issue, and it will be investigated in 25-4068-2. the cause of the low pump flow issue was not determined since product was not returned. the cause of the optical signal issue was not determined since product was not returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12075 | 1220648-2025-25756 | 21224540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224540&pc=OZD | 01/09/2025 | 01/27/2025 |  | 01/23/2025 |  |  | 01/09/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during prep, an automated impella controller alarmed purge disk undetected twice. there was no patient involvement. | [5 empty] | A.1 revised as information was inadvertently omitted from initial manufacturer device report number 1220648-2025-25756. a.3 revised as information was incorrect on initial manufacturer device report number 1220648-2025-25756. e.1 revised us phone number as was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25756. e.2 and e.3 revised as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25756. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25756 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 added code 4121 as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25756. | The investigation into the automated impella controller has been completed. the automated impella controller was returned for investigation. the logs from the complaint date revealed that the console issued a purge disc not detected alarm (event set #402) several times. the console was examined after arriving in field service. a broken purge disc flag was identified. the root cause of this issue was a broken purge disc flag.
12076 | 1220648-2025-25757 | 21224582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224582&pc=OZD | 01/10/2025 | 02/28/2025 |  | 01/23/2025 |  |  | 01/10/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564529 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient with an impella cp implant developed hemolysis during support. an echocardiogram was performed and the pump appeared to be tangled in the papillary muscle. the staff did not try to reposition the device and the decision was made to explant the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and positioning issue has been completed. the product and data logs were not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of positioning issue was most likely patient movement related. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25757. e.4 should have been left blank on the initial manufacturer device report number: 1220648-2025-25757 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12077 | 1220648-2025-25758 | 21224775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224775&pc=OZD | 01/11/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/04/2025 | 01/23/2025 |  |  |  |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561486 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a hematoma and oozing at the anastomosis during impella 5.5 support. it was noted that a direct aortic approach was utilized for impella implant. upon diagnosis, a washout was performed while an additional suture and vascular sealant were applied to manage the condition. support continued uninterrupted. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved by adding a suture and a vascular celant at graft anastomosis. d.6a explant date was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25758 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.3 date of awareness was reported incorrectly on the initial manufacturer device report number 1220648-2025-25758. the date should of been 11-jan-2025.
12078 | 1220648-2025-25759 | 21224811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21224811&pc=PYX | 01/12/2025 | 02/24/2025 |  | 01/23/2025 |  |  | 01/12/2025 | 01/23/2025 |  |  | 04/30/2025 |  | 01/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337996 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella rp insertion, the rp was advanced over wire but unable to pass the pulmonic valve despite multiple attempts. impella wire was removed and exchanged because the 0.027 wire was slightly kinked/bent. sheath flushed again and impella rp advanced over wire and successful advanced into pulmonary artery with pigtail deflecting to right pulmonary artery. impella started. it was further reported that the patient continued to deteriorate and was made a do-not-resuscitate. the patient expired. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use issue since non - abiomed product was used which is not recommended per instructions for use. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25759 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4582 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25759 and a new code was added. code 4629 was inadvertently omitted from initial manufacturer device report number 1220648-2025-25759 and was added.
12079 | 1220648-2025-25689 | 21206947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21206947&pc=OZD | 12/27/2024 | 03/05/2025 |  | 01/22/2025 |  |  | 01/02/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589238 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Incomplete or Inadequate Priming | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported while preparing an impella cp pump to use during a high-risk percutaneous coronary intervention (pci) procedure, the purge solution did not come out of the pump. there were no curves on the display of the automated impella controller (aic), which may have impacted the ability to see the placement signals. it was also observed that the pump delivered blood; however, the pci procedure was continued and successfully completed. the pump was explanted, and the patient's outcome was noted as survived with no harm. | [5 empty] | The investigation for the optical issue and priming issue has been completed. the returned pump was inspected and there were no abnormalities. the 5s parameters were within normal range. the pump was run in the lab and there were no issues reproduced. the pump was primed and delivered purge fluid and purge values were within specification. the data logs from the case showed a spike in purge pressure on case start but showed purge pressure stabilized and was within normal range for the rest of the case after troubleshooting. the logs showed that when the pump was connected, signal-to-noise ratio (snr) was low (<2500) but marginally sufficient to obtain stable optical signal and calibrate new g0. after pump was reconnected (probably in an attempt to resolve purge issue), the snr was extremely low, most likely due to air gap (presumably in pump connector). the snr remained low for the rest of the case. contrast dropped, lamp increased, gain decreased, light increased. the root cause of the optical signal issue was most likely due to an air gap between the automated impella controller and the patient cable plug. no problem was found as the root cause of the priming issue as issue did not reproduce. | Patient-specific information is unavailable at the time of the medwatch writing, and therefore, respective files have been reflected as "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12080 | 1220648-2025-25690 | 21207059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207059&pc=OZD | 01/03/2025 | 02/19/2025 |  | 01/22/2025 |  |  | 01/03/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support; anatomical issues prevented implant of an impella 5.5. the patient was sent to the unit in critical condition. approximately 17 days into support, the patient developed low purge pressure alarms which were unresolved following troubleshooting. discussions with the cardiovascular surgeon surrounding goals of care were made. due to no current exit plan and the complexities of the patient, the decision was made to increase the dextrose in the purge to d20 with 25meq bicarbonate and wait/watch until the impella cp potentially fails. there was an audible hissing/buzzing sound coming from the insertion site; however, there was no visible purge fluid. the cardiovascular surgeon and critical care team put a plan in place to pull the impella cp back into the aorta should the pump fail. explant supplies remained at bedside. two days later it was reported the impella cp device had been explanted. care was withdrawn, and the patient expired. | [5 empty] | The impella device was saved for return. however, it was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. data logs were reviewed. low purge pressure alarms occurred at evening of 1/2 at p-7 unresolved post troubleshooting. no evidence observed to confirm the root cause for reported issue. the pump left on support until patient expired. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12081 | 1220648-2025-25691 | 21207172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207172&pc=OZD | 10/09/2024 | 02/11/2025 |  | 01/22/2025 |  |  | 10/09/2024 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513032 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | [12 empty] | Required Intervention | Life Threatening | Death |  |  |  | VA ECMO. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella device for mechanical circulatory support and experienced hit, high purge pressure and sepsis. the patient would eventually expire. the patient was escalated to an impella 5.5, and was on a respirator, veno-arterial extracorporeal membrane oxygenation (va ecmo) and dialysis therapies prior. nine days into support on the impella 5.5 pump, the physician reported a purge pressure high alarm with purge flows at 1.6-2.6ml and purge pressure at 1060mmhg. attention was drawn to the fact that the pump can stop at any time, and the recommendation made to change the pump; discussion regarding lysis option was made with sharing of the lysis protocol as well. the following day, the physician indicated the pump was left in. the purge flow was at maximum 1ml, and purge pressure noted sometimes fell below 1000mmhg and the next day stated the pump was successfully lysed, the patient was awake and successfully in weaning. however, approximately three days later, the patient was reportedly intubated and sedated and at the beginning sepsis waiting for further decisions. two days later therapies were stopped, and care was withdrawn due to lack of recovery and possible positive outcome. | [5 empty] | The investigation of infection/sepsis and high purge pressure has been completed. no product was returned for analysis. no data logs were returned for analysis. the root cause of high purge pressure and sepsis was not determined as no product or data logs were returned for analysis. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b.2 life-threatening was inadvertently selected on the initial manufacturer device report number and should of not been. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address as it was previously entered incorrectly on the initial manufacturer device report number. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12082 | 1220648-2025-25692 | 21207207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207207&pc=OZD | 01/02/2025 | 09/22/2025 |  | 01/22/2025 |  |  | 01/02/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573372 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Thrombocytopenia | Epistaxis | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | RIGHT VENTRICULAR ASSIST DEVICE | [17 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that the nurse called support for "placement signal in aorta unknown" message. it was reported that the device had proper placement, and there was not a need for repositioning. it was suggested by support that the placement signal alarm be turned off, however the medical team declined. | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) for a bridge to transplant. approximately six days after start of support, the patient developed a hematoma near the impella right axillary artery insertion site. it was noted the patient also has a right ventricular assist device in the same region. over the following next couple of days, the patient was reportedly stable, off all pressers, holding of heparin for supratherapeutic result and underwent surgical evacuation of the hematoma. the following day the surgeon at bedside drained a foot abscess. heparin remained on hold post hematoma with plans to restart the following day. the impella support level was p-6 with 4.2l/min flow. the patient received one unit of red blood cells for log hemoglobin and hematocrit. the patient remained status 7 for heart transplant. additionally, this day, the nurse reported the automated impella controller (aic) had an "aic" failure alert that self-resolved; the team was instructed to exchange out the aic for the backup unit. approximately six days later, the patient experienced a nosebleed after the restarted heparin was increased. the next day, the patient was noted to be heparin-induced thrombocytopenia (hit) positive, and heparin was switch to argatroban. the patient experienced runs of ventricular tachycardia requiring shock from the implantable cardioverter defibrillator. the ventricular tachycardia was noted unrelated to the impella however (without a likely cause provided). days following, the patient remains hit positive and argatroban still running. the patient is now listed as status 1 for orthotopic heart transplant and remains on impella mcs pending transplant. |  |  |  |  | The investigation of access site bleeding and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. e4 should have been left blank on manufacturer device report 1220648-2025-25692. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. the data logs were reviewed, and the signal-to-noise ratio (snr) is widespread, the placement signal rails, the lamp is 100, contrast spikes, gain spikes and light remains stable. the pump flow spikes with the placement signal. without the product back, the issue cannot be confirmed. the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2025-25692. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the automated impella controller serial number (b)(6) with date of event 10-jan-2025 and identifier ending in (b)(6).
12083 | 1220648-2025-25693 | 21207211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207211&pc=OZD | 01/03/2025 | 02/17/2025 |  | 01/22/2025 |  |  | 01/03/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient in cardiogenic shock from an acute myocardial infarction unsuccessful due to the cardiac anatomy. the device was unable to advance past an "extensively" calcified aortic root and the case was aborted. the outcome noted the patient survived and reportedly in stable condition following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issues was most likely patient condition since patient had an extensively calcified aortic root. d6a and d6b were inadvertently omitted from manufacturer device report 1220648-2025-25693. e4 should have been left blank on manufacturer device report 1220648-2025-25693. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12084 | 1220648-2025-25694 | 21207222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207222&pc=OZD | 01/04/2025 | 02/20/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559308 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was a pump stop. the pump was removed and after removal there was a clot near the inlet of the cannula. the pump was replaced and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus and pump stop has been completed. the impella cp was returned for evaluation. upon visual inspection, the impeller blade was found to be broken. no biomaterial was present on the returned pump. the pump was run in the lab for approximately six minutes and the pump stop did not reproduce. data logs were reviewed and real time logs at time of pump stop showed multiple motor current spikes with concurrent speed deviations and interruptions in pump flow occurring around 10 minutes before alarms. immediately prior to ¿impella stopped ¿ motor current high¿ alarms, multiple motor current spikes with speed deviations and interruptions in pump flow were observed. additionally, at time of alarm and pump stop, placement signal spikes to 3000. the pump was able to be restarted multiples times before the pumpwas able to stay on for an additional four hours. approximately three minutes after last pump stop alarm, more motor current spikes were observed and a significant drop in flows was noted after the last motor current spike. the pump flows remained low for the remainder of the case. clinical details noted that the patient was on bipella support and also had an existing transcatheter aortic valve replacement valve (tavr) in place. additionally, a clot was present on the inflow cage upon explant. as the pump was able to be restarted and run for four more hours of support before the pump was explanted after pump stop occurred. the reported failure mode of "pump stop" was investigated under "temporary pump stop". the root cause of the temporary pump stop was most likely due to an interaction with the tavr valve present in the patient while on support. per instructions for use, unintentional interaction of the impeller and tavr valve can result in destruction of the impeller blades. as the necessary information was not provided, the root cause of the thrombus was not determined. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.10 added transcatheter aortic valve replacement (tvar) as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25694. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25694 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 some additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25694 and were added.
12085 | 1220648-2025-25695 | 21207229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207229&pc=OZD | 01/04/2025 | 02/20/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582896 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and three days into impella cp support, the patient was being transported for computed tomography imaging when the purge line was inadvertently pulled and disconnected from the red impella plug. a purge system open alarm appeared and the motor current spiked. flow saturation then developed, and the pump subsequently stopped. the pump was replaced as a result of the noted issue. upon conclusion of support, bleeding was noted from the access site, and it was discovered that the initial impella stick had been too high during implant resulting in an arterial bleed. a vbx stent was placed, and the vessel was successfully closed. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump stop, saturated flow, vascular damage, and purge leak has been completed. no product or logs were returned for evaluation. clinical details noted that when patient was being transported, the clinical staff reported the purge line disconnected from the red impella plug connection. the pump began to alarm "purge system open". a motor current spike was observed and pump saturation was also seen. "impella stopped motor current high" alarm was present and flows decreased to 0. additionally, at conclusion of impella support, arterial bleed from initial impella stick being too high. a stent was placed and vessel was able to close. the root cause of the pump stop could not be definitively determined as no product or logs were returned for evaluation. the root cause of the saturated flow could not be definitively determined. the root cause of the purge leak - pump could not be definitively determined as the leak location could not be definitively determined. the root cause of the vascular damage - peripheral vessel was most likely use related as the initial impella stick was too high. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.7 information was added as it was inadvertently omitted from the initial mnaufacturer device report number 1220648-2025-25695. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25695 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously initial mnaufacturer device report number 1220648-2025-25695.
12086 | 1220648-2025-25696 | 21207241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207241&pc=OZD | 01/04/2025 | 02/07/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia immediately following the implant. the patient presented positive with the flu and underwent a left heart catheterization which revealed 90% occluded left main stenosis with transient loss of flow to the left circumflex resulting in an inferior myocardial infarction, and the decision was made to implant the impella cp device for a bridge to decision on surgical evaluation. once the impella was placed in right common femoral artery, distal runoff angiogram revealed sluggish flow. plans were to watch over next couple of hours. ultimately, a fem/fem bypass was completed; a 6fr retrograde sheath was placed in the left common femora artery connected by a male-to-male adapter to a 5fr antegrade sheath placed in the right superficial femoral artery. distal perfusion verified by doppler. the setup was flushed hourly and doppler. the following day, it was noted the patient remained stable with no events overnight. the impella site was without bleeding and hematoma and the fem/fem bypass remained patent with flushing and doppler every hour. the patient was kept sedated on ventilator until completion of the high-risk percutaneous coronary intervention (pci) scheduled for the following day (as noted, the patient was deemed ineligible for surgery due to heavily calcified aorta and root). the high-risk pci (left main, left anterior descending and circumflex arteries stented) was completed, the impella weaned down during the procedure and removed afterwards without difficulty. the patient's outcome following the explant was noted as survived in stable condition. | [5 empty] | The root cause of the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia.
12087 | 1220648-2025-25698 | 21207290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207290&pc=OZD | 01/04/2025 | 10/02/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585460 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  |  | Death | Yes | Yes | Malposition of Device | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day positioning alarms triggered despite pulsatile motor current. the impella was repositioned and proper placement confirmed under echocardiogram. patient still required chemical support to maintain hemodynamics, possible work up for impella 5.5 was indicated. the following day, waveforms on remained unchanged since the morning when echocardiogram confirmed inlet 3.8cm from aortic valve. placement monitoring was disabled. another echocardiogram was planned with repositioning. next days following, it was noted due to flow saturation, inaccurate position monitoring, and lack of improvement, a decision was made to explant the impella cp without exchange or escalation in care the impella cp device was removed on day four of support, and the patient was medically managed. the patient expired the same day after explant, however. | [5 empty] | The investigation of saturated flow and optical signal issue has been completed. saturated flow: data logs were reviewed and the saturated flows were confirmed on (b)(6) 2025. leading up to the first position alarm there is a drop in motor current, speed deviation, and suction alarm while running at a constant p-level. additionally, the left ventricle placement signal (ps) amplitude becomes variable, often shrinking down to match the aortic ps. however, there is no motor current spike that would be expected during biomaterial ingestion. following this event, the user starts to change p-level often, both position in ventricle and aorta alarms intermittently trigger, and pump flows become saturated depending on the p-level. later in the case, specifically while running constant at p-7 the evening of (b)(6) 2025, motor current/pump flows can be seen to trend up and become saturated. returned product was investigated and there were no abnormalities. there was no sign of biomaterial or blood ingress into the purge gap or even the motor housing when it was torn down. the pump was run in a bucket for 15 minutes and flows were within specification. since data logs did not show conclusive causal evidence and the issue did not reproduce, the root cause of the saturated flows could not be determined. optical signal issue: data logs were reviewed and the positioning alarms were confirmed on (b)(6)2025. leading up to the first position alarm there is a drop in motor current, speed deviation, and suction alarm while running at a constant p-level. additionally, the left ventricle placement signal amplitude becomes variable, often shrinking down to match the aortic ps. following this event, the user starts to change p-level often, and both position in ventricle and aorta alarms intermittently trigger while pump flows become saturated at times. in terms of 5s parameters, gain, light, and lamp were all stable through the case. the signal-to-noise ratio and contrast were consistently low, and this would not cause the positioning alarms, only exacerbate the issue. returned product was investigated and there were no abnormalities. the pump passed laser continuity and there was no sign of biomaterial in the purge gap or anywhere in the motor housing when torn down. when connected to a console, all 5s parameters and pump flows were within specification when running the pump in a bucket of water for 16 minutes. clinical details report that the pumps position was confirmed several times with echocardiography. since data logs do not show conclusive causal evidence, and the issue did not reproduce in the lab, the root cause of the optical signal issues could not be determined. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding.
12088 | 1220648-2025-25697 | 21207299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207299&pc=OZD | 01/04/2025 | 02/17/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540204 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia within hours of support. the impella was implanted without any issues via the right femoral artery. the patient had little pulses to bilateral extremities prior to impella insertion thought to be from diabetes and large amount of vasopressor support. the peel-away sheath was left in place and pressurized. impella secured. reportedly, the physician was aware of no pulses in extremities; the patient was transferred to the unit on impella mcs. however, after approximately seven hours of support the impella cp device was removed due to limited flow in the leg. the patient required increased vasopressor support overnight and reported to be stable following the event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. e4 should have been left blank on manufacturer device report 1220648-2025-25697. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia.
12089 | 1220648-2025-25699 | 21207555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207555&pc=OZD |  | 02/20/2025 |  | 01/22/2025 |  |  | 01/03/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562645 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted an impella cp device via the left femoral artery for mechanical circulatory support (mcs) followed by percutaneous coronary intervention via the right femoral artery. while on impella cp mcs, the patient's left leg appeared to be ischemic with a lower temperature than the right. after a computed tomography scan, it was determined that the patient's leg condition was too severe to be saved. the patient's impella side left leg was amputated. further comments, noted the patient had very sick and especially small (diameter) blood vessels and also a central line was hard to place. the latest update indicated the patient was reportedly in critical condition following these events. no further information was provided. days between implant and ischemia are unknown. | [5 empty] | Patient-specific information, implant/explant dates and event date are unavailable or unclear at the time of this medical device report writing and therefore respective fields have been populated with "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. b.2 revised selections are they were previously not entered correctly on the initial manufacturer device report number 1220648-2025-25699.
12090 | 1220648-2025-25700 | 21207845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207845&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/22/2025 |  |  | 12/23/2024 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and encountered removal difficulty. the impella was successfully implanted, and plain old balloon angioplasty was performed on the left circumflex artery. at the end of the pci, the physician was removing the impella cp device but did not remove the companion sheath from the peel-away sheath. the peel-away sheath was subsequently compromised and a hematoma developed around the access site (determined a non-serious adverse event). the physician subsequently removed the pump and the companion sheath. the access site was perclosed with applied manual pressure. following, the access site was clean dry and intact, the patient was transferred to the unit and reported to be in stable condition following the event. | [5 empty] | The investigation for the reported removal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided noted that during planned pump explant, the attempt to explant the pump prior to removing the companion sheath. the cause of the pump removal issues was most likely use related. hematoma noted after the removal issues; excessive force stated to have been used. the cause of the access site bleeding minor issue was most likely use related as it occurred after the removal. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12091 | 1220648-2025-25701 | 21207911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21207911&pc=OZD | 01/04/2025 | 02/14/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia the following day, approximately 20 hours after start of support. the patient presented to the emergency department with confusion and shortness of breath and was admitted with community acquired pneumonia, urinary tract infection and sepsis. the patient decompensated early requiring intubation and one round of cardiopulmonary resuscitation. the patient was hemodynamically compromised after achieving return of spontaneous circulation and was taken to the cardiovascular lab. right heart catheterization was performed. cardiac output and cardiac index were "low." a pulmonary artery catheter was left in place, and the decision was made to implant an impella cp device. which was completed without complication. the patient, intubated and sedated, was sent to the unit on support. however, the following day in the morning, less that 24 hours later, pulseless extremity was identified. the patient was brought back to the operating room for impella cp device removal and possible surgical intervention. reportedly, the impella was successfully weaned and explanted, the patient received stenting to left iliac artery and thrombectomy. the patient returned to unit in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g.1 revised manufacturing site name and address as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25701. h.6 code 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25701 and was added.
12092 | 1220648-2025-25702 | 21208119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21208119&pc=OZD | 01/04/2025 | 02/14/2025 |  | 01/22/2025 |  |  | 01/04/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582250 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient with st-segment elevation myocardial infarction was unsuccessful due to the cardiac anatomy. the device was unable to pass due to the patient's peripheral artery disease. and the case was aborted. additionally, it was noted the impella access site was bleeding. no blood or blood product were required; however, treatment to resolve the bleeding was not indicated. the patient's outcome following the event is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. d6a and d6b was inadvertently omitted from manufacturer device report 1220648-2025-25702. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12093 | 1220648-2025-25703 | 21208281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21208281&pc=OZD | 01/05/2025 | 02/21/2025 |  | 01/22/2025 |  |  | 01/05/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, hemolysis in urine, labs reporting fine. additionally, no pulses were noted in bilateral lower leg extremities. care was withdrawn and patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the limb ischemia and hemolysis has been completed. data logs were reviewed. there was a motor current and placement signal spike, and speed deviation followed by low pump flows and suction during the time of the reported hemolysis. the motor current remained variable through the case. the cause of the hemolysis issue most likely was biomaterial ingestion. the cause of the low pump flow most likely was due to biomaterial ingestion. the root cause of the limb ischemia could not be determined without additional information.
12094 | 1220648-2025-25704 | 21208456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21208456&pc=OZD | 01/03/2025 | 02/18/2025 |  | 01/22/2025 |  |  | 01/03/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient in cardiogenic shock from an acute myocardial infarction was unsuccessful. the patient presented in profound cardiogenic shock on multiple pressors and inotropes. the impella sheath and wire inserted (access site location and side unknown), and the pump was prepped. however, the patient's anatomy would not accommodate delivery of the impella cp device. reportedly, the patient expired shortly thereafter. no further details were made available. | [5 empty] | Some patient-specific information is unavailable at the time of this mdr writing and therefore, respective fields were reflected "unknown" or left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature.
12095 | 1220648-2025-25705 | 21209004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21209004&pc=OZD | 01/06/2025 | 10/09/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613395 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that the patient expired while on support. the patient was noted to have multiple issues and was ineligible and too sick for a heart, kidney and liver transplant. there was no interruption in support as the automated impella controller (aic) did not appear to have been exchanged or issue impacted the therapy. the device remained in service. per medical safety officer review, the use of the device cannot conclusively be associated with the outcome. | Additional information was received indicating the patient had multiple issues and was ineligible and too sick for heart, kidney and liver transplant. the patient expired while on support. | The complainant reported that the aorta placement signal no longer displays on automated impella controller. the patient remains on support. |  |  |  | The investigation of automated impella controller (aic) - loss of opm data has been completed. console logs from the reported day of the event are consistent with the complaint. the logs showed instances of the following errors: "algs suspended opm signal invalid,"and "invalid_bad_snr_sample_mask" due to signal-to-noise ratio (snr) falling below the minimum threshold. subsequently, peak snr error was observed. long term log and real time log data showed that during the event, contrast values fluctuated between -3000 and 3000, and snr was near zero, indicating that light was not reaching the optical bench detector. the console was not exchanged in response to the optical signal issue. the root cause for the placement signal not reliable alarms, was not determined as the aic was not returned. b.5 additional information was received, and abiomed's medical safety officer review was completed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. additional information received for b1, b2, b5 and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12096 | 1220648-2025-25706 | 21209342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21209342&pc=OZD | 01/06/2025 | 02/19/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533745 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp and there was impella access site bleeding. manual pressure was done, suture, and two units of blood were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the turoy burst was not locked causing the bleed.
12097 | 1220648-2025-25707 | 21209517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21209517&pc=OZD | 01/06/2025 | 02/21/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582354 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, there was bleeding behind the heart that required surgical repair, so the heart needed to be rotated completely to reach the bleeding site. the surgeon pulled the pump back to where it was at 3cm across the valve. when folding over the ventricle to repair the vein behind the heart that was bleeding, the pump perforated the ventricle. the ventricle tissue was incredibly friable. the surgeon then decided to pull the pump out and go back on bypass to repair the ventricle and repair the initial bleeding site behind the heart. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. no product was returned and logs are not applicable. the root cause of the ventricle perforation was most likely patient condition related as the very friable ventricle tissue was perforated by the pump while folding over the ventricle despite the surgeon taking caution and pulling the pump back across the valve beforehand. h6 code 2001 was reported incorrectly on manufacturer device report 1220648-2025-25707. new code was added to health effect - clinical code.
12098 | 1220648-2025-25708 | 21209713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21209713&pc=OZD | 01/06/2025 | 03/12/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577400 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. the impella was dislocated completely in aorta. initially, they pulled the pump back completely into the aorta and let it run on low p-level. suddenly the patient became very critical, had to be resuscitated for 3 minutes. the pump was repositioned with trans-esophageal echo and it worked well. position was not perfect, but well enough to let it run on p6. also, the patient is septic. the patient expired while on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of infection/sepsis could not be determined. the root cause of positioning issue was most likely use related.
12099 | 1220648-2025-25709 | 21210325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21210325&pc=OZD | 01/07/2025 | 02/12/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1291635 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, an automated impella controller (aic) showed ¿purge disc not detected¿ after placing the purge cassette during the procedure. several trials to take out and place in the pc back in were made, but no success. problem solved by replacing the aic. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was tested. a broken purge disc retainer was identified. the purge disc retainer was observed to be broken and was the root cause of the reported purge disc not detected issue. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25709. h.6 code 4114 was entered incorrectly on the initial manufacturer device report number 1220648-2025-25709 as the device was received.
12100 | 1220648-2025-25710 | 21210705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21210705&pc=OZD | 01/01/2025 | 02/24/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583825 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had limb ischemia on their left leg. the vascular surgeon confirmed the vessel was too small for the impella. the impella was removed and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12101 | 1220648-2025-25711 | 21210879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21210879&pc=OZD | 01/06/2025 | 01/22/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Charge | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during the case setup, the automated impella controller continued to lose battery power despite the unit being plugged in. a yellow alarm appeared indicating that the battery power was low and the decision was made to replace the console. there was no patient harm. | [5 empty] | The investigation into the automated impella controller has been completed. the automated impella controller was returned for investigation. the reported charging issue and yellow battery level low alarms were able to be reproduced. batttest results revealed that the health of the batteries were normal but the ac indicator was not present when the console was connected to ac power. after further inspection of the aic, it was discovered that the copper wire on the inside of the power cord plug was detached and exposed. console logs from the reported complaint date of (b)(6) 2025 are consistent with the complaint. there was an instance of battery level low. the root cause of this issue was a damaged power cord plug.
12102 | 1220648-2025-25712 | 21211302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21211302&pc=OZD | 01/07/2025 | 03/07/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582536 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. pump initially ran at normal flow rate. after approximately 10 minutes, a suction alarm appeared, and the flow dropped to 1.5 liters. the suction alarm could no longer be interrupted except for p2. despite repositioning under fluroscopy and echocardiography, automated impella controller (aic) to aic transfer performed. without success. implantation of a new impella cp performed. second system runs without further problems. there was no reported patient harm. | [5 empty] | The investigation for the low pump flow has been completed. data was not returned for review. visual inspection found the impeller blades were broken. no other issues were found on the rest of the pump. the root cause of the low flow was fractured impeller blades but the exact cause of the damaged blades could not be determined due to lack of clinical information. d.9 device date returned/information was added. g.1 added reporting contact address line 1, reporting contact us city, reporting contact state, reporting contact us zip code and reporting contact country as they were inadvertently omitted from the initial manufacturer device report number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12103 | 1220648-2025-25713 | 21212899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21212899&pc=OZD | 01/06/2025 | 03/14/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had an unexpected controller failure alarm while on patient support. aic was exchanged and there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) ¿ controller error (yellow alarm) has been completed. the data analysis confirmed the console shut down unexpectedly. during the device analysis the investigator was unable to reproduce the reported failure mode upon running the optical test pump. the root cause for the unexpected controller shutdown was not determined due to the inability to reproduce. h6 code 925 was reported incorrectly on manufacturer device report 1220648-2025-25713. new code was added to component code.
12104 | 1220648-2025-25714 | 21213508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213508&pc=OZD | 01/01/2025 | 03/07/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541713 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella cp support. during transfer to a new site, the patient felt ill and hit his face against the steering wheel. this caused the patient to suffer a broken jaw and massive bleeding for the duration of support. heparin was suspended in an attempt to manage the bleed, but this resulted in an increase of purge pressure and eventually a pump stop. the pump had to be restarted multiple times and the decision was made to replace the device. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel, pump stop, and high purge pressure has been completed. there was no product or data logs returned. the clinical stated that the patient hit his face against the steering wheel causing massive bleeding for the duration of support. heparin was reduced to try to suspend the bleeding the root cause of the vascular damage - peripheral vessel is most likely due to the patients condition as the patient suffered a severe head injury causing persistent bleeding throughout impella support. it was stated that due to the heparin being stopped in the purge for bleeding, the impella had an increase in purge pressure requiring the device to be turned off and on again twice and they had to replace it. this statement is most likely referring to the pump stopping. due to lack of product or logs returned, the root cause of the high purge pressure and pump stop cannot be confirmed. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25714 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 revised to reflect instructions for use that were inadvertently omitted from initial manufacturer device report number 1220648-2025-25714.
12105 | 1220648-2025-25715 | 21213545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213545&pc=OZD | 12/27/2024 | 03/12/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548988 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and the shaft kinked upon insertion. the impella 5.5 was aborted and the impella cp was inserted. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through the vasculature. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25715. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25715 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12106 | 1220648-2025-25716 | 21213561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213561&pc=OZD | 01/04/2025 | 03/06/2025 |  | 01/22/2025 |  |  | 01/07/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574593 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support when the angle of the placement sheath became shallow, causing the blood vessel at the tip of the placement sheath to dissect. the impella was removed and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the vascular damage issue was most likely use related, the angle of the placement sheath became shallow, causing the blood vessel at the tip of the placement sheath to dissect. b3 date of the event was revised to reflect the correct date on manufacturer device report 1220648-2025-25716. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25716 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12107 | 1220648-2025-25717 | 21213634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213634&pc=OZD | 01/09/2025 | 02/24/2025 |  | 01/22/2025 |  |  | 01/09/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1415626 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella support, the automated impella controller failed to turn off. the staff had to forcibly shut the console down despite the connection cable being unplugged. there was no patient involvement at the time of the noted issue. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) shutdown or restart issue has been completed. according to that data and device analysis the root cause of the aic's inability to shut down correctly was due to an aic software anomaly. e4 should have been left blank on manufacturer device report 1220648-2025-25717.
12108 | 1220648-2025-25718 | 21213644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213644&pc=OZD | 01/05/2025 | 02/04/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1626081 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported during preparation for an emergency case, the purge disc holder on the automated impella controller (aic) was noted to be broken. the aic was exchanged with no report or indication of interruption in therapy or harm to the patient. | [5 empty] | The investigation for the automated impella controller (aic) purge disc/flag issues has been completed. data logs were reviewed which found that the user were able to insert the disc and connect the pump. during purge priming, with the cassette being inserted, the console displayed ¿purge disc not detected alarm¿ twice, indicating most likely the retainer and flag were broken. this aligns with the reported failure mode ¿purge disk holder has broken off¿. the aic was returned for evaluation finding that the purge retainer was broken. the purge disc retainer was confirmed to be broken. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25718. e.1 revised name prefix/title, first name, middle name and last name as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25718. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25718 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address section as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25718. h.6 code 2199 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25718 and a new code was added. | Patient-specific information and initial report name are unable at the time of the mdr writing and therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
12109 | 1220648-2025-25719 | 21213648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21213648&pc=OZD | 12/25/2024 | 02/21/2025 |  | 01/22/2025 |  |  | 01/06/2025 | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025506107 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and developed shortness of breath and subsequently experienced cardiopulmonary arrest with a demised outcome. the patient was on independent support by the impella 5.5 (recently weaned off from extracorporeal membrane oxygenation) and had developed shortness of breath while sitting during rehabilitation on the seventh day of impella mcs. an angiography revealed a pulmonary embolism in the main pulmonary artery which was difficult to treat leading to a do not attempt resuscitation. the patient was diagnosed with heparin induced thrombocytopenia (hit) at the time of the above conditions and it was believed that the hit had contributed to the pulmonary embolism. heparin was in the purge line leading up to the event. the patient would expire two days later. reportedly, the cause of death was pulmonary embolism, and the cause of hit was the use of heparin. it is unknown of the hit contributed to the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia and thrombus could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ b.2 revised selections as an incorrect selection was previously selected on the initial manufacturer device report number 1220648-2025-25719 and a new selection was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25719 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted initial manufacturer device report number 1220648-2025-25719.
12110 | MW5165142 | 21219851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21219851&pc=OZD | 01/09/2025 | 01/13/2025 |  | 01/22/2025 |  |  |  | 01/22/2025 |  |  |  |  | 01/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  |  |  | 548298 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Nurse | [6 empty] | 74 YR |  | Male | 104 KG | Non Hispanic | White | Injury | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [28 empty] | Initial submission | [9 empty] | Impella put in patient. when it was turned on, the machine and the pump worked, but the sensor did not. impella removed and a new device put in. no harm to the patient. impella was put in as a protective device. s/n of impella (b)(6). the rep wanted the device released to him so the company can run studies. device to be released to (b)(4) rep. of abiomed.
12111 | 1220648-2025-25657 | 21196327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196327&pc=OZD | 01/02/2025 | 02/19/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551078 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Malposition of Device | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for medical circulatory support and experienced a pump leak and runs of ventricular tachycardia. additionally, patient support also required exchanging of the automated impella controller (aic). two days prior to the impella 5.5 device implant, the patient returned for a follow-up visit post coronary percutaneous coronary intervention (totally occluded left anterior descending artery, thrombectomy and three drug-eluting stents) and was found to be tachycardiac. the impella 5.5 was implanted without incident, and the patient was sent to the unit for support. however, approximately four days after start of support, the nurse reported a purge reservoir leak (purge cassette change ruled out faulty cassette). the decision was made not to immediately remove the pump as the patient was planned for upcoming left ventricular assist device placement. of note, same day there was report of a successful exchange of the aic for an unresolved aic power disconnected" alarm. support continued. however, two days later the nurse call to discuss the position of the impella 5.5 pump, as the patient had more frequent runs of ventricular tachycardia. there was no suction, low flow or hemolysis. reportedly, patient movement triggered the ventricular tachycardia. echocardiogram was completed. images revealed pump in a good position but potentially too close to papillary muscle or mitral valve chordae. risks/benefits of repositioning were discussed; however, there is no information that indicates the pump was repositioned. eleven days later the impella 5.5 device was explant and heart transplant completed. the patient reportedly remained intubated post operatively and was on an intra-aortic balloon pump for increase coronary pressure. | [5 empty] | The investigation of purge leak ¿ pump and vf/vt/af/at has been completed. the pump was not returned for investigation. clinical information provided mentions that there was a possible purge reservoir leak. without the pump or pictures, the failure cannot be investigated. therefore, the root cause of the purge leak issue was not determined since product was not returned. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the automated impella controller serial number (b)(6) with date of event 02-jan-2025 and identifier ending in (b)(6).
12112 | 1220648-2025-25658 | 21196560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196560&pc=OZD | 08/26/2024 | 10/03/2025 |  | 01/21/2025 |  |  | 08/26/2024 | 01/21/2025 |  |  | 11/30/2025 |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435243 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Death | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient (unknown age and ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced positioning issues and would subsequently expire. the impella pump was successfully implanted and patient on the unit with extracorporeal membrane oxygenation (ecmo). approximately four days after start of support, the physician reported a recurring alarm ¿impella in aorta¿ and "impella position unknown." the position was noted "good" at 4.9cm in the ventricle, pulsatile curves. the patient was with his/her own ejection fraction of approximately 10-15% and on extracorporeal membrane oxygenation therapy. support continued until reportedly the patient expired approximately 19 days later. the decision was made to withdraw care. no further information regarding the cause of death was provided, and there is not enough evidence to exclude the impella device as an associated factor. | A complainant reported a patient was on an impella 5.5 along with extracorporeal membrane oxygenation support. while on support there were recurring alarms ¿impella in aorta¿ and "impella position unknown" reported. no further information was provided. care was withdrawn and the patient expired. |  |  |  |  | B.2 required intervention was selected incorrectly on the initial manufacturer device report number 1220648-2025-25658. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25658 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not returned; therefore, an evaluation, analysis of the device was not possible. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. product was fixed from 0550-0006 to 2000163 as reported to as investigated because the reported complaint was against the purge cassette, not the pump. product was not returned for review. the root cause of the pc leak was unable to be determined as limited clinical details were provided and no product was returned for review. failure mode positioning issues was removed since it was reported that positioning alarm were triggering despite the pump being in proper position. data logs were reviewed and pump in aorta alarms were confirmed. pump optical metrics were stable with no placement signal not reliable alarms. since the site reported that the pump was in good position, the cause of the false position in aorta alarms was patient condition related. the pump was not explanted due to the optical signal issue, positioning alarms reported on 8/26 and pump explanted (b)(6). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
12113 | 1220648-2025-25659 | 21196570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196570&pc=OZD | 11/21/2024 | 01/23/2025 |  | 01/21/2025 |  |  | 11/21/2024 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567960 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | [11 empty] | Required Intervention | Death | Disability | Life Threatening | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced an access site complication and stroke with a demised outcome. the patient presented with congestive heart failure exacerbation. an intra-aortic balloon pump was placed and showed worsening hemodynamics in which the decision was made to cannulate for veno-arterial extracorporeal membrane oxygenation (va ecmo). following, the patient was transported to the cardiovascular operating room for an axillary impella 5.5 device placement. the impella implant procedure was successfully completed, patient extubated and transferred to the intensive care unit. approximately seven days after start of the support, an access site hematoma developed. an evacuation was performed to the hematoma, and mcs continued. five days later, the patient was severely diaphoretic, possible infection from prior ecmo groin site; however, results were negative. the patient had an incision and drainage washout two days following, and a couple of days thereafter, the patient experienced left sided weakness noted to be from pressing on a nerve during the ecmo groin site washout procedure. three days following the left-sided weakness, the patient had slurred speech with maintained left sided weakness. computed tomography showed the patient had a stroke on the brain stem. the patient was intubated to maintain the airway but would later expire this same day. the cause of death was not provided; therefore, unknown and the impella could not be excluded as a possible contributing factor. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12114 | 1220648-2025-25660 | 21196588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196588&pc=OZD | 01/02/2025 | 02/13/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571895 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device as elective placement for mechanical circulatory support (mcs) and at time of removal difficulty was encountered requiring surgical intervention. the patient was transferred from outside hospital and found to have severe aortic insufficiency with dilated ascending. the patient underwent aortic valve replacement/ascending with axillary impella 5.5 placement. after 15 days of mcs, the patient was taken the operating room of the impella removal. the patient was intubated and prepped. impella 5.5 device was weaned to performance level p-2. and then turned off. sutures were cut, the impella was attempted to be removed, and resistance was encountered. multiple unsuccessful attempts were to get impella out of graft. while back force was being applied to remove impella catheter, it broke in half about 1mm above the motor housing. the cannula was free floating in left ventricle. consequently, the patient underwent emergent sternotomy and cannulated for bypass. heart arrested. tube graft opened and impella was removed. the graft was closed., and the heart deaired. intra-aortic balloon pump was placed and patient separated from bypass. the patient was reported to have survived the event and last known to be in critical condition. | [5 empty] | The impella device was not received from the customer. abiomed representative was unable to retrieve the pump as the site retained the pump for evaluation by pathology. therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (detached inflow/outflow - great vessel) has been completed. the pump was not returned for investigation. as per the clinical information provided, the cannula was completely detached during removal of impella and was left floating in the heart chamber. without the product, the failure could not be investigated. therefore, the root cause of the product damage issue was not determined since the product was not returned for investigation. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25660 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4617 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25660 and was added.
12115 | 1220648-2025-25661 | 21196616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196616&pc=OZD | 01/02/2025 | 02/20/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility a patient (unknown race and ethnicity) was implanted with an impella 5.5 device, and during use, approximately four days after start of support, the automated impella controller (aic) displayed ¿aic power disconnected¿ alarm although console was plugged in. hospital staff checked the outlets, and all were functioning. regardless of which outlet the aic was plugged into, said alarm triggered without resolution. consequently, the console was switched out to another aic with no report or indication of interruption in support. there was no harm to the patient as a result of the event. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported complaint date are consistent with the complaint and show that there were numerous instances of unusual ac power disconnected alarm observed in the logs. these alarms would consecutively appear then resolve themselves consistently. this console was tested after arriving in field service. the reported ac power detection issue could not be reproduced. the console was able to detect ac power as normal while testing in the lab. the only abnormality observed on this console were slightly bent prongs on the power cord plug. the root cause of the issue could not be determined due to the inability to reproduce. extensive testing was conducted to ensure appropriate voltages across the components associated with the ac power system, but no issues were identified. before sending this console back to the customer, field service was instructed to replace the power supply as a precaution, replace the power cord, verify power system functionality, and perform a full functional check on the console and optical system. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the impella 5.5 serial number (b)(6) with date of event 02-jan-2025 and identifier ending in (b)(6).
12116 | 1220648-2025-25662 | 21196687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196687&pc=OZD | 01/02/2025 | 02/14/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584645 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an attempt to implant an impella cp device pre-percutaneous coronary intervention (pci) was unsuccessful. angiogram of the right femoral artery access site was performed, and reportedly, the site appeared suitable for the impella placement although calcium burden was noted. access was obtained via ultrasound and micro-puncture technique. an 8fr dilator utilized, but when the physician tried to insert 10fr dilator, it was unable to advance in a safe manner. multiple attempts were made; however, attempts were unsuccessful. impella cp implant was aborted. the physician placed a terumo 9fr sheath and proceeded with pci. it was reported there was no harm to the patient (demographics unknown) as a result of this event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, impella insertion aborted due to calcium burden as per the clinical description provided. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number and a new code was added. a medical device problem code was added as it was inadvertently omitted from initial manufacturer device report number. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number and were added. | Patient-specific information and the initial reporter's name are unknown, unavailable at this time and respective field(s) reflect "unknown" or have been left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12117 | 1220648-2025-25663 | 21196699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196699&pc=OZD | 01/02/2025 | 03/05/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580246 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that the patient unfortunately suffered from suspected dead bowel and was on support until children could come to visit before expiring. patient expired on support 1/18 at 1830. per the medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome, delay in needed therapy. the user for some reason unknown had difficulty passing the device beyond the innominate/subclavian injunction; there was buckling. there is no allegation regarding manufacturing/quality issue at this time. | The complainant reported that multiple complications were encountered during impella 5.5 support. during delivery of the device, it was noted that resistance was encountered at the innominate/subclavian junction resulting in the catheter kinking. despite the noted kink, the delivery was ultimately successful and support was initiated. roughly three days into support, oozing was noted at the impella access site. heparin was stopped to manage the condition and the patient remains on support. |  |  |  |  | The investigation of access site bleeding and delivery issues has been completed. the impella device was not returned for evaluation. the failure mode: product damage (cannula kink) changed to failure mode: delivery issue since the cannula kink occurred during insertion. the root cause of delivery issue cannot be determined since the complaint device not returned for analysis and insufficient clinical data provided. the root cause of access site bleeding - major issue cannot be determined since the compliant device not returned for analysis and insufficient clinical data provided. b2 outcome revised to reflect death which inadvertently was not included on initial manufacturer device report number 1220648-2025-25663. b5 revised to reflect the patient outcome and mso statement which inadvertently was not included on initial manufacturer device report number 1220648-2025-25663. d6b revised to reflect the explant date which inadvertently was not included on initial manufacturer device report number 1220648-2025-25663. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25663 as it is unknown if the initial reporter also sent the report to the food and drug administration. h1 revised to death which inadvertently was not included on initial manufacturer device report number 1220648-2025-25663. h6 revised to reflect patient outcome of death which inadvertently was not included on initial manufacturer device report number 1220648-2025-25663. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12118 | 1220648-2025-25664 | 21196731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196731&pc=OZD | 01/02/2025 | 02/06/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support and one day into impella 5.5 support, a controller error alarm appeared on the automated impella controller (aic). the waveforms on the top of the screen went blank while the bottom readings remained visible. the issue self-resolved after a few minutes, but the decision was ultimately made to replace the console. there was no harm to the patient. | [5 empty] | The investigation for the optical signal issues has been completed. the console log confirmed the reported failure mode, given there was a controller error alarm triggered during the clinical procedure as well as multiple "algs suspended opm crc error" messages. real time logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25664 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25664 and a new code was added. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
12119 | 1220648-2025-25665 | 21196750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196750&pc=OZD | 01/03/2025 | 02/18/2025 |  | 01/21/2025 |  |  | 01/03/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559307 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that four days into impella cp support, a patient developed significant bleeding from their impella access site. the patient's hemoglobin levels dropped resulting in the patient being given four units of packed red blood cells, one unit of platelets, and one unit of fresh frozen plasma. manual pressure was applied to the site and the heparin drip was discontinued to manage the condition. the pump was subsequently weaned and explanted the following day. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25665. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12120 | 1220648-2025-25666 | 21196781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196781&pc=OZD | 01/02/2025 | 02/05/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584926 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and developed limb ischemia. the patient would eventually expire. the patient was in the intensive care unit (icu) for 12 hours and had issues with blood pressure as the vasopressors were increasing. the decision was made to place impella cp device for added support. the right femoral artery was accessed using micro-puncture and ultrasound. perclose x2 were placed. it was noted 6500 units heparin was given. the impella was successfully implanted, dressing was placed, and patient was sent to icu without incident. however, approximately 4 hours later, the patient's leg became cyanotic, the right ventricle was struggling, and the impella could not run above p4 without suction. right ventricle support was discussed. however, the patient was not expected to survive throughout the night. the patient would eventually expire next morning overnight. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
12121 | 1220648-2025-25667 | 21196783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196783&pc=OZD | 01/03/2025 | 02/28/2025 |  | 01/21/2025 |  |  | 01/03/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589784 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella 5.5. unable to pass through vasculature. procedure aborted due to patient anatomy. there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella was unable to pass through the vasculature. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25667 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
12122 | 1220648-2025-25668 | 21196801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196801&pc=OZD | 01/03/2025 | 02/17/2025 |  | 01/21/2025 |  |  | 01/03/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539563 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the sterile sleeve tore. the pump was exchanged due to the sterile sleeve tear and needing to advance device. there was no reported patient harm. | [5 empty] | The investigation of product damage (sterile sleeve damage) and positioning issues has been completed. the root cause of the product damage (sterile sleeve) issue was not determined since product was not returned. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. e4 should have been left blank on manufacturer device report 1220648-2025-25668. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12123 | 1220648-2025-25669 | 21196812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196812&pc=OZD | 01/05/2025 | 02/28/2025 |  | 01/21/2025 |  |  | 01/05/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589632 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a 5.5 pump being placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. due to anatomic issues the 5.5 implant was aborted and the team instead placed a cp. the patient survived the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25669 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12124 | 1220648-2025-25670 | 21196825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196825&pc=OZD | 01/06/2025 | 02/14/2025 |  | 01/21/2025 |  |  | 01/06/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582295 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had a hemorrhagic bleed to the cerebellum. the patient was taken to the operating room for an external ventricular drain and craniotomy. the patient expired after 150.07 hours of impella support. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25670. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12125 | 1220648-2025-25671 | 21196830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196830&pc=OZD | 01/06/2025 | 02/21/2025 |  | 01/21/2025 |  |  | 01/06/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560112 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when implanting the impella cp initially had great waveforms and flows. the physician repositioned the pump during sheath exchange. shortly after sheath exchanged the waveforms suddenly changed followed by a decrease in flows, ultimately below 1 lpm, and motor current was essentially flat, angiographically pump remained across the valve in what appeared to be an adequate position. the impella cp was exchanged for a new one, successfully. the patient survived the event. | [5 empty] | The investigation for the low pump flow has been completed. the impella device nor the data logs were returned for evaluation. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25671 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 a health effect - impact code was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25671. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12126 | 1220648-2025-25672 | 21196836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196836&pc=OZD | 01/07/2025 | 01/29/2025 |  | 01/21/2025 |  |  | 01/07/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) fell to the floor. the aic was removed from service. there was no patient involvement. | [5 empty] | A regulatory report is no longer necessary since manufacturer device report 1220648-2025-25672 was found to be a duplicate to manufacturer device report 1220648-2025-25487. additional information will be documented in manufacturer device report 1220648-2025-25487. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12127 | 1220648-2025-25673 | 21196862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196862&pc=OZD | 01/07/2025 | 02/18/2025 |  | 01/21/2025 |  |  | 01/07/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565636 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella cp. team was unable to advance sheath all the way due to vascular anatomy and peripheral artery disease (pad). unable to place impella. insertion procedure aborted. the patient remains stable. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the doctors were unable to advance sheath all the way due to vascular anatomy and peripheral artery disease and were unable to place impella and aborted insertion procedure. e4 should have been left blank on manufacturer device report. h6 code 2422 and 4627 were reported incorrectly on manufacturer device report. new codes were added to health effect - clinical code and health effect - impact code.
12128 | 1220648-2025-25674 | 21196876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21196876&pc=OZD | 01/07/2025 | 03/11/2025 |  | 01/21/2025 |  |  | 01/07/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581022 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an impella cp device was implanted preoperatively and two days into impella cp support, the system triggered high purge pressure and purge system blocked alarms. the purge cassette was replaced during troubleshooting, but the issue persisted. heparin dose was increased in response to the condition and ongoing monitoring had been outlined for support. upon follow up, it was verified that the purge issue resolved with the noted intervention. support continued for two more days until the patient suffered a cardiopulmonary arrest while waiting for ventricular septal perforation repair and passed away. medical safety review of the event noted there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. however, the impella is unlikely related to the reported death. the patient was continued on impella support, and the device was not removed because of the purge pressure issue. the patient was subsequently taken to the operating room for the cardiac surgical procedure: ventricular septal perforation repair where he met his demise. | [5 empty] | The investigation of high purge pressure has been completed. there were no data logs returned for the complaint. the clinical details indicated that there was high purge pressure, two days after implant. the cassette was replaced but that did not resolve the issue. the heparin concentration was increased to 50 (u/l) and the purge flow rate improved from 1 milliliter (ml) to 3ml and there was no problem with pump operation. the next day, the patient went into cardiopulmonary arrest which was not related to the impella so the pump was emergently explanted as the patient expired. the returned pump was cut at the catheter and no red handle was returned to confirm the serial number of the pump. there were deformations on the blades of the impeller but none were fractured. regarding the high purge pressure, there was no kinks seen in the returned portion of the catheter. there was a slight film in the purge gap but that was cleared after the pump was soaked. as part of the investigation the catheter was then cut near the motor housing and no kink in the purge line was seen or signs of blood ingress. a syringe with fluid was used in the purge line and fluid expelled very easily from the purge gap. due to lack of sufficient product returned as well as data logs, the root cause of the high purge pressure could not be determined. d.9 date device returned/information was added. e.1 address line 2 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25674. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25674 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25674. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12129 | 1220648-2025-25675 | 21197097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21197097&pc=OZD | 01/14/2025 | 02/27/2025 |  | 01/21/2025 |  |  | 12/26/2024 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025577852 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support due to needing a heart transplant and while on support, the patient had arrythmia issues. the pump position was checked and support continued. additionally, the patient received the transplant and the doctor decided to place pump in surgical mode and reinsert into donor heart for 12-24 hrs. after explant, there was a pulse discrepancy in the upper extremities. the patient was taken for a right brachial/ulnar/radial artery thromboembolectomy. the procedure was successful. staff assumed the clot was dislodged when removing the 5.5. the patient survived the event. | [5 empty] | The investigation of vf/vt/af/at, limb ischemia, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12130 | 1220648-2025-25676 | 21197367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21197367&pc=OZD | 01/07/2025 | 02/25/2025 |  | 01/21/2025 |  |  | 01/07/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589952 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient had a dissection of left external iliac artery related to re-insertion of dilator into 14fr peel-away sheath without wire guidance. the perforation was managed by a covered stent and balloon tamponade. the patient survived the event. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the dilator was not returned for investigation. as per the clinical information provided, the dilator was inserted post pump explant and without any wire support which caused the perforation of the iliac artery. therefore, the root cause of the vascular damage issue was most likely use related to improper technique. h.6 code 3160 was reported inocrrectly on the initial manufacturer device report number 1220648-2025-25676. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25676 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12131 | 1220648-2025-25677 | 21198794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21198794&pc=PYX | 01/16/2025 | 02/25/2025 |  | 01/21/2025 |  |  | 01/02/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548077 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella rp flex support and during support, the placement signal gradually increased over the course of two hours. during this time, the impella flows dropped and eventually were reading 0.0l/m. the decision was made to explant the impella. there was no patient harm. | [5 empty] | B3 date of event was revised as it was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25677. g4 pma/510(k)number was revised as it was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25677. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12132 | 1220648-2025-25678 | 21199085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21199085&pc=OZD | 01/04/2025 | 02/17/2025 |  | 01/21/2025 |  |  | 01/04/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582290 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had limb ischemia. a percutaneous internal bypass was placed and flow was restored. support continued. additionally, while weaning the pump, the patient went into atrial fibrillation requiring addition of vasopressors. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of limb ischemia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia and vf/vt/af/at was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿.
12133 | 1220648-2025-25679 | 21199635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21199635&pc=PYX | 01/04/2025 | 09/27/2025 |  | 01/21/2025 |  |  | 01/04/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561571 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The investigation uncovered placement signal issues. | The complainant reported that following implant, an impella rp flex device experienced a significant decrease in pump flow. it was noted that heparin was not being supplied due to bleeding at the extracorporeal membrane oxygenation site and it was believed a thrombus had entered the pump. heparin was restarted to address the issue and support continued. a couple days later, the patient developed hemolysis. care was withdrawn and the patient passed away. it is unknown if the complications contributed to the patient's passing. |  |  |  |  | The investigation of thrombus, low pump flow, and hemolysis has been completed. the impella device was not returned for evaluation. low pump flow failure mode was changed to placement signal issue since the characteristics determine a false low flow which is a characteristic of sensor drifting. product was not returned for investigation. data logs show while the pump flow was drifting down, the placement signal was observed to be drifting upwards. the placement signal (ps) drifted from 100mmhg to 300mmhg in the span of 90 minutes and then eventually starts to stabilize. motor current, cvp values and optical signal were stable at the time. no psnr alarms were present at the time. these characteristics indicate a drifting dp sensor leading to false flow values. therefore, the root cause of the placement signal issues was determined to be most likely dp sensor drift. data logs were investigated. at the time of the reported failure, there were no motor current spikes or suction alarms observed. placement signal was trending down to normal from the previous rise since last 34 hours. pump flow was seen getting better from 2l/min to 2.5l/min at p9. the expected flow at p9 is (3.3l/min to 3.7l/min). eventually the ps stabilizes and acts similar to cvp and motor current. the patient had received few units of blood which helped the flow to get better. therefore, the root cause of the hemolysis issue was most likely patient condition related. as the necessary information was not provided, the root cause of the thrombus was not determined". | B5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-25679. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12134 | 1220648-2025-25680 | 21201031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21201031&pc=OZD | 01/06/2025 | 03/12/2025 |  | 01/21/2025 |  |  | 01/08/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593395 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an attempted impella cp implant and during the implant, the guidewire bent and kinked. this made access difficult and the implant was aborted. the patient was stable. | [5 empty] | The investigation of product damage (guidewire damage - great vessel) and delivery issues has been completed. during the case, the guidewire bent and was kinked. the guidewire was returned. investigation of the guidewire confirmed kink. the cause of the delivery issue was kinked guidewire due to returned guidewire with visual kink. product damage (guidewire damage - great vessel) failure mode (fm) was removed since there was no fracture to the guidewire. b3 date of event revised to reflect the correct date on manufacturer device report 1220648-2025-25680 g1 reporting contact first and last name was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25680 g2 report source; health professional was missing from manufacturer device report 1220648-2025-25680. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿.
12135 | 1220648-2025-25681 | 21201115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21201115&pc=PYX | 01/05/2025 | 02/24/2025 |  | 01/21/2025 |  |  | 01/05/2025 | 01/21/2025 |  |  | 01/31/2026 |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517159 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, the pump had decreased flows. there were also motor current spikes at 2351 and 2352. the pump was explanted and biomaterial was observed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the data logs were reviewed, and - as reported - there was a biomaterial ingestion signature the night of 1/4. suddenly, there was a motor current (mc) spike above 1600 ma, a speed deviation, step up in dp ps, increase in purge pressure, and drop in pump flows. shortly after, there was another smaller mc spike with speed deviation. the low pump flows were sustained for the remainder of the case. returned product was investigated and biomaterial was seen to be wrapped around the impeller after soaking the pump. clinical details reported that the patient's anticoagulation levels were subtherapeutic leading up to the complication and thrombus was noted in the inflow cage at explant. based on all the above-mentioned details, the root cause of the low pump flow was determined to be biomaterial. thrombus: since the investigation into low pump flow determined that the root cause was biomaterial, the above investigation is representative of the investigation into thrombus as well. g2 report source; study was missing from manufacturer device report 1220648-2025-25681. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12136 | 1220648-2025-25682 | 21202073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202073&pc=PYX | 01/08/2025 | 02/25/2025 |  | 01/21/2025 |  |  | 01/03/2025 | 01/21/2025 |  |  | 03/31/2025 |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024394908 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella rp flex support and during support, the nurse pulled the blanket off the patient. the blanket got caught on the device and pulled the cannula out of the pulmonary artery all the way into the inferior vena cava/femoral vein. the pump was explanted and the patient remained stable. | [5 empty] | The investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post-sterile inspection checks. clinical information provided indicated that the pump was pulled back across valve by nurse during support. blanket caught on the device due and pulled the cannula out of the pulmonary artery, all the way into the inferior vena cava/femoral vein. no product was received for evaluation. the cause of the positioning issue was most likely use related due to the nurse pulling on the blanket causing the pump to dislodge. corrected information provided in e4. e4 should have been left blank on the initial manufacturer device report 1220648-2025-25682. | B3 date of event was revised as it was inadvertently reported incorrectly on manufacturer device report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12137 | 1220648-2025-25684 | 21202118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202118&pc=PYX | 01/05/2025 | 01/21/2025 |  | 01/21/2025 |  |  | 01/05/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025585351 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that roughly ten days into impella rp flex support, the patient developed a gastrointestinal bleed. heparin was present in the purge line at the time of the noted condition and systemic heparin was halted to manage the condition. two units of packed red blood cells were provided to counteract the bleed and the condition stabilized. support with the implanted pump was uninterrupted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12138 | 1220648-2025-25685 | 21202171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202171&pc=OZD | 01/05/2025 | 03/10/2025 |  | 01/21/2025 |  |  | 01/05/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587328 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the purge pressure rose. the purge cassette was replaced, however the issue persisted. the was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. no product/logs were returned. a picture of the noted mark on the catheter was provided, which showed a compressed/flattened 2-3cm length of catheter at the 75 cm marking. the root cause of the high purge pressure was most likely a kinked purge line since the provided image of catheter appears to be kinked and clinical details suggest excessive force being used when securing the pump e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
12139 | 1220648-2025-25686 | 21202248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202248&pc=OZD | 01/07/2025 | 02/07/2025 |  | 01/21/2025 |  |  | 01/07/2025 | 01/21/2025 |  |  | 02/28/2026 |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493142 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support placed for a patient admitted in acute myocardial infarction / cardiogenic shock. the pump had an access site bleed. the bleed was treated with blood products, and the patient was evaluated for a retroperitoneal bleed. the pump remained on for support despite the bleed. the patient survived the event. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. the pump passes all post sterile inspection checks. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12140 | 1220648-2025-25683 | 21202282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202282&pc=OZD | 01/04/2025 | 02/18/2025 |  | 01/21/2025 |  |  | 01/04/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577866 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. a femstop was placed on the access site. heparin was not seen to be given to the patient during support and it is unknown if this was due to the bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 outcomes attributed to adverse event; other serious was selected incorrectly on manufacturer device report 1220648-2025-25683. e4 should have been left blank on manufacturer device report 1220648-2025-25683. h6 code 4614 was reported incorrectly on manufacturer device report1220648-2025-2568. new code was added to health effect - impact code. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12141 | 1220648-2025-25688 | 21202417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202417&pc=OZD | 06/13/2024 | 02/14/2025 |  | 01/21/2025 |  |  | 01/08/2025 | 01/21/2025 |  |  | 02/28/2026 |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473980 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured thrombocytopenia with a possible relationship to the impella 5.5. on (b)(6) 2024, platelet count 105. on (b)(6) 2024, platelet count dropped to 67. heparin-induced thrombocytopenia panel negative. on (b)(6) 2024, platelet count lowest at 60. platelet count steadily rose until on (b)(6) 2024, it was within normal limits. the patient is stable. the patient survived the event. | [5 empty] | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25688. g2 report source; study was missing from manufacturer device report 1220648-2025-25688. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 for use during cardiogenic shock, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12142 | 1220648-2025-25687 | 21202921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21202921&pc=OZD | 01/06/2025 | 03/12/2025 |  | 01/21/2025 |  |  | 01/08/2025 | 01/21/2025 |  |  |  |  | 01/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576818 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that roughly seventeen days into impella cp support, the pump stopped unexpectedly. the device was removed as a result of the noted failure. there was no noted patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the device inserted without anissue. clinical information reported that 17 days later, the pump stop occurred. the pump was explanted 19 days after implant. a service loaner automated impells controller was used so data logs could not be recovered to confirm event dates. however, all reported dates occur more than 8 days after implant date. the design life of cp pumps is 8 days. the product was returned. abnormality was observed in the form of a large purge gap, indicating bearing wear. the cause of the pump stop was most likely bearing wear due to the large purge gap seen on the returned product. the device was used beyond the 8 day cp design life. d.9 date device returned/information was added. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
12143 | 1220648-2025-25656 | 21190803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21190803&pc=OZD | 01/02/2025 | 02/07/2025 |  | 01/20/2025 |  |  | 01/02/2025 | 01/20/2025 |  |  |  |  | 01/20/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580384 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for a bridge to transplant and approximately 20 days after start of support, the patient's sterile sleeve was observed to be damaged. prior the white blood cells count was raised even with all lines being removed and no signs of infection. the graft had been cultured. the patient was placed on antibiotics. the decision made to remove right axillary 5.5 and graft and place new left axillary 5.5 and graft as the patient still awaits transplant (listed as status 2). | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issues has been completed. device history record review confirmed the pump passed all post sterile inspection checks. no device was received for evaluation. the root cause of the product damage (sterile sleeve damage) was not determined as no product or images were returned. because insufficient clinical evidence was provided, the root cause of the infection was unable to be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12144 | 1220648-2025-25632 | 21178050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178050&pc=OZD | 01/02/2025 | 02/14/2025 |  | 01/17/2025 |  |  | 01/02/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591367 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a younger patient admitted in with acute myocardial infarction, cardiogenic shock and ventricular fibrillation arrest. the pump had persistent low flows that were treated by volume and blood delivery. additionally, the patient had a sheath placed for limb perfusion. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow and limb ischemia has been completed. the product and data logs were not returned for review. based on clinical description, the root cause of low flow was most likely patient condition related. the root cause of the limb ischemia could not be determined without additional information.
12145 | 1220648-2025-25633 | 21178109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178109&pc=OZD | 12/31/2024 | 02/18/2025 |  | 01/17/2025 |  |  | 12/31/2024 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574645 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump being delivered to support a patient admitted in acute myocardial infarction/cardiogenic shock. the patient's anatomy prompted a 14 french sheath to be delivered at a 90 degree angle. the angulation caused the sheath to kink and the pump failed to be delivered and the team observed the inlet and pigtail to become uncoiled and remain in the sheath. the entire system was removed. the pump was removed and a portion remained in the sheath. the long sheath was instead placed with a new pump. there was no patient harm. | [5 empty] | The investigation of product damage (detached inflow/outflow - peripheral vessel) and introducer damage ¿ mechanical has been completed. the product was returned for investigation. a data log review was not required for this case. the returned pump inlet cage was separated from the cannula, hanging on unwind nitinol coil inside the introducer, where a kink was reported. no further investigation was required. as per the given clinical information, due to the patient's anatomy, the doctor had to access at nearly a 90-degree angle, causing a kink in the impella peel-away sheath, which prevented the catheter from advancing. when removing the impella, the inlet and pigtail uncoiled and remained in the sheath, while the motor and cannula were removed. the sheath was taken out with the inlet and pigtail still inside. the cause of the product damage (detached inflow cage) issue was use related to applied excessive force, based on returned product investigation. the cause of the introducer damage issue was use related to steep angle of insertion due to patient anatomy. | The pump and introducer have been returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12146 | 1220648-2025-25634 | 21178208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178208&pc=OZD | 10/21/2024 | 02/06/2025 |  | 01/17/2025 |  |  | 12/23/2024 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522064 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella cp device used: ¿the patient developed hemolysis. admission total hgb 15.1 g/dl, nadir 9.5 g/dl ((b)(6)2024). peak ldh 1130 u/l ((b)(6) 2024). plasma free hgb remained within normal limits, as did total bilirubin¿. | [5 empty] | The investigation of hemolysis has been completed. data logs showed the pump was in improper position. this event remained for about 1.25 hours then resolved. no other issues were found on the logs. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12147 | 1220648-2025-25635 | 21178299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178299&pc=OZD | 10/23/2024 | 02/06/2025 |  | 01/17/2025 |  |  | 12/23/2024 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517669 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured recurring non-sustained ventricular tachycardia (vt) with a "possible" relationship to the impella device used: ¿the patient developed recurring episodes of non-sustained ventricular tachycardia and frequent premature ventricular contractions.¿ it was reported that the patient/event has not recovered/not resolved. also, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿the patient developed hemolysis. admission total hemoglobin 14.7 g/dl, nadir 9.0 g/dl ((b)(6) 2024); peak ldh 2867 u/l ((b)(6) 2024); admission total bilirubin 2.0 mg/dl, peak 6.6 mg/dl ((b)(6) 2024).¿ it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation for the hemolysis and the tachycardia has been completed. data logs were reviewed which confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in ventricle/aorta. the next day, pump was in proper position. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of the tachycardia could not be determined without additional information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12148 | 1220648-2025-25636 | 21178485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178485&pc=OZD | 01/02/2025 | 02/13/2025 |  | 01/17/2025 |  |  | 01/02/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575915 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant and it was aborted due to patient anatomy. multiple attempts were made to get the impella in. however, they were unsuccessful and the implant was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because insertion was aborted due to patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12149 | 1220648-2025-25637 | 21178511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178511&pc=OZD | 01/02/2025 | 02/03/2025 |  | 01/17/2025 |  |  | 01/02/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1725570 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) displayed purge volume low alarm. the nurse reported purge fluid bag change, cassette change, and de-airing did not resolve alarm. troubleshooting completed. aic transfer resolved the issue and there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. according to the device and data analysis the root cause of this issue was likely a defective purge assembly that was misreporting purge fluid volume. h8 usage of device was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25637
12150 | 1220648-2025-25638 | 21178638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178638&pc=OZD | 08/31/2024 | 02/10/2025 |  | 01/17/2025 |  |  | 12/23/2024 | 01/17/2025 |  |  | 09/30/2025 |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444721 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella device used: ¿the patient developed hemolysis. admission total hemoglobin 11.1 g/dl, nadir 7.2 g/dl ((b)(6) 2024); peak ldh 1314 u/l ((b)(6) 2024); admission total bilirubin 0.9 mg/dl, peak 1.5 mg/dl ((b)(6) 2024).¿. the patient recovered with sequelae. loqi also reported that the patient endured bleeding at impella cp access site right femoral artery with a "definite" relationship to the impella device used: ¿the patient developed bleeding from the right femoral artery at the impella cp access site. admission total hgb 11.1 g/dl, nadir 7.2 ((b)(6) 2024). heparin was held due to bleeding, femostop was applied, and dressings were changed multiple times.¿. the patient recovered with no sequelae. it was noted that care was withdrawn and the patient expired. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issues has been completed. device history record review confirmed that the pump passed all post sterile inspection checks. data logs were reviewed, and it was confirmed that no suction alarms occurred. no motor current spike was observed out of p-level change. many impella position in aorta alarms were observed around the time issue reported followed high purge pressure alarms and flow was about 3l/min at p-2. the root cause of the reported issues cannot be determined since the complaint device not returned for analysis and insufficient data provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12151 | 1220648-2025-25639 | 21178660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178660&pc=OZD | 12/31/2024 | 02/21/2025 |  | 01/17/2025 |  |  | 12/31/2024 | 01/17/2025 |  |  | 01/31/2026 |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468501 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and following the implant, same day, it was noted the purge pressure was rising and high the following day. lysis therapy was initiated to counteract the condition and support continued uninterrupted. approximately four days later, a computed tomography scan noted that the patient had experienced a stroke. the patient's condition was monitored until planned explant four days later which was successfully completed. | [5 empty] | The investigation for the stroke and the high purge pressure has been completed. there were no data logs returned. the product was returned, and high purge pressure was reproduced. there was a small amount of biomaterial seen in the purge gap as well as around the impeller. more notably, a kink in the catheter was seen near the motor housing. high purge pressure reproduced when attempting to purge the device with the provided cassettes. the catheter was stripped near the kink to see if there was a kink in the line. there was a significant kink in the purge line. once the catheter was cut on the red handle side, purge fluid was able to expel from the catheter. the catheter on the motor side was then cut to remove the kink in the line and purge fluid was pushed through the purge line with a syringe and fluid expelled from the purge gap. based on the returned product, the root cause of the high purge pressure was most likely due to a kink in the purge line. the root cause of the stroke could not be determined without additional information. added d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12152 | 1220648-2025-25640 | 21178796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178796&pc=OZD | 03/03/2024 | 02/03/2025 |  | 01/17/2025 |  |  | 12/23/2024 | 01/17/2025 |  |  | 10/31/2025 |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406477 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured non-sustained ventricular tachycardia with a "possible" relationship to the impella device used: ¿the patient experienced 18 beats of non-sustained ventricular tachycardia, which activated the aicd to shock the rhythm back to the intrinsic atrial fibrillation rhythm. echo confirms impella 5.5 placement, electrolytes were normal but potassium and magnesium were replaced regardless. concern for likely relatedness of progression of stage d heart failure.¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. h.5 revised as previously entered incorrectly on initial manufacturer device report number 1220648-2025-25640. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12153 | 1220648-2025-25641 | 21178908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21178908&pc=OZD | 01/02/2025 | 05/16/2025 |  | 01/17/2025 |  |  | 01/02/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484305 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the impella stopped, displayed impella defective alarm flashing bright and dark yellow, soft buttons were unresponsive. the impella cp was replaced and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop has been completed. the root cause of the pump stop is due to an open data cable line a the joint between the patient cable pin 6 and the grey line in the patient cable as was found with testing of returned product and data review. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12154 | 1220648-2025-25643 | 21179117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21179117&pc=OZD | 01/05/2025 | 02/17/2025 |  | 01/17/2025 |  |  | 01/05/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant but the wire kept coming out of the left ventricle during cardiopulmonary resuscitation (cpr). the patient went into pulseless electronic activity and all attempts to save the patient stopped. the patient never regained a pulse and the doctor was unable to recross the aortic valve. the patient expired after 15-20 minutes of cpr. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk. percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature.
12155 | 1220648-2025-25644 | 21179191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21179191&pc=OZD | 01/01/2025 | 02/12/2025 |  | 01/17/2025 |  |  | 01/01/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592042 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following removal of the peel away sheath, the patient began to bleed from their access site. forward tension sutures and figure 8 sutures were placed, but the bleed persisted. hemostasis was finally achieved when femstop was applied to the site. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
12156 | 1220648-2025-25642 | 21179293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21179293&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/17/2025 |  |  | 12/21/2024 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580399 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while in the intensive care unit, an echocardiogram was done and noted tamponade. the patient was emergently taken back to cardiovascular operating room for chest exploration and direct aortic approach 5.5. the 0.035 wire was removed and was replaced with 0.018 wire. diagnostic catheter was removed and the impella was backloaded onto 0.018 guidewire and placed into left ventricle using double barrel technique. after successful repair of left ventricle perforation and successful implant of impella 5.5 direct aortic approach chest packed sternum and skin closed. the patient eventually expired after care was withdrawn. | [5 empty] | The device or the data logs were not returned for analysis. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to small anatomy of the patient. the root cause of ventricle perforation is not determined because the timing of the event is unclear as per the provided clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12157 | 1220648-2025-25645 | 21179507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21179507&pc=OZD | 12/26/2024 | 02/21/2025 |  | 01/17/2025 |  |  | 12/26/2024 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588008 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp device was pre-operatively implanted in a patient for mechanical circulatory support during a high-risk surgery and high purge pressure was encountered. the patient was presented to the hospital due to suspected acute coronary syndrome and multiple branch lesions. ventricular septal perforation was discovered and the impella cp device was implanted. immediately after the impella cp pump was turned on, the purge blockage alarm sounded. the purge cassette was replaced without resolution. consequently, the impella cp device was explanted and replaced. the patient remained on support for approximately nine days before expiring due to multiple organ failure. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. the first 30 minutes temporary impede of adequate support with the initially placed pump cannot be ruled out as a contributing factor. however, it is noted, recognized the replacement pump was inserted (used without issues), and the patient subsequently died nine days later because of multi-organ failure due to other premorbid conditions. | [5 empty] | The investigation for the high purge pressure has been completed. the data logs indicate there were motor current spikes one minute after starting the pump and the purge pressure immediately rises. it rises for 8 minutes until it triggers high purge pressure alarms. troubleshooting is done by replacing the cassette, but it remains high. the case lasts for 32 minutes and then the pump is explanted. data logs confirm priming was complete and the pump is started a minute after. the returned pump had biomaterial in the purge gap. high purge pressure reproduced. there was a blockage isolated to the motor and the purge line was seen to be an amber color indicating there was blood in the line. the pump was torn down and biomaterial was seen inside the motor. the root cause of the high purge pressure is due to biomaterial based on data log and product analysis. instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12158 | 1220648-2025-25646 | 21180231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180231&pc=OZD | 01/02/2025 | 02/04/2025 |  | 01/17/2025 |  |  | 01/02/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583400 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient experienced oozing from their impella site and a small hematoma developed. manual pressure was applied and the dressing was changed on multiple occasions to manage the condition. one unit of blood was given to the patient and support with the implanted pump remained uninterrupted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12159 | 1220648-2025-25647 | 21180533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180533&pc=OZD | 01/03/2025 | 02/14/2025 |  | 01/17/2025 |  |  | 01/03/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has a controller error alarm. the aic was removed from service and returned. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console(aic) controller error-yellow alarm has been completed. data logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. there were 26 opm crc errors noted during this case, with one instance resulting in a controller error (took long enough to resolve that the error condition triggered). there were also bad file descriptors next to these opm crc errors that indicates the epc was not seeing successful communication to the rlm. reviewing the rlm logs for these times, the rlm was rebooted which would have caused the aic to see unsuccessful data transfer to the rlm until the usb was restarted. console was sent back for investigation but the loss of optical data failure mode was not reproduced. the exact same repeated rlm reboots was not reproduced but the rlm was noticed to have an abnormal boot sequence which likely points to partition switching and usd card corruption. additionally, the module would not connect to ic_test and would power cycle when attempted to which is another symptom of corruption. during testing, the serial communication for live monitoring of rlm logs also failed and would not recover. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
12160 | 1220648-2025-25648 | 21180682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180682&pc=OZD | 08/06/2024 | 02/07/2025 |  | 01/17/2025 |  |  | 12/20/2024 | 01/17/2025 |  |  | 10/31/2025 |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413176 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | Heart Failure/Congestive Heart Failure | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured cardiogenic shock and 3.5 months following this event, the patient endured heart failure with a "probable" relationship to the impella device used. the patient was implanted with an impella 5.5 pump for mechanical circulatory support, impella assisted intervention to the left internal mammary, left anterior descending, and saphenous vein graft-obtuse marginal arteries. approximately eight months following explant, the patient presented to the hospital with an out of the hospital cardiac arrest complicated by cardiogenic shock requiring intubation and sedation with treatment. eleven months following explant, reportedly the patient experienced heart failure. no further details were provided. it is unknown how the device may have contributed to the events and the patient's outcome is unknown as well. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause was not determined.
12161 | 1220648-2025-25649 | 21180690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180690&pc=OZD | 01/05/2025 | 02/07/2025 |  | 01/17/2025 |  |  | 01/05/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583430 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced some oozing and developed a hematoma around their impella cp access site following implant. manual pressure was applied and the perclose devices/sutures were tightened to manage the condition. support was uninterrupted. the patient survived the event. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post sterile inspection checks. the root cause of the access site bleeding was most likely patient condition since hematoma occurred after arriving to intensive care unit (icu). | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12162 | 1220648-2025-25650 | 21180803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180803&pc=OZD | 01/06/2025 | 02/21/2025 |  | 01/17/2025 |  |  | 01/06/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580948 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to experience low flows one day into patient support. the pump was subsequently replaced to resolve the noted failure. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. the product and data were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12163 | 1220648-2025-25651 | 21180846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180846&pc=PYX | 01/01/2025 | 04/11/2025 |  | 01/17/2025 |  |  | 01/07/2025 | 01/17/2025 |  |  | 01/31/2026 |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517161 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Renal Failure | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that the impella rp motor spiked and flows decreased during a procedure the patient was undergoing. the impella was removed. there was no harm to the patient. | Additional information was received on (b)(6)2025 via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of "thrombus in device": no adverse event was reported as a result of the event. | Additional information was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening renal failure with a "probable" relationship to the impella device used: ¿(b)(6) 2025 creatinine peaked to 2.5. crrt started, converted to prrt. (b)(6) 2025 last hd and creatinine dropped to 0.8¿. the patient recovered with no sequelae. | Additional information was received via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of "thrombus in device". no adverse event was reported as a result of the event. |  |  | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The following updates were made to manufacturer device report 1220648-2025-25651 to reflect new information received. based on new information received b1 updated to both adverse event and product problem. b5 updated with additional information based on new information received h1 type of reportable event was updated to serious injury h6 health effect - clinical code updated to reflect new information g2 report source; study was missing from manufacturer device report 1220648-2025-25651. | B.5 additional information was received. h.6 added health effect - clinical code to reflect new information. b.3 revised date of event due to new information received since the previous manufacturer device reports number 1220648-2025-25651 have been submitted. | The following information was updated to reflect additional information received. b.1 updated to both adverse event and product problem. b.2 updated to reflect other serious. b.5 updated with additional information. h.1 type of reportable event updated to serious injury. h.6 added health effect - clinical code and health effect - impact code to reflect new information d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25651. g.2 study selection was added to report source as it was missing from the initial manufacturer device report number 1220648-2025-25651.
12164 | 1220648-2025-25652 | 21180853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21180853&pc=OZD | 01/06/2025 | 02/07/2025 |  | 01/17/2025 |  |  | 01/06/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient was found to have no pulses in the lower extremity and there was significant bleeding from groin site. surgeon stated the patient was clamped down on pressors all day and likely clotted off the leg due to vasoconstriction. the pump was explanted and replaced. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12165 | 1220648-2025-25653 | 21181861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21181861&pc=OZD | 01/03/2025 | 02/20/2025 |  | 01/17/2025 |  |  | 01/03/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568904 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on day one of impella cp support developed bleeding from the impella cp access site. hemoglobin dropped and the patient received one unit of blood products. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since the act was above the recommended range which is between 160-180. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12166 | 1220648-2025-25654 | 21182217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21182217&pc=OZD | 01/06/2025 | 02/24/2025 |  | 01/17/2025 |  |  | 01/06/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479162 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reporting that the connection between the sterile sleeve and the lock mechanism closest to the red plug of an impella 5.5 pump was damaged. support with the implanted pump remained uninterrupted. there was no noted patient harm. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. the pump was not returned for investigation. pictures of the tear were not returned as well. without the pump or the pictures, the failure cannot be further investigated. therefore, the root cause of the product damage issue was not determined since product was not returned and insufficient clinical information was provided. d6b explant date added d4 expiration date was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25654. d4 unique identifier (udi) # was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25654. h4 device manufacture date was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25654.
12167 | 1220648-2025-25655 | 21182499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21182499&pc=OZD | 01/03/2025 | 02/12/2025 |  | 01/17/2025 |  |  | 01/03/2025 | 01/17/2025 |  |  |  |  | 01/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126030 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller failure alarm. the aic was exchanged. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the root cause for controller failure was not determined due to the inability to reproduce during the device analysis and data review.
12168 | 1220648-2025-25610 | 21167707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21167707&pc=OZD | 01/01/2025 | 02/18/2025 |  | 01/16/2025 |  |  | 01/01/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591308 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported being unable to pass impella cp due to narrowing in left femoral artery (lfa). ballooned lfa and switched to long impella sheath to ease impella access. impella inserted over .018 and activated on auto at 1154 with 3.5 flow. after impella was placed in left ventricle, some lower flows were observed, and a kink was noticed on impella above the elbow. the physician removed and replaced with a new impella. there was no harm to the patient. | [5 empty] | The investigation into the reported issues has been completed. the returned device was visually inspected. cannula kink was observed near outflow cage. no biomaterial and no broken blade observed. the root cause of the reported issues was the patient's anatomy related since it was not possible to pass the impella due to narrowing in left femoral artery/patient's vessel condition. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12169 | 1220648-2025-25611 | 21167810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21167810&pc=OZD | 12/26/2024 | 04/03/2025 |  | 01/16/2025 |  |  | 12/26/2024 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563799 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella cp support was decompensating as impella was not offloading the left ventricle. the cp was exchanged for the impella 5.5. the patient is stable and survived the event. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to the patient's death. thus, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. no product was returned. there was no low pump flow seen in the data logs. no motor current spike, placement signal trend, or positioning alarms seen. the logs only go from 12/25 2:44 to 12:25 3:19. due to insufficient information, the root cause of the low pump flow cannot be determined. | B.1 revised from the initial manufacturer device report number 1220648-2025-25611 to reflect both adverse event and product problem. b.2 revised from the initial manufacturer device report number 1220648-2025-25611 to reflect the patient outcome death and added date of death. b.5 added patient outcome information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25611. d.6b revised explant date was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25611. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25611 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 type of reportable event was revised from the initial manufacturer device report number 1220648-2025-25611 to reflect the patient outcome of death. h.6 health effect - impact code 1802 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25611.
12170 | 1220648-2025-25612 | 21168163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21168163&pc=OZD | 11/18/2024 | 06/22/2025 |  | 01/16/2025 |  |  | 01/02/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849082 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and intermittent placement signal was encountered throughout the case with stable motor current. reportedly there were no issues encountered during the pci, and the patient was reported to be in stable condition. the investigation of the impella cp device was completed and it was determined event may be due to the automated impella controller (aic). therefore, a request has been made for return of the aic evaluation and analysis. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) loss of opm data has been completed. console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following errors: "algs suspended opm signal invalid, and "invalid_bad_snr_sample_mask" due to snr falling below the minimum threshold. the first placement signal not reliable alarms (#1036, due to low signal-to-noise ratio) alarm was observed just after 6 minutes of pump connected, pump was removed and reconnected into the console, however multiple instances of placement signal not reliable alarms (#1036, due to low signal-to-noise ratio) were observed. ltlog data indicates that the unreliable placement signal persisted for half of the case, with varying intensity, which correlated with low snr. this led to placement signal not reliable alarms (#1036, due to low signal-to-noise ratio) rtlog data also showed pulsating contrast values and snr values. console was not returned to field service for further investigation. the root cause for the placement signal not reliable alarms, could not be determined as aic was not returned. corrections to the initial mfr report: d4 updated model number. | The investigation of automated impella controller (aic) - loss of opm data has been completed. the console was received for investigation. data analysis: the console logs from the reported day of the event are consistent with the complaint. the logs show multiple instances of the following errors: "algs suspended opm signal invalid, and "invalid_bad_snr_sample_mask" due to signal not reliable (snr) falling below the minimum threshold. the first placement signal not reliable alarms (1036, due to low signal-to-noise ratio) alarm was observed just after 6 minutes of pump connected, pump was removed and reconnected into the console, however multiple instances of placement signal not reliable alarms (1036, due to low signal-to-noise ratio) were observed. ltlog data indicates that the unreliable placement signal persisted for half of the case, with varying intensity, which correlated with low snr. this led to placement signal not reliable alarms (1036, due to low signal-to-noise ratio) real time (rt) log data also showed pulsating contrast values and snr values. device analysis: the console was returned for further investigation. the reported complaint could not be reproduced when running a known-good pump and purge cassette. however, upon inspecting the analog output of the ccd from the optical bench, intermittent distortions in the signal (envelope/fringe) were observed. when the ccd array signal was not distorted, the measured "5s" parameters were within specifications. root cause: the root cause for the placement signal not reliable alarms was due to the defective optical bench or its components. d9 device returned to manufacturer date added. h6 type of investigation, investigation findings, and investigation conclusions revised based on the returned console investigation results.
12171 | 1220648-2025-25613 | 21169032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21169032&pc=OZD | 01/01/2025 | 02/12/2025 |  | 01/16/2025 |  |  | 01/01/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580984 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, no distal doppler pulses were noted. the impella cp was replaced with impella 5.5 to prevent issues to the leg. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25613 and was added.
12172 | 1220648-2025-25614 | 21169655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21169655&pc=OZD | 01/03/2025 | 02/18/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533935 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the impella sheath was inserted with large amount of bleeding when attempting to do single access technique. the intervention for the bleed is unknown. care was withdrawn after 24.05 hours of support, and the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12173 | 1220648-2025-25615 | 21170309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170309&pc=OZD | 01/02/2025 | 10/24/2025 |  | 01/16/2025 |  |  | 01/02/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563966 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella cp support was oozing from the impella access site. systemic heparin was held, and one unit of blood was administered. the patient continued oozing, so the cp was removed and impella 5.5 was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was returned for evaluation. the root cause of the access site bleeding major was not definitively determined as limited clinical information was provided and no issue with returned product was observed d9 device returned to manufacturer date added. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is determined to be anticoagulation management because the partial thromboplastin time values were very high compared to the therapeutic values during the time of bleed. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25615 and was added. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25615. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12174 | 1220648-2025-25616 | 21170375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170375&pc=OZD | 01/01/2025 | 02/06/2025 |  | 01/16/2025 |  |  | 01/01/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582184 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had oozing/bleeding at the access site. a femstop was applied to the access site and anticoagulant was withheld. care was withdrawn after 141.18 hours of impella support and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12175 | 1220648-2025-25617 | 21170472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170472&pc=OZD | 11/29/2024 | 02/18/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572780 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An impact study was received that found that the patient had a hematoma at the access site, premature ventricular contractions (pvc) and intermittent ventricular tachycardia (vt), and ventricular fibrillation (vf) arrest. the impact study found that there was a relationship between the patient's injuries and the impella or impella procedure. for the patient's hematoma, one unit of red blood cells was given to the patient. the incision was reopened, and a moderate amount of clot was expressed. the site was irrigated with antibiotic irrigation. the incision was approximated with two nylon sutures. an iodoform wick was placed through the incision. the wounds were dressed. for the patient's vt, the patient was given intravenous amio, lido and defibrillation to sinus rhythm. when the patient went into vf arrest, the patient was defibrillated and the flow on the impella was decreased. the patient survived the event. | [5 empty] | The investigation of access site bleeding, vf/vt/af/at, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
12176 | 1220648-2025-25618 | 21170547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170547&pc=OZD | 01/02/2025 | 02/05/2025 |  | 01/16/2025 |  |  | 01/02/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was placing an impella cp to support a patient admitted in with st-elevation myocardial infarction. the patient was placed on impella cp support via the femoral and was found to have limb ischemia. the ischemia prompted the pump to be explanted after just over two hours. the patient outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
12177 | 1220648-2025-25619 | 21170628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170628&pc=OZD | 01/04/2025 | 02/17/2025 |  | 01/16/2025 |  |  | 01/04/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583369 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had limb ischemia. the patient was unstable and expired after 70.70 hours of impella support. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12178 | 1220648-2025-25620 | 21170719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170719&pc=OZD | 01/02/2025 | 02/20/2025 |  | 01/16/2025 |  |  | 01/02/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606114 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump being placed for the support of a patient admitted in with acute myocardial infarction and cardiogenic shock. the pump's .018 wire/impella cp was placed within the left ventricle (lv) and the team had ventricular tachycardia (vt) that required shock. the shock caused the vt to resolve. the impella cp was placed and support was initiated but the team caused the pump to become malpositioned and pulled back across the valve. the medical doctor scrubbed back in to redeliver the pump to the lv but the pump kinked and this prompted the team to replace the pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of vf/vt/af/at, positioning issues, and product damage (cannula kink) has been completed. returned complaint device visually inspected. cannula kink observed near outlet and inlet cages. no other issues observed. the root cause of the positioning issue was use related since the scrub tech went to dress and clean site, in doing so impella catheter was pulled back across the valve and the touhy had not been tightened all the way. the root cause of the product damage (cannula kink) was use related since the wireless re-access attempted, resulted in cannula kink. as the necessary information was not provided, the root cause of the vt/vf was not determined. g1 reporting contact fax number was inadvertently omitted from manufacturer device report 1220648-2025-25620. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12179 | 1220648-2025-25621 | 21170916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170916&pc=OZD | 01/03/2025 | 01/30/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197227 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic), while on patient use, fell over and hit the ground. the impella connect board came off. the impella console was sitting on its stand and as the nurse raised up the patient¿s bed, the bed caught the underside of the console and pushed the console off the stand. this happened because the bolts that hold the console to the stand were not in the proper position to hold the console and the stand together. after the fall, there was no interruption in the function of the impella and the patient had no adverse effects. the bedside nurse and the charge nurse switched out the patient¿s aic successfully. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | After internal review it has been discovered that manufacturer device report 1220648-2025-25621 is a duplicate of manufacturer device report 1220648-2025-25575. no further information will be reported in manufacturer device report 1220648-2025-25621. see manufacturer device report 1220648-2025-25575 for any further information.
12180 | 1220648-2025-25622 | 21170987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21170987&pc=PYX | 01/03/2025 | 02/19/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025570397 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella rp flex support and during the implant, sequential dilation was performed and while inserting the 23 french sheath, resistance was met. the sheath would not advance so a 23 french dilator was used and the sheath was inserted. the sheath had perforated the internal jugular vein. the sheath was removed and manual pressure was held. insertion was then performed through the left femoral access. the patient expired after 34.24 hours of impella support. | [5 empty] | The investigation has been completed. an introducer and a dilator were returned for evaluation. upon visual inspection, damage to the tip of the introducer sheath was observed. additionally, the tip of the dilator also showed damage. clinical details noted that when inserting the sheath into the rij, resistance was met and the physician pushed in order to attempt to place the introducer sheath. the patient was noted to have unique neck anatomy with a short neck and deep tissue. the sheath was able to be placed successfully in left ij. the root cause of the vascular damage - peripheral vessel was most likely due to introducer sheath tip split. d.9 date device returned/information was added. | G6 type of report; 30 day was inadvertently not selected manufacturer device report 1220648-2025-25622-1. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12181 | 1220648-2025-25623 | 21171552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21171552&pc=OZD | 01/04/2025 | 02/20/2025 |  | 01/16/2025 |  |  | 01/04/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592118 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [5 empty] | CONTINUOUS RENAL REPLACEMENT THERAPY. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that one day into impella cp support, a patient developed hemolysis attributed to hyperlipidemia from sedative and the patient's chronic pathology. continuous renal replacement therapy was initiated shortly after and the decision was ultimately made to explant the pump the following day. the patient survived the event. | [5 empty] | The investigation for the hemolysis and the renal failure has been completed. the pump was not returned for investigation. data logs were returned and investigated. there were numerous "impella position unknown" alarms and "impella in aorta" alarms observed. no other motor current spike or abnormalities were found. clinical information mentioned that the pump was pulled back seven centimeters and was repositioned later to remove slack to the three centimeter mark. pump repositioning did not resolve the hemolysis and the patient experienced hemolysis for the next few days until explant. the root cause of the hemolysis issue was not determined since the product was not returned and due to insufficient evidence per log and clinical review. the cause of the renal failure issue was not determined since the product was not returned and due to insufficient evidence per log and clinical review. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12182 | 1220648-2025-25582 | 21171954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21171954&pc=OZD | 12/31/2024 | 03/20/2025 |  | 01/16/2025 |  |  | 12/31/2024 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566820 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and following the implant would experience an access site bleeding complication. the patient came in for septic shock versus cardiogenic shock and coded in the emergency department. the intubated patient coded and was sent to the catheterization lab and coded again on the table. one shock was delivered. the patient was in cardiogenic shock and was found to have multiple vessel disease. an impella cp device was immediately inserted via the right femoral artery. a temporary pacer wire was placed via the right femoral vein, and the physician intervened the circumflex with a stent, right coronary with two stents and ramus (transluminal angioplasty) arteries. however, it was reported the physician was unable to intervene the chronic totally occluded left anterior descending artery. following intervention, the temporary pacer wire was removed, a swan-ganz catheter was inserted, and the patient was transported to the unit remaining on impella mcs. a short time thereafter, the right groin was oozing. manual pressure was applied, and a femostop was added to control the bleeding. of note, anticoagulation was not used during support, and it is unknown if this was due to the patient bleeding. the next day, the femostop remained applied, oozing stopped, and the patient had developed a fever. the physician and family discussion occurred and the decision was made to change the patient's status to do not resuscitate. supportive care was continued. the impella was running at support level p-7. however, the following day the family decided on comfort care, and the patient expired. | [5 empty] | The statement in the b5 narrative "the user facility reported a patient" refers to a physician in a hospital setting with a patient that experienced an adverse event while an impella device was in use. the physician reported the adverse event directly to abiomed. there was no user report that was submitted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12183 | 1220648-2025-25624 | 21174920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174920&pc=OZD | 12/30/2024 | 02/12/2025 |  | 01/16/2025 |  |  | 12/30/2024 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was a question of if the patient had limb ischemia. the patient's pulses were absent on their right leg. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 revised as some information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25624. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25624 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12184 | 1220648-2025-25625 | 21174922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174922&pc=PYX | 12/31/2024 | 10/29/2025 |  | 01/16/2025 |  |  | 01/01/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025546008 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. lower than expected flows were noted towards the end of therapy. there was no patient harm. | The investigation uncovered placement signal issues. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product returned. data review: data logs show that on the last day of support the placement signal (ps) begins to drift upwards from around 60 to 250mmhg. flows thus decrease. motor current (mc) and central venous pressure (cvp) are stable during this time. this indicates dp sensor drift. there are no other log abnormalities. this sensor issues causes the flows to be wrongly displayed as low to the user, but stable mc indicates pump to be flowing. the cause of the placement signal issues was most likely dp sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. e4 should have been left blank on the initial report.
12185 | 1220648-2025-25626 | 21174926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174926&pc=OZD | 01/02/2025 | 02/14/2025 |  | 01/16/2025 |  |  | 01/02/2025 | 01/16/2025 |  |  | 07/31/2025 |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442860 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was attempted to be inserted but unsuccessful. the patient had calcified iliac¿s. the case was aborted and there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to inserted due to calcification in the iliac vessels. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12186 | 1220648-2025-25627 | 21174927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174927&pc=PYX | 01/03/2025 | 05/23/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575358 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Difficult to Remove | Material Separation | [5 empty] | Hemolysis | Hematuria | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that on day four of impella rp support labs were hemolyzing and urine was bright red. additionally, echocardiography was done and showed thrombus around the inlet up into the internal jugular and interfering with rp function. decision was made to remove rp. prior to rp removal physician decided to place 29fr cannula in, alongside the rp cannula, entering the right subclavian and terminating in the pulmonary artery. there was difficulty passing the device alongside the rp motor housing and cannula. physician was reminded that the rp catheter was 22 fr. there was concern for both devices in the same vessels and pathways. physician went to remove rp, and the catheter detached from the motor housing. the patient¿s sternum was already open, vascular surgery was consulted and decision made to remove the rp cannula through an aortotomy in the right atrium. rp flex cannula removal was successful. the patient survived the event. | [5 empty] | The investigation results were amended for product damage (detached inflow/outflow - peripheral vessel) and thrombus investigation results were added. the investigation results for hemolysis was previously reported in manufacturer device report number 1220648-2025-25627-1. the failure mode removal issues was changed to product damage (detached inflow/outflow - peripheral vessel) the root cause of the product damage (detached inflow/outflow - peripheral vessel) issue was most likely due to a use-related issue as evidenced by clinical details of 29f cannula placed alongside pump. as the necessary information was not provided, the root cause of the thrombus was not determined. b1 was inadvertently left blank on manufacturer device report number 1220648-2025-25627-1. b2 was inadvertently left blank on manufacturer device report number 1220648-2025-25627-1. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25627 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 medical device problem code 1562 was inadvertently omitted on manufacturer device report number 1220648-2025-25627-1. h6 codes 3331, 114, and 50 were erroneously entered on manufacturer device report number 1220648-2025-25627-1. new codes entered investigation findings and investigation conclusions. | Additional information provided in h6 and h10. the investigation into the reported issue shas been completed. pump was not received for investigation. device history record review confirmed that the pump passed all post-sterile inspection checks. data logs were investigated and there was a motor current spike followed by low flows at the time of reported hemolysis. purge pressure was also observed to be rising a little. these are symptoms of biomaterial buildup. clinical information mentions that echo showed thrombus at the inlet. therefore, the root cause of the hemolysis issue was most likely biomaterial. also, it was noted that the catheter was detached from the motor housing while removing the pump. the physician had placed 29f cannula alongside the rp pump in the same vessel which made the explant difficult. the cause of the product damage (detached inflow/outflow - peripheral vessel) issue was not determined since product was not returned for investigation. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12187 | 1220648-2025-25629 | 21174970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174970&pc=OZD | 03/12/2024 | 02/20/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  | 04/30/2024 |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218579 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Atrial Fibrillation | Hemolysis | Tachycardia | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiac arrhythmia with a possible relationship to the impella device used: 2/12 round 2130 rapid atrial fibrillation (af) and so started intravenous amiodarone. 3/13 unsuccessful direct current cardioversion 3/13 for paroxysmal af. later developed sustained ventricular tachycardia (1min) status post shock, sinus rhythm, rapid af, started on lidocaine bolus/drip. the patient recovered with sequelae. additionally, the loqi database reported upon a patient with hemoglobin drop with a possible relationship to the impella device: worsened anemia likely surgical. hemoglobin down to 6.8 from 8.5, received one unit of packed red blood cells. the patient survived the event. | [5 empty] | The investigation of hemolysis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25629. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12188 | 1220648-2025-25628 | 21174972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174972&pc=OZD | 01/03/2025 | 02/19/2025 |  | 01/16/2025 |  |  | 01/03/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585519 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella cp was exchanged due to the patient developing access site bleeding. impella support was for four hours. the patient survived the event. | [5 empty] | The investigation has been completed. no device was received for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. | E4 should have been left blank on manufacturer device report 1220648-2025-25628. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12189 | 1220648-2025-25630 | 21174978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174978&pc=OZD | 01/04/2025 | 02/13/2025 |  | 01/16/2025 |  |  | 01/04/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594145 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Perforation | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered following delivery of an impella 5.5 pump for mechanical circulatory support. upon removing the guidewire, the pump was inadvertently pulled back across the aortic valve. the device was re-advanced successfully; however, the staff was unable to properly position the pump towards the apex of the left ventricle. while attempting to reposition the device, the patient's blood pressure began to drop, and a chest exploration revealed a rupture of the left ventricular free wall. a mass transfusion protocol was initiated, and the patient was placed on cardiopulmonary bypass. the free wall rupture was repaired and the pump was explanted. the patient¿s outcome is unknown. | [5 empty] | The investigation for the positioning issue and the ventricle perforation has been completed. no product was returned and logs are not applicable. clinical details state the tissue of the patient's left ventricle was noted to be very necrotic at this point and the md said that the free wall rupture was "fileted/torn open." md did not believe the rupture to be directly caused by the impella but rather due to the condition of the patient's tissue. the root cause of the ventricle perforation was most likely patient-condition related. furthermore, after the pump was pulled into the aorta during guidewire removal, the md chose to attempt advancing the pump slowly back over the valve, which was successful, but then there were complications with correctly positioning inside the ventricle. the root cause of the positioning issue was most likely due to the wireless re-access attempted after the pump had been inadvertently pulled into the aorta. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12190 | 1220648-2025-25631 | 21174984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21174984&pc=OZD | 01/05/2025 | 03/11/2025 |  | 01/16/2025 |  |  | 01/05/2025 | 01/16/2025 |  |  |  |  | 01/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556094 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following implant, an impella cp pump triggered an impella in aorta alarm and purge system blocked alarm. positioning was verified to be correct via fluoroscopy and sonography and no notable kinks were discovered in the purge line. lysis protocol was discussed with the staff, but the decision was ultimately made to replace the pump. there was no noted patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue and high purge pressure has been completed. the data logs were not available. the returned product was investigated and there was remnants of biomaterial in the purge gap and around the impeller that remained even after soaking the pump. the pump was connected to a console and purged. initially, high purge pressure alarms reproduced, but once the pump was started, the pressure resolved. a window was cut in the catheter and there was no blood ingress. clinical details report that the tissue plasminogen activator (tpa) protocol was provided, though it's unclear if it was used. this could explain why the high purge pressure was able to resolve after restarting the pump. based on returned product and clinical details provided, the root cause of the high purge pressure was most likely biomaterial. the returned product was investigated and there were remnants of ingested biomaterial in the purge gap. there were no other visual abnormalities. the pump passed laser continuity, and when connected to the console, all 5s parameters were within specification. when the pump was connected to the console and purging, high purge pressure initially reproduced. while running the pump, all optical signal metrics were stable. clinical details report that the impella position in aorta alarm triggered at the same time as purge system blocked alarms. additionally, pump position was confirmed with fluoro. since the root cause of the high purge pressure was most likely biomaterial, the two complications occurred at the same time, and pump position was confirmed to be good, the positioning alarms were likely false triggers due to ingested biomaterial. based on returned product and clinical details provided, the root cause of the optical signal issue was determined to most likely be biomaterial. d.9 date device returned/information was added. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
12191 | 1220648-2025-25576 | 21157451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157451&pc=OZD | 12/20/2024 | 02/20/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471020 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient 38-week pregnant patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and while on supported the automated impella controller experienced a persistent battery failure alarm. the patient presented to the emergency department one day prior with palpitations and chest pain electrocardiogram showed st-elevation myocardial infarction versus myocarditis. catheterization was completed and ruled out blockages. the following day the patient was implanted with an impella 5.5 device now on venoarterial extracorporeal membrane oxygenation and impella mcs (ecpella). the patient was noted on continuous renal replacement therapy as well. the following day in the morning tissue plasminogen activator (tpa) was started due to elevated purge pressure above 1100 and purge flow around 2.4. with tpa, flows were noted to be improving. the patient continued on ecpella support with work up for orthotopic heart transplantation with hopes of recovery. now 12 days on support, the nurse reported automated impella controller (aic) had a battery failure. the battery was at 50%, it was plugged back in, but the alarm continued. it was indicated the aic was replaced. the patient was eventually decannulated and impella 5.5 device was successfully weaned and explanted in the cardiovascular operating room after a total of 14 days on support. the outcome of the patient following the event is unknown. however, there was no report of harm resulting from the aic failure. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the impella 5.5 device serial number (b)(6) with date of event 20-dec-2024 and identifier ending in (b)(6) . | The investigation for the console (aic) battery failure-red alarm has been completed. data logs were reviewed finding that during the support, the console displayed ¿battery failure (transport connection blown/missing) ¿ alarm,¿ event #360. this confirms the reported failure mode. after a few minutes, the console also displayed the ¿battery level low (50%) ¿ alarm,¿ event #23. the console was returned for evaluation. the reported failure mode of battery failure was reproduced. battery 2¿s lcd screen was blinking. the batttest utility via telnet confirmed the cell imbalance with cell 2 of battery 2. upon replacing the original batteries with known good ones, there was no battery failure alarm, and the batteries were able to charge. the root cause of the battery failure alarm was a single cell failure.
12192 | 1220648-2025-25577 | 21157531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157531&pc=OZD | 12/29/2024 | 02/13/2025 |  | 01/15/2025 |  |  | 12/29/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565481 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a gastrointestinal (gi) bleed during impella cp support. heparin was in the purge line leading up to the condition. the patient was given three units of packed red blood cells to counteract the bleed and resulting low hemoglobin levels. three days later, two units of packed red blood cells and two units of platelets were given to the patient. care was withdrawn the following day upon the family's request and the patient passed away. it is unknown if the gi bleed contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the gastrointestinal bleed (gi bleed) has been completed. data logs are not relevant to the issue and product was not returned for investigation. clinical details do not report any excessive force or patient conditions that would explain the cause for a gi bleed on the 11th day of support. since clinical details were limited and product wasn't returned, the root cause of the gi bleed could not be determined. the patient was on cp support for 11 days when the complication occurred. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12193 | 1220648-2025-25578 | 21157655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157655&pc=OZD | 12/30/2024 | 02/14/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. the physician was unable to pass impella through bifurcation. the procedure was aborted. additionally, upon removing from packaging, it was noted that the 0.018¿ wire was bent. the decision was made to use platinum plus wire. there was no harm to the patient. | [5 empty] | The investigation for the delivery issue and the product damage has been completed. product was not returned for investigation. clinical information mentioned that the pump was unable to pass through the bifurcation of the artery and pump implant was aborted due to the difficult vasculature. therefore, the root cause of the delivery issue was patient condition related to patient anatomy. without the pictures the failure cannot be further investigated. therefore, the root cause of the product damage (guidewire damage - peripheral vessels) issue was not determined since product/images were not returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12194 | 1220648-2025-25579 | 21157671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157671&pc=OZD | 12/31/2024 | 02/21/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584761 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately three days after start of the support, a fluid leak was discovered at the infusion filter of an implanted impella cp pump. no alarms were noted on the automated impella controller, and the functionality of the pump remained unaffected. the staff was advised to replace the pump due to the leak, but the physician declined due to the patient's instability. the pump remained in use despite the noted leak. there was no report or indication of harm to the patient as a result of the event. eventually, the patient was successfully weaned and the impella cp device was explanted. | [5 empty] | The investigation of purge leak ¿ pump has been completed. there was no product or data logs returned. without the product, the root cause of the purge leak - pump could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
12195 | 1220648-2025-25580 | 21157817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157817&pc=OZD | 12/30/2024 | 02/28/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519805 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 had no issue with the insertion but a few minutes after starting the impella, blood and purge solution ran out of the red impella plug. it was advice to the medical doctor (md) to change the impella. the md chose to keep the impella and add extracorporeal membrane oxygenation support. the next day no leak was observed around the impella plug. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the failure mode "product damage (hole in catheter) was added to address the blood leaking from the red handle. the product was returned for investigation. there was blood seen in the red handle and in the catheter. the system was purged overnight but data logs could not be obtained because the impella was defective due to the blood in the handle. the pump did not have any purge leaks and was able to keep a steady purge pressure of ~550mmhg. the root cause of the purge leak could not be reproduced. the clinical stated there was bleeding from the red handle. the catheter was looked at and a hole was found near the 35cm mark. this hole most likely occurred during insertion, since the bleeding was noted shortly after starting the pump. the root cause of the product damage (hole in catheter) is due to use.
12196 | 1220648-2025-25581 | 21157830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21157830&pc=OZD | 12/31/2024 | 02/12/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful as the physician was unable to pass the lesion into the distal aorta. the sheath was exchanged. however, delivery was still unsuccessful. the impella cp catheter was removed, and due to extensive amount of time spent on access and lack of progression, the team made the decision to abort the impella implant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass bifurcation on proximal section of the iliac.
12197 | 1220648-2025-25583 | 21158101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158101&pc=OZD | 01/01/2025 | 02/13/2025 |  | 01/15/2025 |  |  | 01/01/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580457 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient woke up in a confused state and pulled impella back into aorta prompting team take patient back to operating room for a new impella 5.5 device. additionally, at the end of the case, blood was noticed dripping from red impella plug. according to the provider, there may have been lots of torquing of the device. the patient survived the event. | [5 empty] | The investigation for the access site bleed and the positioning issue has been completed. the product was returned and evaluated. no logs are available for pump. there was blood leaking out of the red plug of the returned pump. a hole was then found in the catheter between the 30 and 35 cm markings, which is typically where suturing would take place around the catheter. the red plug was opened and the inside was covered in blood. however, the pump did pass electrical testing and was able to be recognized by the console without alarms, meaning that neither the motor cable lines or the communication cables had been damaged at this point. there were no reported issues with the pump's functionality in the field. clinical details state the pump that was pulled into the aorta by the patient was referring to a previous pump. it further states that access site bleed was actually product damage (hole in catheter) since the pump was found to have blood in the red plug and a hole in the catheter. the root cause of the hole in the catheter was most likely use-related as there was a puncture hole found in the catheter, likely around the suturing site (30-35 cm). added d9 device return date. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12198 | 1220648-2025-25584 | 21158193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158193&pc=OZD | 12/31/2024 | 02/17/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584651 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and approximately nine days after start of the support, the patient experienced an access site hematoma in the axillary region. consequently, heparin was withheld and the patient was given one unit of red blood cells. the patient was reported to be stable following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12199 | 1220648-2025-25585 | 21158422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158422&pc=OZD | 01/01/2025 | 02/05/2025 |  | 01/15/2025 |  |  | 01/01/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574171 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support and experienced lower extremity ischemia the following day. the patient was brought to the catheterization lab for an emergent heart catheterization and went into ventricular tachycardia/ventricular fibrillation arrest at inception of the procedure and was subsequently defibrillated 12 times. the decision made to place an impella cp in the setting of acute myocardial infarction-cardiogenic shock, which was successfully completed, auto flows within range followed by two drug-eluting stent implantations to the right coronary artery. the patient was transferred to the intensive care unit in stable condition. the following day, pedal pulses were undetectable. however, popliteal pulses were present. aortogram with run-off revealed right femoral artery flows that stopped at knee level, and bilateral peripheral artery disease. the plan was for the patient to receive two units of packed red blood cells, followed by fem-pop and possible removal of the impella in the operating room. the following day, the patient was taken to the operating room, and an embolectomy was performed at knee level. pedal pulses were undetectable still. the pump was weaned and explanted and pedal pulses were still undetectable via doppler. the patient was noted to have survived post explant of the device. however, further treatment to the patient for the ischemia is unknown at this time. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. d.5 revised as operator of device should not of been left blank on initial manufacturer device report number 1220648-2025-25585.
12200 | 1220648-2025-25586 | 21158617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158617&pc=OZD | 01/01/2025 | 02/12/2025 |  | 01/15/2025 |  |  | 01/01/2025 | 01/15/2025 |  |  | 06/30/2025 |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced limb ischemia followed by thrombus and ventricular tachycardia. the patient returned to the catheterization lab for restenting of stents placed approximately five days prior that restenosed. it was noted that the patient did not take the prescribed plavix medication. the patient was transferred back to the unit despite an ejection fraction of 15% and a left ventricular end-diastolic pressure of 35. on the unit, the patient went into pulmonary failure and was taken back to the catheterization lab to unload with an impella cp. following the impella implant, limb ischemia was experienced, and the patient underwent an external femorofemoral bypass surgery. on day three of impella support, the patient had a loss of the right pedal pulse. the patient was taken to the operation room and a thrombectomy aspiration performed on penumbra device of bilateral legs. flow past impella sheath was confirmed with contrast. impella was going to be explanted. however, the patient went into ventricular tachycardia, was shocked, and impella performance level was adjusted. a transesophageal echocardiogram was performed for placement, and the position was optimized. in the operating room, the patient had another run of ventricular tachycardia and was shocked again. the patient was cannulated for venoarterial extracorporeal membrane oxygenation while remaining on impella mcs. on day six of impella support, no pulse and discoloration to right lower extremity was noted and was unsuccessfully managed. there was no additional intervention. a clot in bilateral limbs and reperfusion failed. the outcome of the patient following the event is unknown. however, the last plan of care update, day seven of support, noted the patient still has an ejection fraction of 15% and has no option for permanent left ventricular assist device or transplant. the patient was successfully weaned and the impella cp device was explanted the following day. the patient was noted to have survived the explant. | [5 empty] | The investigation for the reported limb ischemia, ventricular tachycardia (vt), and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia, vt, thrombus could not be determined as the device was not returned nor sufficient clinical details. h.6 code 4641 was entered incorrectly on the initial manufacturer device report number 1220648-2025-25586. a new code was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12201 | 1220648-2025-25587 | 21158714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158714&pc=OZD | 01/03/2025 | 01/17/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591208 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia, hemolysis, and bleeding. the patient would subsequently expire. the patient presented to the emergency department after having ventricular fibrillation arrest at home. in the emergency department, the patient coded multiple times, was intubated, and transferred to the catheterization lab for an angiogram. right coronary artery disease was noted. cardiogenic shock with unstable hemodynamically. the decision was made to implant an impella cp device prior to percutaneous coronary intervention. the xext day following the implant, pulses in the right leg became weak and the right foot appeared mottled. the leg began to swell, and a fasciotomy was subsequently completed to restore pulse to the leg. four days later, pulse was once again absent in the patient's leg. the patient had been given ten units of blood as a result of the condition and intervention. coinciding with the ischemia, the patient also developed hemolysis with four units of bleed being given to help manage the hemolysis. the support level was lowered to reduce suction that may have been contributing to the hemolysis. as a result of the prior interventions, the patient experienced bleeding complications from the fasciotomy sites. multiple blood products were given to counteract the bleed and heparin levels were reduced. the patient later passed away after seven days of support. it is unknown if the complications contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B.2 life-threatening should had not been selected on the initial manufacturer device report number 1220648-2025-25587.
12202 | 1220648-2025-25588 | 21158908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158908&pc=OZD | 01/02/2025 | 02/05/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had an acute brain infarction. the patient was do-not-resuscitate and expired. | [5 empty] | The investigation of stroke/cva has been completed. the root cause of stroke cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12203 | 1220648-2025-25589 | 21158982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21158982&pc=OZD | 11/19/2024 | 02/12/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025532011 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a patient in need of care escalation from an impella cp. the patient was in acute myocardial infarction/cardiogenic shock. the patient suffered an access site bleed. the bleed was treated by evacuation and one unit of red blood cells. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided.
12204 | 1220648-2025-25590 | 21159250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159250&pc=OZD | 01/16/2024 | 04/01/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  | 07/31/2025 |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403300 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to impella support, the patient was cannulated with extracorporeal membrane oxygenation (ecmo) and was intubated. it was reported that after the impella cp with serial number (b)(6) (pump 1) was implanted and connected to the automated impella controller (aic) with serial number (b)(6) (aic 1), an alarm occurred for impella defective. the impella cp was explanted and replaced with a new impella cp pump with serial number (b)(6) (pump 2), and the impella defective alarm was not resolved. an additional aic with serial number (b)(6) (aic 2) was connected to the second pump, and the impella defective alarm was not resolved. the aic was replaced again with a third unit, the alarm was resolved, and support was initiated. after 6 days of ecmo and impella support, a cerebral bleed was discovered, and therapy was discontinued. it was noted that the cerebral bleed was not impella related. the patient was given systemic heparin during impella use. it was noted that the patient expired while on impella cp support. there was a delay in needed support when pump 1 would not be recognized by aic 1. additionally, there was a delay of support as the aic 1 and aic 2 would not recognize pump 1 or pump 2. it cannot be said the period of time with loss of hemodynamic support did not contribute to the patient's overall outcome. the patient experienced a cerebral vascular accident while on support with pump 2. while the physician stated it was not impella related, but there is no alternative cause identified or provided as the likely cause of such event. additionally, the patient was given systemic heparin for impella use which may have contributed to the cerebral bleeding. pump 1, aic 1, aic 2, and pump 2 cannot be ruled out as contributing factors in the overall event. | [5 empty] | This is one of four medical device reports associated with the event. refer to manufacturing report number 1220648-2024-06292 for automated impella controller 1 with serial number (b)(6). refer to manufacturing report number 1220648-2024-06294 for automated impella controller 2 with serial number (b)(6). refer to manufacturing report number 1220648-2024-06290 for pump 1 impella cp with serial number (b)(6). the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12205 | 1220648-2025-25591 | 21159319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159319&pc=OZD | 12/31/2024 | 03/07/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  | 12/31/2025 |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450498 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | White | Malfunction | No | Yes | Increase in Pressure | Failure to Pump | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following successful insertion, an impella cp device failed to start. during the initiation attempt, a pump blocked alarm occurred coinciding with a placement signal not reliable alarm. the device was removed as a result of the failure. there was no noted patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump did not start and optical signal issue has been completed. the failure mode (fm) "high purge pressure" was added for the voc for the "blocked alarms." there were no data logs returned to confirm if the pump was able to be started or was attempted to be ramped up to get it started. the pump was returned and there was a significant kink in the catheter near the motor housing. a laser continuity test was performed and light was seen from this location in the catheter indicating there was a broken fiber line in this location. the root cause of the optical signal issue is due to a damaged fiber line at the location of the catheter kink. in addition to that, there was a kink in the purge tubing. this kink was most likely leading to the "purge system blocked" alarms. without purge fluid expelling into the motor, that can lead to blood ingress. in the purge gap there was slight evidence of biomaterial. the root cause of the high purge pressure is most likely due to the kinked purge line seen near the motor. the pump was soaked and then the pump was ran in the lab. it initially would not start but after a few tries it was able to run without any pump stop alarms. however, the flow was saturated indicating there was ingress into the motor. the motor was torn down and at the bottom of the impeller, buildup could be seen. the impeller shaft had discoloration under the impeller with a slight red color as well as it was not shiny indicting there was material on it. the root cause of the pump did not start is most likely due to biomaterial based on the returned product. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25591 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12206 | 1220648-2025-25593 | 21159326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159326&pc=OZD | 01/02/2025 | 02/05/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564035 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following insertion of an impella 5.5 pump for mechanical circulatory support, some suction was encountered at support level p6. during troubleshooting, the pump was repositioned and subsequently triggered episodes of ventricular tachycardia. the support level was lowered to p4 and stability was achieved. support continued with the implanted pump following the adjustments to support. the patient survived the event. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12207 | 1220648-2025-25592 | 21159327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159327&pc=OZD | 12/30/2024 | 02/14/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  | 10/31/2025 |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423621 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced access bleeding requiring blood transfusions. the patient was taken to the catheterization lab and a 99% occlusion in the right coronary artery was observed. two stents were placed, and an intra-aortic balloon pump was unsuccessfully attempted. it was then decided to place the impella cp, which was successfully implanted. however, three days after the start of support, access site bleeding was observed, requiring two units of packed red blood cells. next steps were to stent the left anterior descending artery and look at the left circumflex artery with thereafter next steps to wean and explant the impella cp. overnight, the patient's urinary output decreased, and the patient continued to decompensate. no intervention was taken to culprit lesion, and the patient would expire the same day. family decided to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. b.2 revised as a thrid selection was incorrectly reported on the initial manufacturer device report number 1220648-2025-25592.e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25592 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12208 | 1220648-2025-25594 | 21159332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159332&pc=OZD | 12/30/2024 | 02/14/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560877 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support while waiting for heart transplant and approximately 61 days into support, a kink in the catheter at the proximal end by the red plug for identified. the nurse "untwisted" the catheter to mitigate risk of damage, the support continued with the device and no adverse effect to the patient. the patient was successfully transplanted 10 days later and over a period of 48 hours the patient was successfully weaned and impella cp device was explanted. | [5 empty] | The investigation for the product damage has been completed. product was not returned for review. the root cause of catheter kink was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12209 | 1220648-2025-25595 | 21159348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159348&pc=OZD | 12/30/2024 | 01/24/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541587 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | It was reported the field service took call from the user facility and stated during support to a patient (unknown demographics), the automated impella controller issued a purge cassette error message. consequently, the aic was exchanged as advised by field service. there was no report or indication of support interruption and/or adverse effects to the patient. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue - other has been completed. according to data and device analysis the reported purge cassette error issue was determined to be a purge disc not detected issue caused by a broken purge disc flag. e4 should have been left blank on manufacturer device report 1220648-2025-25595.
12210 | 1220648-2025-25596 | 21159364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159364&pc=OZD | 12/30/2024 | 02/18/2025 |  | 01/15/2025 |  |  | 12/30/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591487 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and shortly after the implant, there were high purge pressures with purge blocked alarms. staff tried troubleshooting, however the pump was explanted and replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of high purge pressure has been completed. the pump was returned for investigation. the data log showed a four-minute gradual rise in purge pressure, followed by an active high purge pressure alarm, and then pump flow saturation. no motor current spike or optical signal abnormalities were reported during the case. no evidence of physical damage was noted on the returned product. the pump was soaked in warm water for 15 minutes, and biomaterial was found inside the purge gap. the pump was primed, and high purge pressure was reproduced. the pump was able to start in a bucket of water, and the purge pressure began to improve. additionally, some biomaterial was observed in the purge gap after the pump was run in a bucket of water. the cause of the high purge pressure issue was biomaterial, based on returned product investigation. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25596 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12211 | 1220648-2025-25597 | 21159797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21159797&pc=OZD | 10/23/2024 | 01/29/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542835 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an automated impella controller (aic) - 5.5 support ongoing for a patient bridging to transplant or left ventricular assist device. the aic failed to display appropriate flows and so the team replaced the aic. no harm or injury and pump remains on support with the 2nd console. in the investigation of the console the data logs revealed new information. the console was not the malfunctioning product, rather the pump was at fault. | [5 empty] | The investigation for the low pump flow has been completed. no product was returned. there are fluctuations in motor current, placement signal and the pump flow without any changes in p-level. there is a small spike in motor current, followed by lower pump flows. at this time, suction alarms are occurring. flows are variable for 9 more hours and there are positioning alarms. the console is switched 9 hours later and the pump is running fine for the remainder of the case. the root cause of the low pump flow is most likely due to ingested biomaterial. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12212 | 1220648-2025-25598 | 21160098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21160098&pc=OZD | 01/02/2025 | 10/01/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585467 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the device exhibited an optical sensor issue. | Us complainant reported the complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. staff believed the vessel may have been nicked by a scalpel during initial access. two units of red blood cells were given to the patient. additionally, impella mean arterial pressure (map) appeared to be about 20 mmhg lower than arterial line map during support. support continued. the patient survived the event. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. no issues were found with the returned product, and the pump functioned per the specifications. a deteriorating trend in the placement signal was reported during the initial case run, but it recovered, showing a slow downward trend for another 18 hours. a positive left ventricle offset was applied during the initial trend of the placement signal. no placement signal trend was noted correlating with the placement signal trend. the cause of the optical signal issue was most likely patient condition related, based on returned product investigation and data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | Additional information provided in h3, h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post-sterile inspection checks. the pump was returned for investigation. no issues were found with the returned product, and the pump functioned per the specifications. clinical details suggest a discrepancy between the a-line and impella map. there is also concern over vascular damage could be from the scalpel used during the procedure, which caused access site bleeding that resolved after applying manual pressure and administering 2 units of blood. the cause of the vascular damage issue was not determined due to insufficient clinical information. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12213 | 1220648-2025-25599 | 21160374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21160374&pc=OZD | 01/03/2025 | 02/18/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a hematoma and experienced some oozing from their access site following impella cp implant. manual pressure was held and forward tension sutures were placed to manage the condition. support was continued with the implanted pump following the intervention. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. according to the clinical details the root cause of the access site bleeding is determined to be patient condition because of diseased left and right main coronary artery disease, left anterior descending, and left circumflex artery. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12214 | 1220648-2025-25600 | 21160705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21160705&pc=OZD | 01/03/2025 | 02/19/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182674 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has a power issue and a system self-check failed. the aic was removed from service. no patient harm has been reported. | [5 empty] | E.1 the initial reporter name is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - a/c power issues and aic - system self-check error has been completed. the console logs are consistent with complaint and show communication failures, causing console crashes and reboots. eventually the console was unable to boot-up. inspection of the console electronics - in particular power battery manager (pbm) board - did not reveal any corrosion or mechanical damage. attempts to retrieve firmware from directly from pbm1 microcontroller were unsuccessful and the microcontroller was non-responsive (pbm2 firmware and microcontroller were fine). the cause of the aic issue was a defective pbm board.
12215 | 1220648-2025-25601 | 21160729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21160729&pc=OZD | 01/03/2025 | 02/19/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576549 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the impella alarmed pump stop and motor dropped to 0 with 0 flow. immediately following impella was able to ramp itself up but unable to flow above p4 -2.2l without suction. the position was checked with echocardiography and repositioned; at this time, thrombus is suspected. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the root cause of low pump flow issue cannot be determined since the complaint device not available for analyzing and insufficient clinical data provided. data logs reviewed: no mc spike observed out side of p-level changes. suctions occurred. a suction alarm occurred followed by drop in flow to zero then ramp it up and consist on p4 -2.2l without suction. no positioning alarms observed. failure mode temp pump stop added to the investigation since the impella alarmed pump stop and motor dropped to 0 with 0 flow. the root cause of temp pump stop issue cannot be determined since the complaint device not available for analyzing and insufficient clinical data provided.
12216 | 1220648-2025-25602 | 21161081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21161081&pc=OZD | 12/31/2024 | 02/13/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587406 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 placement was aborted due to the patient¿s vascular anatomy. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12217 | 1220648-2025-25603 | 21164028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164028&pc=OZD | 12/31/2024 | 02/12/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580356 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. at end of case and repo sheath had been advanced, site looked good but no doppler pulse was present and post fem shot showed no flow distally. the patient had weaned on levo and was off epi so decision was made to remove impella due to perfusion concerns. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25603. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12218 | 1220648-2025-25604 | 21164081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164081&pc=OZD | 12/31/2024 | 02/14/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025585984 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. it was documented that multiple attempts were made, but ultimately the patient's severe stenosis prevented a successful delivery. upon catheter removal, multiple kinks were visible on the cannula and the decision was made to replace the pump. the second device was successfully implanted after the axillary artery was ballooned and stented. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, 5.5 insertion proved very difficult due to abnormal patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12219 | 1220648-2025-25605 | 21164082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164082&pc=OZD | 12/31/2024 | 02/12/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574134 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and it was aborted due to patient anatomy. patient was noted to be stable and no further details were provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. e4 should have been left blank on manufacturer device report 1220648-2025-25605.
12220 | 1220648-2025-25606 | 21164086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164086&pc=OZD | 12/31/2024 | 02/13/2025 |  | 01/15/2025 |  |  | 12/31/2024 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a large hematoma at the impella access site on day one of impella cp support. heparin was stopped and hematoma resolved. on day three of cp support hematoma noted. heparin stopped, manual pressure held, one unit of packed red blood cells was administered, and one platelet was given. the patient was taken to the cath lab and after left anterior descending artery intervention the impella was removed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25606 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12221 | 1220648-2025-25607 | 21164090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164090&pc=OZD | 01/02/2025 | 02/13/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574102 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was being prepped for an impella implant and upon removing the guidewire from the packaging, the wire bent and kinked in two spots. a new kit was opened and a different wire was used successfully. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the product was returned, and the guidewire was bent in two places. it was bent near the distal tip and the other end. due to limited clinical information provided, the root cause could not be determined. d.9 date device returned/information was added. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
12222 | 1220648-2025-25608 | 21164171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164171&pc=OZD | 01/02/2025 | 02/12/2025 |  | 01/15/2025 |  |  | 01/02/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571166 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was removed after five hours of support due to distal limb ischemia. the patient is stable and survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25608 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12223 | 1220648-2025-25609 | 21164192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21164192&pc=OZD | 01/03/2025 | 02/10/2025 |  | 01/15/2025 |  |  | 01/03/2025 | 01/15/2025 |  |  |  |  | 01/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569264 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient was brought to computed tomography scan for stroke-like symptoms and was found to have had a stroke. the patient survived the event. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12224 | 1220648-2025-25539 | 21147531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147531&pc=PYX | 12/29/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  | 08/31/2025 |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395773 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an impella rp implant procedure to a patient who was post-operative cardiothoracic surgery, the tip of the 23 french x 30 centimeter dilator was noted to be torn/damaged and was unable to use for advancement of the peel away sheath into the body. the decision was made to use a dilator and sheath from the impella rp flex. the outcome of the procedure was noted as aborted. however, it is unknown what actions were taken to treat the patient as well as the patient's outcome following the event. there was no report or indication of any adverse effects to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of introducer damage ¿ mechanical has been completed. the cause of the introducer damage was not determined since product/images were not returned and insufficient clinical details were provided.
12225 | 1220648-2025-25540 | 21147609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147609&pc=PYX | 12/29/2024 | 02/07/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  | 03/31/2025 |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024390088 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to delivery an impella rp flex to support right heart failure, and the case was aborted. subsequent actions to treat the patient (unknown patient-specific information) was not provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
12226 | 1220648-2025-25541 | 21147727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147727&pc=OZD | 12/27/2024 | 02/05/2025 |  | 01/14/2025 |  |  | 12/27/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541807 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient¿s left leg was noted to be cold and mottled. the impella cp was residing in the left common femoral artery. an ultrasound of the left leg showed no arterial flow distal to the impella access site. the decision was then made to remove the impella. arterial flows and pulses were restored after removal. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12227 | 1220648-2025-25542 | 21147752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147752&pc=OZD | 12/26/2024 | 02/05/2025 |  | 01/14/2025 |  |  | 12/26/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction presented to the catheterization lab for impella cp implant and upon insertion the patient went into sustained ventricular fibrillation with a demised outcome. left heart catheterization was completed and showed clot from left main (lm) through the entire left anterior descending artery (lad) with a heavily calcified right coronary artery and pulmonary arterial end-diastolic pressure of 71. for the impella cp implant, the groin sized up to 14fx13cm peel away. pigtail and j wire were used to gain left ventricle before abiomed .018 wire placed and impella inserted with flows ranging between .5-1.5 in auto. upon insertion, the patient went into sustained ventricular fibrillation, was shocked and received cardiopulmonary resuscitation. at this point, the decision was made to place an impella rp. after the impella rp had been opened and prepped, the physician opted to place the patient on veno-arterial extracorporeal membrane oxygenation (va-ecmo) due to cardiac standstill and oxygenation requirements. minimal flows were able to be achieved after placement and at this point the patient had exhibited signs of pulmonary hemorrhage. single access was performed after ecmo placement where physician had attempted to use the penumbra and series of balloons on lm-lad but was unsuccessful. ultimately, the patient expired on support in catheterization lab; the team determined there was no other intervention that could be delivered to the patient. | [5 empty] | The investigation for the reported ventricular fibrillation (vf) has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the vf was not determined. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25542. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25542 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 several health effect - impact codes were inadvertently omitted from the initial manufacturer device report number 1220648-2025-25542 and were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12228 | 1220648-2025-25543 | 21147757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147757&pc=OZD | 12/29/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584940 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient admitted with a st-elevation myocardial infarction with was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and soon thereafter implant, the patient experienced ventricular tachycardia. the impella was noted to be in good position pointed towards the apex, but patient would go into ventricular tachycardia. ultimately the decision was made by physician to have the impella pointed towards mitral valve at 4.2cm as this orientation did not cause ventricular tachycardia. while repositioning the impella 5.5, the patient went into ventricular tachycardia requiring multiple cardioversions. the patient continued on impella 5.5 mcs for an additional 10 days before being successfully weaned and reportedly implanted with a durable left ventricular assist device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined.
12229 | 1220648-2025-25544 | 21147761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147761&pc=OZD | 12/29/2024 | 02/28/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592560 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and thrombus in the femoral artery occurred with the use of the impella cp. the patient had a very high bleeding, and the decision was made to use a femostop to resolve the bleeding. the thrombus was first seen, after the surgical opening of the femoral artery and a hematoma was the result of the use of the femostop. there was no treatment of the hematoma. following, approximately ten hours after start of support, the impella cp device was explanted due to malposition and the patient's need for higher support. the patient was reported to be stable following the event. | [5 empty] | The investigation for the access site bleed, positioning, and thrombus has been completed. product was not returned for review. the root cause of access site bleeding and thrombus was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12230 | 1220648-2025-25545 | 21147767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147767&pc=OZD | 12/30/2024 | 02/25/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583673 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was initially implanted with an impella cp device and was escalated to an impella 5.5 device for mechanical circulatory support. approximately four days after start of support the patient experienced retroperitoneal bleeding. the source of the bleed could not be ascertained. information received from the attending physician indicated that there was a possibility of aortic wall damage when the impella cp was removed, but the computed tomography scan could not determine either a location of dissection or bleeding. the causal relationship, however, between the bleeding and the impella 5.5 is unknown. the patient would subsequently expire. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. the cause of the vascular damage - peripheral vessel was not determined since the product was not returned and there was insufficient clinical information. b.2 revised as some selections were reported incorrectly on the initial manuacturer device report number 1220648-2025-25545. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25545 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4614 was reported incorrectly on the initial manuacturer device report number 1220648-2025-25545.
12231 | 1220648-2025-25546 | 21147783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147783&pc=OZD | 12/30/2024 | 02/13/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538088 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported post left heart catheterization (lhc) three days prior, the patient returned to the hospital catheterization lab for the second lhc with atherectomy and stent to distal left main and proximal left anterior descending arteries with impella assistance. however. during the protected percutaneous coronary intervention, a dissection occurred. the physician obtained right femoral artery access, dilated up, and then placed the peel-away sheath. as the physician was working to insert the impella cp, the femoral artery was showed on fluoroscopy and it realized the femoral artery was dissecting. the physician removed all devices, applied manual pressure, utilized perclose (already placed), and then placed a femostop to address the dissection. consequently, the case was aborted due to anatomy complications and femoral artery dissection. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. as per the clinical information provided, the dissection was observed after implant during flouroscopy which was due to the patients anatomy. therefore, the root cause of the vascular damage issue was patient condition related to patient anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b.1 and b.2 revised as selections were inadvertently omitted from initial manufacturer device report number 1220648-2025-25546. d.4 revised catalog number as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25546. d.6a and d.6b revised as dates were inadvertently omitted from initial manufacturer device report number 1220648-2025-25546. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25546as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4627 was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25546. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2025-25546 and were added.
12232 | 1220648-2025-25547 | 21147790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147790&pc=OZD | 12/30/2024 | 02/13/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546792 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support and the following day after start of support, there was a call from the bedside regarding position in the aorta and high purge pressure alarm. the position was confirmed under echocardiogram. the impella connect suggested clot ingestion. there was a motor current spike and brief saturated flow that resolved. the device was eventually explanted and replaced approximately two days later for clot ingestion. the outcome of the patient following the event is unknown. however, there is no report or indication of any adverse effects to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of saturated flow and high purge pressure has been completed. there was no product or data logs returned. the clinical description stated that there was an ingestion signature with high purge pressure and saturated flow but then the saturated flow resolved. tissue plasminogen activator was started but the pump was ultimately exchanged due to clot ingestion. however, due to lack of product or data logs returned, the root cause of the high purge pressure could not be confirmed. the root cause of the saturated flow could not be determined.
12233 | 1220648-2025-25548 | 21147837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147837&pc=OZD | 12/30/2024 | 02/04/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320592 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller had a controller failure alarm during use supporting a patient in the intensive care unit. it was noted the device remained in use and there was no harm to the patient. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) controller failure has been completed. before the complaint date, the first instance of pbm communication occurred at boot up. this happened intermittently for subsequent boot ups and cases. the last imc log available and there were not logs available for the complaint date. the console was returned for investigation and the failure mode reproduced. the pbm1 communication trace was found to be corroded on the pbm. additionally, there were pbm communication issues in the imc logs as a result. the cause of the pbm communication failure was a contaminated communication pcb trace.
12234 | 1220648-2025-25549 | 21147922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147922&pc=PYX | 12/30/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  | 07/31/2025 |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024385522 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock was initially implanted with an impella cp device, escalated to an impella 5.5 and placement of an impella rp device for bipella support. the impella rp was placed through the right femoral vein and unable to advance. the impella rp kinked onto self. the decision made to abort the impella rp device implant and attempt with an impella rp flex which was successfully implanted. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
12235 | 1220648-2025-25550 | 21147948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21147948&pc=OZD | 12/30/2024 | 09/22/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548162 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was brought to the catheterization lab for biventricular implantable cardioverter defibrillator placement and went into ventricular tachycardia storm and was shocked 60 times. the decision was made to place an impella cp device, which was successfully completed. however, following, the implant, low flows were encountered. flows were initiated on auto at 3.6 liters per minute. the physician made the decision to keep the peel away in place even after discussing the risks associated with doing so. the patient was transported to the unit with a plan to transfer to a higher level of care. after arriving to the unit, the impella started giving different alarms. it was noted the patient was not moving. the pump would read impella in ventricle, then impella in aorta, then back in ventricle. thereafter, the aorta signal completely went out for approximately ten seconds and the left ventricle systolic waveform read 3000. then the waveform would come back and match up with the arterial line. following there were suboptimal flows as well. the staff was uncertain if there was a clot ingested, if the optical sensor was inaccurate or if the patient's heart was "really bad shape." the pump was eventually explanted after a total of six hours on support. the patient was reported to be in critical condition following the event. the patient was transferred to an outside hospital for the higher level of care. | [5 empty] | The investigation of low pump flow has been completed. upon visual inspection, no biomaterial was present on returned pump. no cannula kinks or impeller damage was observed. pump was run at p-9 for approximately 7 minutes and low flow did not reproduce. data logs were reviewed and lt logs show that when pump was running at p-6, flows were lower than expected range. multiple suction alarms occurring throughout support. rt log at beginning of case show variable motor current and pump flow with low pulsatility. rt log at time of pump running at p-6 shows suction alarms and a small step up in motor current with a speed deviation and step up in pump flow. clinical details noted that pump was having suboptimal flows throughout support. the root cause of the low pump flow was most likely due to biomaterial. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. upon visual inspection, no abnormalities were noted. optical parameters of returned pump were checked, and no hard failures of the optical sensor were observed. pump passed laser light continuity testing. no biomaterial was present on returned pump. pump was connected to console and no placement signal not reliable (psnr) alarm was reproduced. data logs were reviewed and lt log from field shows three instances of ¿impella in aorta¿ and ¿impella in ventricle¿ occurring at the same time. the first two instances of positioning alarms, ¿impella position unknown¿ alarms also occurred. no hard failures of the optical parameters were observed during case run. at times of ¿impella in aorta¿ and ¿impella in ventricle¿ alarms, optical parameters were affected and became more variable. rt log at time of first instance of ¿impella in aorta¿ and ¿impella in ventricle¿ alarms, all optical parameters were affected. additionally, motor current was variable at time of alarms. rt log at second instance of ¿impella in aorta¿ and ¿impella in ventricle¿ alarms shows a small step up in motor current and pump flow. additionally, a small motor current spike is observed seven minutes after step up in motor current with a small spike in pup flow and small speed deviation. both instances had placement signal become more pulsatile after motor current was affected. clinical details noted that pump was on p-auto and multiple alarms were occurring simultaneously. the root cause of the optical signal issue was most likely due to biomaterial. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-25550. d10 concomitant medical products and therapy dates updated. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12236 | 1220648-2025-25551 | 21148209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148209&pc=OZD | 12/27/2024 | 02/07/2025 |  | 01/14/2025 |  |  | 12/27/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551314 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the decision was made to take patient to cath lab to explant impella out of concern for hemolysis and place intra-aortic balloon pump with continued extracorporeal membrane oxygenation support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post sterility inspection checks. data log review confirmed occurrence of the reported failure mode and clinical narrative. logs showed pump started on p-9 but then ran on p-1 through the rest of the case. additionally, there were noted suction and positioning alarms triggered during the case but resolved. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12237 | 1220648-2025-25552 | 21148376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148376&pc=OZD | 12/16/2024 | 02/10/2025 |  | 01/14/2025 |  |  | 12/27/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured anemia with a "possible" relationship to the impella device used: ¿hgb 9.3 gm/dl on admission. post impella implant (b)(6) 2024- hgb 4.4. no clear evidence of bleeding. transfused 3 units prbcs and 2 units plasma. the patient expired.¿ the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12238 | 1220648-2025-25553 | 21148804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148804&pc=OZD | 12/29/2024 | 02/04/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584546 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had a large hematoma at the impella access site. heparin was paused and possible a wash out would be needed. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12239 | 1220648-2025-25554 | 21148915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148915&pc=OZD | 12/28/2024 | 02/28/2025 |  | 01/14/2025 |  |  | 12/28/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587351 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support when a kink was visualized on the catheter shaft and an orange-red light was seen. the position was confirmed and there was no reported patient harm. | [5 empty] | The investigation for the product damage has been completed. the product was returned and evaluated. there was a slight kink in the catheter found just past the edge of the kink protector. an optical fiber break at that location was confirmed with the laser continuity test, which confirmed the reported issue in the field. the 5s parameters also indicated a fiber break with signal-to-noise ratio at nearly zero. data logs are not available, but it is known that the issue did not present until the second day of support. the root cause of the optical signal issue was a damaged fiber line as a kink and optical fiber break were found in the catheter just past the kink protector. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12240 | 1220648-2025-25555 | 21148973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148973&pc=OZD | 12/30/2024 | 02/25/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519895 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the attempted implant of an impella 5.5 device for mechanical circulatory support (mcs) to a patient (demographics unknown) with cardiomyopathy was unsuccessful due to the patient's anatomy. the implant procedure was aborted. further details including the patient's status are unknown. however, the was no report or any indication of any adverse effects to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy.
12241 | 1220648-2025-25556 | 21148995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21148995&pc=OZD | 12/29/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562127 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella 5.5. after cut down and the graft was placed, the physician accessed the left ventricle with ease but was unable to advance the impella 5.5 past the anastomosis as the artery was very small. after multiple attempts, change in position, switching to a v18, the decision was made to abort the impella 5.5 insertion and opted for an impella cp. the patient¿s outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, unable to advance 5.5 past the anastomosis as artery was very small.
12242 | 1220648-2025-25557 | 21149121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149121&pc=OZD | 12/18/2024 | 02/05/2025 |  | 01/14/2025 |  |  | 12/18/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582340 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 and during support, the patient had bleeding at the access site. a wound vac was placed at the access site. support was continued and the staff monitored. additionally, the patient had ectopy when in certain positions. there were no impella alarms. the staff ordered echocardiography to check positioning. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding and ectopy has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the ectopy was not determined.
12243 | 1220648-2025-25558 | 21149387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149387&pc=OZD | 12/28/2024 | 02/27/2025 |  | 01/14/2025 |  |  | 12/28/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581930 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that high purge pressure developed during impella cp support. possible causes were reviewed and lysis protocol was sent to the staff to manage the issue. support continued uninterrupted. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the pump was not returned for investigation. data logs indicate that the doctor took more than 4 minutes to complete the purge bag/cassette change procedure. additionally, within two minutes, a 'purge cassette failure' alarm appeared. the cassette change procedure was repeated, and a 'purge system blocked' alarm was seen after the completion of the procedure. it was evident that the doctor tried to check the purge line connection by disconnecting and reconnecting it. this is where the 'purge line open' alarm was seen at the beginning of the high purge pressure. the cause of the high purge pressure issue was most likely use related to longer purge cassette/bag change procedure. as per the instructions for use, the recommended time for the cassette bag change procedure should be less than 2 minutes. g1 manufacturing site name and address revised to reflect proper manufacturing facility on manufacturer device report 1220648-2025-25558.
12244 | 1220648-2025-25559 | 21149452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149452&pc=OZD | 12/30/2024 | 02/27/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576845 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the purge pressure suddenly increased while the purge flow decreased. a high purge pressure and purge flow blocked alarm message appeared on the automated impella controller in response to the sudden change in support. resolution options were discussed with the site and the decision was made to initiate a tissue plasminogen activator protocol. support was continued following the intervention. upon follow up, it was documented that the patient passed away the following day due to their overall condition. it is unknown if the impella related malfunction contributed to the patient's passing. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella 5.5 was returned for evaluation. the pump was returned. cut at catheter and only pump cannula and motor housing side was returned. upon visual inspection, biomaterial was present in the purge gap after pump was soaked in warm water to remove dried blood. data logs were reviewed and long term log showed an increase in purge pressure over a period of approximately four hours with a simultaneous decrease in purge flow. purge pressure remained high for the remainder of the case. clinical details noted that "high purge pressure" alarm occurred on day five of support, the pump had dextrose 5 percent in water with sodium bicarbonate running in the purge solution. tissue plasminogen activator protocol was initiated in attempt to resolve alarms. the root cause of the high purge pressure was due to biomaterial. b.1 revised as product problem was inadvertently omitted on the manufacturer device report number 1220648-2025-25559.
12245 | 1220648-2025-25560 | 21149644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149644&pc=OZD | 12/30/2024 | 02/13/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525058 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [5 empty] | CONTINUOUS RENAL REPLACEMENT THERAPY | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis after one day of impella cp support. the patient's labs were hemolyzed and evidence of renal failure was present. continuous renal replacement therapy was initiated to counteract the condition and the decision was made to wean and explant the pump. the patient survived the event. | [5 empty] | The investigation for the hemolysis and renal failure has been completed. data logs were reviewed, there were no motor current spikes outside of p-level changes, purge trends or positioning issues noted. pump flows were within specification throughout the case, but they were unstable/variable at times and 16 suction alarms triggered on 1/1. additionally, there was a sustained drop of the placement signal between 12/30 and 12/31 that did not improve again until 1/3. returned product was investigated and only part of the catheter and the cannula were returned, so the pump wasn't able to be run in the lab to attempt to reproduce the issue. however, there appeared to be no damage to the pump that would suggest it was the cause. there was some biomaterial residue in the purge gap, but as mentioned above, there were no signs of biomaterial ingestion in the data logs that would suggest this caused the hemolysis reported from the first day of support. clinical details report that the patient's cvp dropped between the first two days of support along with the patient's labs becoming hemolyzed and urine output turning red. by 1/3, the complication had resolved, cvp had improved, and urine output was normal. this aligns with the trend seen in the ps on data logs. based on all the information available, the root cause of the hemolysis was determined to most likely be due to patient condition. the root cause of the renal failure was not determined due to insufficient clinical details. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12246 | 1220648-2025-25561 | 21149717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149717&pc=OZD | 12/28/2024 | 02/10/2025 |  | 01/14/2025 |  |  | 12/28/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, a crack was found in purge sidearm after transporting the patient. decision was made to replace the impella. there was no reported patient harm. | [5 empty] | The investigation for the product damage has been completed. the pump was not returned for investigation, and the data log was not returned either. according to the clinical report, side arm damage was found damage during transportation, accompanied by a low purge pressure alarm. the clinical team indicated signs of a purge leak on the pump, as the low purge pressure alarm had been triggered. so the product damage failure mode was updated to purge leak-pump failure mode. the cause of the purge leak issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12247 | 1220648-2025-25562 | 21149724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149724&pc=OZD | 12/31/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 12/31/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194665 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, a controller error alarm was occurring on the automated impella controller (aic). the placement signal and left ventricle waveform disappeared. the impella 5.5 continued to provide support with good flows and normal motor current. no event started the alarm and it self-resolved after about two minutes. a new controller was brought to the room. however, the aic was not switched but just kept next to the active aic. no patient harm has been reported. | [5 empty] | The investigation for the automated impella controller (aic) controller failure-red alarm has been completed. the console logs from the reported day of event are consistent with the complaint and confirmed the reported failure mode given there were multiple controller error alarms ¿opm comm crc error¿ triggered during the clinical procedure. real time logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console logs also indicated frequent interruptions of universal serial bus communication from aic to rlm. the console was tested but the placement signal issue was not reproduced. during console inspection, the rlm was found to be malfunctioning (frequent interruptions of data streaming) due to compromised power and data connection caused by loose input/output panel internal connector (unrelated to the complaint). the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12248 | 1220648-2025-25563 | 21149820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21149820&pc=OZD | 01/01/2025 | 02/13/2025 |  | 01/14/2025 |  |  | 01/01/2025 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591472 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Vascular Dissection | Insufficient Information | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after removing the impella cp, the percloses were not able to be successfully deployed. the vascular surgeon was called for repair. the patient is stable. the patient survived the event. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the data logs are not relevant to the investigation and the product was not returned. clinical details allege that once both the pump and sheath had been successfully removed, the physician had difficulty deploying the perclose, so vascular was called in for repair. with effort, they were able to obtain hemostasis with new percloses. based on the clinical details provided, the root cause of the vascular damage was determined to most likely be a use issue. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b.1 revised as product problem was inadvertently selected on initial manufacturer device report number 1220648-2025-25563. b.2 revised as selection was inadvertently omitted from initial manufacturer device report number 1220648-2025-25563 h.6 health effect - clinical code 4580 was reported incorrectly on manufacturer device report 1220648-2025-25563. a new code was added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report 1220648-2025-25563. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
12249 | 1220648-2025-25564 | 21150038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21150038&pc=OZD | 12/30/2024 | 02/07/2025 |  | 01/14/2025 |  |  | 12/30/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient underwent an emergent catheterization on day prior with no percutaneous coronary intervention. the team was awaiting surgical consult and during the wait time the patient experienced ventricular tachycardia. the patient was brought back to the catheterization lab for an impella cp implant until plan of care was determined. the next day after start of support, hemolysis was noted. however, there was not treatment. the following day, the patient was successfully weaned and the impella cp device was explanted. the patient was reported to be in stable condition following the event. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. pump passed all post sterile inspection checks. no device was received for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12250 | 1220648-2025-25565 | 21150315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21150315&pc=OZD | 12/28/2024 | 02/07/2025 |  | 01/14/2025 |  |  | 12/28/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591310 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. due to the patient's vasculature, the staff was unable to advance the pump into position. as a result of the noted resistance, the decision was made to abort the implant. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
12251 | 1220648-2025-25566 | 21150601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21150601&pc=OZD | 12/31/2024 | 02/19/2025 |  | 01/14/2025 |  |  | 12/31/2024 | 01/14/2025 |  |  | 02/28/2026 |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484434 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was implanted to support a patient with acute myocardial infarction / cardiogenic shock. during the implant, the cp had purge pressure low and motor current high alarms. the impella stopped and then restarted. while the impella was being placed the patient coded and shortly after placement the patient expired. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and high purge pressure has been completed. the impella device was not returned for evaluation. the failure mode "high purge pressure" is changed to "purge leak - pump" as the clinical details stated there were low purge pressure alarms. the data logs show that the case lasted for 20 minutes. the pump ran for 17 minutes and then experienced a high motor current pump stop. low purge pressure was seen at the beginning of the case but resolved after 1 minute and remained stable. no other purge issues are noted. there were positioning/suction alarms seen and the motor current/placement signal are not pulsatile before the pump stops. the motor current drops from 400ma to 200ma and pump and the pump was turned off, and then 8 seconds later there is a high motor current pump stop as it was attempted to be turned on. the pump is restarted and then there are suction alarms after. it runs for 28 more seconds and then the pump is removed. the clinical details state that the patient was coding upon impella placement. the root cause of the temporary pump stop could not be determined. the root cause of the purge leak - pump could not be determined since product was not returned.
12252 | 1220648-2025-25567 | 21151546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21151546&pc=OZD | 12/31/2024 | 03/03/2025 |  | 01/14/2025 |  |  | 12/31/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569084 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted but the cp flow was not able to go above 1l in auto. a kink of impella cannula was also visualized. the impella was removed, and a new one was placed. the patient survived the event. | [5 empty] | The investigation of low pump flow has been completed. the data logs confirmed low pump flow when pump started and remained to the rest of the case. visual inspection found a kink on the cannula about 15 millimeters from the of cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was a cannula kink due to patient condition. d.9 date device returned/information was added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12253 | 1220648-2025-25568 | 21151803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21151803&pc=OZD | 12/29/2024 | 02/14/2025 |  | 01/14/2025 |  |  | 12/29/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577745 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed oozing at the impella access site. sutures were added. it was also reported that labs were hemolyzed. one unit of prbc was given after hemoglobin dropped. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the access site bleed and the hemolysis has been completed. returned complaint device was visually inspected. canula kink was found near outlet observed. no sharp edge or broken blade found. no biomaterial observed. data log reviewed finding no motor current (mc) spike observed outside of p-level changes. no positioning issue observed. suction alarms occurred. clinical stated that additional sutured added to stop bleeding. hemolysis issue reported, echo with pump measured deep and pulled back 2cm, no hemolysis reported and patient reported to be stable. the root cause of access site bleeding most likely was use related since extra sutures needed to achieve hemostasis. the root cause of hemolysis issue most likely was improper positioning related since echo with pump measured deep and pulled back 2cm, and on (b)(6) 2024 no concern or issue stated and patient reported to be stable. cannula kink may occurred during removal. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12254 | 1220648-2025-25569 | 21152040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21152040&pc=OZD | 12/31/2024 | 02/19/2025 |  | 01/14/2025 |  |  | 12/31/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582271 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during the implant, the repositioning sheath kinked and it was causing difficulty repositioning the pump. also, during insertion, the patient went into ventricular fibrillation (vf). the patient was defibrillated and support continued. furthermore, the patient developed a hematoma with oozing at the impella access site. the doctor thought this occurred from the sheath exchange. manual pressure was held and heparin was withheld. the patient expired after 25.50 hours of impella support due to having ventricular tachycardia and the patient's family withdrawing care. | [5 empty] | The product was not returned for investigation. a data log review was not required for this case run. as per the given clinical information, the pump was placed successfully using a long peel-away sheath. the patient went into ventricular fibrillation (vf) after pump placement. additionally, the doctor reported difficulty repositioning the catheter due to a kink in the repositioning sheath caused by the angle of insertion. also, hematoma were noted at the impella insertion site, attributed to sheath exchanges, and the doctor applied manual pressure and re-dressed the site. the introducer damage was removed and added to the positioning issue after the investigation of clinical details. the cause of the positioning issue was use related to steep angle of insertion, which cause the kink in reposition sheath. the cause of the access site bleeding issue was use related since re-dressed the site and maintained the angle match. resolve the issue of growing hematoma. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12255 | 1220648-2025-25570 | 21152304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21152304&pc=OZD | 01/01/2025 | 03/03/2025 |  | 01/14/2025 |  |  | 01/01/2025 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493969 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella cp support. while on support, pump stop was noted on console. nurse reported they were in the process of transitioning patient to comfort measures only. patient expired shortly after. nurse reported pump stop was not the cause. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. no product was returned for analysis. the data logs showed a motor current high pump stop. there was a motor current spike with a concurrent drop in motor speed, drop in purge flow, and a rise in purge pressure. the logs showed failed restart attempts. the clinical details do not provide any additional information on pump stop but report that it did not occur during patient movement. the root cause of the pump stop was not determined as no product was returned and limited clinical information related to failure was provided. | B.1 revised to add adverse event due to medical safety officer review since the previous manufacturer device report number 1220648-2025-25570 was submitted. b.2 revised to add death and date of death due to medical safety officer review since the previous manufacturer device report number 1220648-2025-25570 was submitted. b.5 abiomed's medical safety officer review was completed and information was added since the previous manufacturer device report number 1220648-2025-25570 was submitted. h.1 revised to reflect death due to medical safety officer review since the previous manufacturer device report number 1220648-2025-25570 was submitted.
12256 | 1220648-2025-25571 | 21154188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21154188&pc=OZD | 09/20/2024 | 02/03/2025 |  | 01/14/2025 |  |  | 12/23/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella device used and would eventually expire. the patient was on medical floor waiting for surgical aortic valve replacement planned and began to decompensate. the patient was transferred to the intensive care unit and on dobutamine, intubated for airway protection, and then to the cardiovascular lab for an impella cp device placement. prior to the impella implant procedure, the patient had a left ventricular end-diastolic pressure (lvedp) of 40 and 20% ejection fraction. a biventricular aortic valve replacement was performed followed by the impella cp device implant via the left femoral artery. the impella implant was completed without complications, position verified, and the patient's hemodynamics immediately responded. the lvedp was now 20. the next day the nurse reported urine output changed to a tea color. the impella cp pump performance level was reduced to p-7. aortic valve replacement surgery was planned for the next monday (approximately three days later). the following day, the patient was weaning down at performance level p-6 with good outputs and cardiac index. urine output and color improved. however, two days later, the patient decompensated upon preparation for the operating room and expired. it was reported regarding the hemolysis, the patient had not recovered/the event had not resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12257 | 1220648-2025-25572 | 21154193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21154193&pc=OZD | 09/03/2024 | 02/05/2025 |  | 01/14/2025 |  |  | 12/26/2024 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548678 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular fibrillation with a possible relationship to the impella device used: patient went into ventricular fibrillation at 13:04 on (b)(6) 2024. patient was shocked and returned to normal sinus rhythm (nsr). patient went into ventricular fibrillation at 13:58 on (b)(6) 2024. patient was shocked with return to nsr. at 14:01 patient went back into ventricular fibrillation and was shocked again with return to nsr. the patient decompensated and expired. impella support was for 14 hours. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12258 | 1220648-2025-25573 | 21154210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21154210&pc=PYX | 01/01/2025 | 09/24/2025 |  | 01/14/2025 |  |  | 01/01/2025 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561573 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The investigation uncovered a placement signal not reliable alarm during data log review. | Us complainant reported the patient was on impella rp flex support and during support, there were high purge pressures. tissue plasma activator was added to the purge and support continued. the patient survived the event. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a placement signal not reliable alarm occurred nine hours into the case. the signal-to-noise ratio (snr) goes to 0, lamp steps up, contrast decreases, gain increases, and light remains stable. these 5s parameters are consistent with a sensor break. the product was not returned. the root cause of the optical signal issue is due to a damaged sensor. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated. | The investigation of high purge pressure has been completed. the pump was not retrieved for investigation, but the data logs were submitted and analyzed. the analysis indicated a motor current spike followed by an increase in purge pressure lasting over 5 minutes, which triggered high purge pressure alarms and purge system blocked alarms. the pump flow was reduced, and according to clinical information tissue plasminogen activator (tpa) was added to the purge, which resolved the high purge pressure alarms. these findings suggest symptoms consistent with biomaterial buildup. therefore, the root cause of the high purge pressure issue was most likely due to biomaterial. instructions for use as follows: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25573 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 revised to reflet the instructions for use that was inadvertently not included on the initial manufacturer device report number 1220648-2025-25573. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12259 | 1220648-2025-25574 | 21154273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21154273&pc=OZD | 01/01/2025 | 02/17/2025 |  | 01/14/2025 |  |  | 01/01/2025 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577597 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella cp. the procedure aborted due to patient anatomy and severe calcification. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, unable to pass impella due to patient severe peripheral artery disease (pad). e4 should have been left blank on manufacturer device report 1220648-2025-25574. h6 code 2422 and 4627 were reported incorrectly on manufacturer device report 220648-2025-25574. new codes was added to health effect - clinical code and health effect - impact code.
12260 | 1220648-2025-25575 | 21154277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21154277&pc=OZD | 12/31/2024 | 02/12/2025 |  | 01/14/2025 |  |  | 01/01/2025 | 01/14/2025 |  |  |  |  | 01/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197227 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported upon placing an automated impella controller (ai) on a stand during patient support, the console fell to the floor. it was noted by the nurse that the aic had fallen due to wobbly screws. pump support was uninterrupted, but the decision was made to replace the console. there was no report or indication of patient harm. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. the console logs were reviewed as a precaution but did not contain any data relevant to this complaint. console was inspected and tested after arriving in field service. the only damages observed from the reported fall was a damaged rlm. the rest of this console was structurally sound. the console functioned normally while testing with a known good pump and purge cassette kit. the aic cart was not sent back for investigation. the reported issue with the aic cart being wobbly was likely use related and not related to a defect of the cart itself. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2025-25575.
12261 | 1220648-2025-25509 | 21136734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21136734&pc=OZD | 12/24/2024 | 02/05/2025 |  | 01/13/2025 |  |  | 12/24/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541721 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma around the access site. the sheath was repositioned, and manual pressure was placed, but due to the initial stick, the physician stated that they were unable to post close the site. the bleed remained difficult to manage, and the decision was made to explant the device. femstop was placed to achieve hemostasis. a new impella cp pump was subsequently implanted via the opposite femoral artery for ongoing support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details provided.
12262 | 1220648-2025-25510 | 21136807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21136807&pc=OZD | 12/25/2024 | 02/10/2025 |  | 01/13/2025 |  |  | 12/25/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544907 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the procedure was aborted due to patient anatomy. the pump could not pass the innominate bifurcation due to tortuosity. the decision was made to abort the implant and placed an impella cp instead. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12263 | 1220648-2025-25511 | 21136899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21136899&pc=OZD | 12/26/2024 | 02/10/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient eventually expired on support. | The complainant reported a patient was on impella cp support. while on support, a patient had multiple runs of ventricular tachycardia and ventricular fibrillation requiring multiple shocks from patient implantable cardioverter defibrillator. the doctor believed the impella was too deep and potentially causing arrhythmias. the impella cp was pulled back and believed to be in a better position away from heart structures. there was no harm to the patient and the patient remains on support. |  |  |  |  | The investigation for the reported positioning issue and ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue and vt could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b2 selected death and added date of death, b5 patient outcome of death, d6b-date of explant, h1-updated to death, h6 impact code 1802. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12264 | 1220648-2025-25512 | 21137022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137022&pc=OZD | 12/26/2024 | 02/06/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483540 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and while repositioning the impella, a small perforation was noted in the left ventricle. the doctor was pulling the pump back, where the elbow was at the level of the aortic valve, when a small perforation was noticed. bleeding was quickly controlled, and perforation was repaired. the doctor reported the perforation was ¿small, about the size of the wire.¿ they were unsure if the impella 5.5 catheter or wire caused the perforation. blood products were given to the patient and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. the cause of the ventricle perforation issue was most likely improper positioning of the device, based on given clinical details.
12265 | 1220648-2025-25513 | 21137107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137107&pc=OZD | 12/26/2024 | 02/13/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563834 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had bleeding from the impella access site. a suture was placed and bleeding resolved. additionally, the patient had limb ischemia in the right lower extremity, and it was noted that the duct-dependent pulmonary circulation was clotted. the patient now has palpable pulses. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation for the limb ischemia and access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by placing a suture at the bleeding site.
12266 | 1220648-2025-25515 | 21137119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137119&pc=OZD | 12/26/2024 | 02/10/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567212 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that three days into impella 5.5 support, the patient developed a hematoma coinciding with oozing at their impella site. three units of blood were given to treat the condition and heparin was placed on a twenty-four hour hold. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post sterility inspection checks. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
12267 | 1220648-2025-25514 | 21137134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137134&pc=OZD | 12/19/2024 | 01/15/2025 |  | 01/13/2025 |  |  | 12/19/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116071 |  |  |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) will not go into maintenance mode. rlm keeps restarting. there was no report or indication of patient involvement. | [5 empty] | The automated impella controller was returned for evaluation, analysis and when tested the complaint was reproduced: the module would not fully boot-up, serial terminal interface would show scrambled data (¿garbage¿), and rlm kept rebooting. inspection of the miniusb connection between backstrap cable and rlm did not reveal any irregularities. inspection of the rlm and console¿s electronics (4-in 1 carrier, i/o [input/output] port assembly) did not produce any evidence as to the cause of rlm failures. a series of tests, performed with spare (known-good) 4-in-1, i/o assembly, rlm, microsd card and backstrap cable, allowed the determination that the only combination of components that allowed the system to work, was one that did not include original i/o panel and original backstrap cable. a product history review was completed, and no action was required as a result of this review. in conclusion, it was determined that both original i/o panel and original backstrap cable components were in some way compromised and contributed to the rlm failure. the module itself operated within specification. | The investigation for the console (aic) shutdown issue has been completed. the aic was returned for evaluation. when console was tested the the lab, the issue was reproduced: the module would not fully boot-up, serial terminal interface would show scrambled data (¿garbage¿), and rlm kept rebooting. inspection of the miniusb connection between backstrap cable and rlm did not reveal any irregularities. inspection of the rlm and console¿s electronics (4-in 1 carrier, i/o port assembly) did not produce any evidence as to the cause of rlm failures. a series of test, performed with spare (known-good) 4-in-1, i/o assembly, rlm, microsd card and backstrap cable, allowed us to determine, that the only combination of components that allowed the system to work, was one that did not include original i/o panel and original backstrap cable. it¿s been concluded that both these components were in some way compromised and contributed to the rlm failure. the module itself operated within specification. the cause of the issue was a defective cable. added d9-device available for evaluation checked yes corrections to the initial mfr report: d2 common device name updated d4-updated udi number.
12268 | 1220648-2025-25516 | 21137428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137428&pc=OZD | 11/07/2024 | 02/07/2025 |  | 01/13/2025 |  |  | 12/21/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening hemolysis with a possible relationship to the impella device used: persistent drop in hemoglobin despite packed red blood cell transfusions -positive hemolysis labs likely 2/2 impella causing hemolysis. hemoglobin baseline 11.7 nadir 6.1. the patient outcome is unknown. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right axillary artery access site hematoma with a probable relationship to the impella device used: patient intubated and sedated. blood noted from bilateral internal jugular line sites, patient's nose, right axillary impella insertion site (with small hematoma upon palpation). the patient outcome is unknown. the patient remains on impella support. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the access site bleed and the hemolysis. data logs generally showed the pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved. the patient was on dialysis during impella support. it was unknown if hemolysis was related to impella or patient condition as no product was returned for review. therefore, the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. d.6b explant date was added. h.4 revised as date was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25516. h.6 code 4643 was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25516.
12269 | 1220648-2025-25517 | 21137448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137448&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122738 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the aorta position waveform disappeared intermittently during impella support. the automated impella controller was replaced and the issue resolved. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. according to device and data analysis the cause of the automated impella controller (aic) issue was a software issue. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2025-25517. g2 report source; foreign was missing from manufacturer device report 1220648-2025-25517.
12270 | 1220648-2025-25518 | 21137483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137483&pc=OZD | 12/23/2024 | 01/28/2025 |  | 01/13/2025 |  |  | 12/24/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the patient had no distal pulse down the right leg. impella was removed the next day. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. e.1 email was reported incorrectly on the initial manufacturer device report number 1220648-2025-25518. the correct email is unknown. phone number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25518. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12271 | 1220648-2025-25519 | 21137557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137557&pc=OZD | 12/26/2024 | 02/05/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, while adjusting the pump, the pigtail was caught in a chordae. the doctor pulled the pump across the aortic valve and was unable to reposition. the pump was explanted and replaced. the patient remained stable and there was no reported patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12272 | 1220648-2025-25520 | 21137785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137785&pc=OZD | 06/21/2023 | 10/31/2025 |  | 01/13/2025 |  |  | 06/21/2023 | 01/13/2025 |  |  | 03/31/2025 |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342457 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Insufficient Information | [12 empty] | Other | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was admitted to the hospital where it was discovered she had a total occlusion of the left anterior descending coronary artery (lad). the patient's lad was opened, the patient remained stable. prior to leaving the catheterization lab, the patient went into ventricular tachycardia and fibrillation where she was rescued by her automatic implantable cardioverter defibrillator. the patient was transferred to the icu where she continued to decline and decompensate overnight. the family wished to escalate care of the patient, and the physician decided to place the impella to wean the inotrope support as a last effort, and an impella cp device was implanted for mechanical circulatory support. it was reported upon connecting the impella cp to the automated impella controller (aic) with the serial number (b)(6), there was a message for "optical sensor not supported". the pump was disconnected and reconnected several times, however this did not resolve the issue. the aic was swapped with a backup unit, and the pump was successfully connected. it was noted after impella supported was initiated, the patient was in good and stable condition and was brought back to the icu. in the icu, the aic alarmed for placement signal low. stat echocardiography was ordered for the patient, and it was observed the impella cp was too deep within the ventricle. the physician was able to successfully pull the impella back per best practices, and "great" waveforms on the aic were noted. as reported on (b)(6) 2023, the patient continued to decline despite maximal mechanical and medical support with worsening liver and renal function. there was a need for escalation of vasopressor and inotropic support on the morning of (b)(6) 2023 with poor prognosis. it was noted that the impella was placed and functioning well, however, due to multiple metabolic derangements the patient continued to declined. the collective decision was made by the cardiology and pulmonary teams in conjunction with the family to transition the patient to do not resuscitate comfort care. therefore, the decision was made to withdraw mechanical circulatory support and the impella was removed at the bedside by dr. (b)(6). patient tolerated the removal well and hemostasis achieved. the patient later expired. although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise after a malposition event of the impella which may have contributed to lower than expected support. the malposition of the impella may have been resultant of the optical sensor issue reported on the first aic that was in use. | [5 empty] | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-25212 for the automated impella controller with sn (b)(6). the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: based on the clinical details, case start was reported on ic7877 with "optical sensor not supported" message during case start. disconnecting and reconnecting the pump did not resolve the issue. after switching the pump to a backup console (ic7786), it was correctly recognized and optical placement signal was available. there was no problem was found with the pump and it is apparent the console was the potential cause of the issue. positioning issues: the clinical details note that the pump was noted to be deep when an echo was done to check a placement signal low alarm and was pulled back successfully. the root cause of the positioning issues was not determined due to insufficient clinical details and unreturned product. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated.
12273 | 1220648-2025-25521 | 21137911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21137911&pc=OZD | 12/27/2024 | 02/13/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579637 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted and started but flows never exceeded 1.0liters per minute. the impella cp was exchanged for a new one. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. returned complaint cp pump was visually inspected. a big chunk of biomaterial observed around impeller and purge gap. no cannula kink observed. no broken blade found. data logs confirm low flows and suction alarms. motor current (mc) spike observed. the root cause of the low pump flow issue was most likely biomaterial ingestion.
12274 | 1220648-2025-25522 | 21138123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21138123&pc=OZD | 12/25/2024 | 02/07/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547784 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support when there was suction events. the suction resolved without the use of blood products. however, the doctor did not want to reposition the pump. the pump was left as is and support was continued. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related.
12275 | 1220648-2025-25523 | 21139144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21139144&pc=OZD | 12/28/2024 | 02/14/2025 |  | 01/13/2025 |  |  | 12/28/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583084 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that eight days into support, an impella 5.5 pump began to experience high purge pressure. the purge cassette was replaced, but the issue persisted. as a result of the ongoing issue, the staff was advised to initiate a tissue plasminogen activator protocol to lower the purge pressure. there was no patient harm and support with the implanted pump remained ongoing. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. no product was returned. the data logs show that purge pressure was stable until there was an air in purge alarm at 12/28 at 3:22 pm and de-airing was done. this led to a drop in purge pressure and there was no purge flowing as the lines were not clicked on. this lasted for 30 minutes as the were changing the bag, cassette, and de-airing but they kept getting air in purge alarms. a minute after they completed the procedure, the purge pressure rapidly rose to 1050 mmhg over the course of 1 minute, triggering the high purge pressure alarms. they attempted to de-air and change the cassette again, but the purge pressure remained high for the rest of the case of the case, approximately 19 hours. without sufficient clinical information and lack of product returned, the root cause of the high purge pressure cannot be determined since it is unknown why the air in purge alarm was occurring leading to long troubleshooting that did not resolve which eventually led to high purge pressure.
12276 | 1220648-2025-25524 | 21139401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21139401&pc=OZD | 12/28/2024 | 02/13/2025 |  | 01/13/2025 |  |  | 12/28/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580373 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, there were multiple rapid successions of a variety of alarms: pump stopped, preventing retrograde flow, pump failure, controller failure, and purge blocked. the team attempted to reprime and change the controller which, did not work. ultimately a new pump was needed. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. the data logs showed that the pump had alarms but no spike to 3000. it was unable to be started and was attempted a few times. the product was returned, and the pump had dried dextrose around the purge gap. it was soaked in tergazyme and the dextrose was removed. additionally, the red lumen was still in the pump. the red lumen was removed and the pump was able to operate within specification with no alarms. the impella cp instructions for use (ifu) states "remove the easyguide lumen by gently pulling the label in line with the catheter shaft while holding the impella. if you suspect that a portion of the red lumen remains in the catheter, do not use the impella catheter." the root cause is due to a red lumen issue as it was not removed based on the returned product.
12277 | 1220648-2025-25525 | 21139699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21139699&pc=OZD | 12/23/2024 | 02/25/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537493 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella 5.5. the patient had a history of arch replacement. after descending replacement for aortic dissection, cardiopulmonary bypass became difficult, and an attempt was made to insert a 5.5 catheter, but it got caught on the artificial blood vessel in the arch and could not be advanced, so insertion was abandoned. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition because the impella was unable to pass through the vascular arch.
12278 | 1220648-2025-25526 | 21139711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21139711&pc=OZD | 10/28/2024 | 01/13/2025 |  | 01/13/2025 |  |  | 12/30/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478158 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and after the implant, there were optical signal issues. the waveform continued to appear and disappear. since there were no problems with the dynamic support of the impella itself, the pump and automated impella controller (aic) were not replaced. the site would like the pump and the aic analyzed. the investigation was completed and found that the controller error and loss of placement signal was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. no patient harm has been reported. | [5 empty] | The impella device was returned and evaluated. the logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. the real time logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. the controller error and loss of placement signal was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
12279 | 1220648-2025-25527 | 21141999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21141999&pc=OZD | 12/27/2024 | 02/27/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  | 11/30/2025 |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424494 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was exchanged due to thrombus. a high purge pressure system blocked alarm occurred. troubleshooting was done unsuccessfully. duration of impella support was for five hours and the impella cp was exchanged for a new one. the procedural outcome was the patient the survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of high purge pressure and thrombus has been completed. the pump was returned for investigation. no damage was noted on the returned product. the impeller was clean, but some biomaterial was found in the purge gap. high purge pressure was reproduced during testing, and the pump was also unable to start. the catheter was cut near the motor housing, and purge splashing was observed from the catheter, which proved a blockage in the motor housing. additionally, blood was found clotted on the motor side of the purge line. data log shows the purge pressure rising over 22 hours with a similar gradual trend in motor current. the pump flow also became saturated at the end of the case run. the cause of the high purge pressure issue was biomaterial, based on returned product investigation. d9 date device returned to manufacturer added.
12280 | 1220648-2025-25528 | 21143412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21143412&pc=OZD | 12/27/2024 | 02/12/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support the patient had access site bleeding. manual pressure was held, femstop applied and heparin was withheld. additionally, the patient lost distal pulses to left femoral artery. the decision was made to explant the cp and implant an impella 5.5. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25528 and was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12281 | 1220648-2025-25529 | 21145544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145544&pc=OZD | 12/26/2024 | 02/12/2025 |  | 01/13/2025 |  |  | 12/26/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531550 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  | Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that twenty two days into patient support, an impella cp pump experienced a high motor current and suddenly stopped. attempts to restart the pump were unsuccessful and the decision was made to explant the device. following removal, a clot was discovered hanging off the tip of the pigtail. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the pump stop and thrombus has been completed. no product was returned. log review showed a slight gradual increase in motor current, purge pressure, and pump flow in the last two days of the case. the purge pressure did not exceed 800 mmhg. no major motor current spikes were observed in this time. the p-level was changed from p-5 to p-7, and after being at p-7 for ~30 minutes, there was a small sudden motor current increase followed by a pump stop. there were then multiple unsuccessful attempts to restart the pump. there was a clot found on the pigtail during pump explant. the pump stop also occurred on day 23 of support, which is beyond the 8 day impella cp indication for use. the root cause of the pump stop was not determined since the product was not returned. the root cause of the thrombus could not be determined without additional information. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
12282 | 1220648-2025-25530 | 21145545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145545&pc=OZD | 12/27/2024 | 02/04/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568653 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed bleeding and a hematoma around their axillary access site during impella 5.5 support. one unit of packed red blood cells were given to the patient to manage the condition. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25530.
12283 | 1220648-2025-25531 | 21145607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145607&pc=OZD | 12/25/2024 | 02/19/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553810 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in reported that impella cp was inserted to support a patient with acute myocardial infarction and cardiogenic shock. the impella was removed due to the patient improvement. after removal the patient when into shock immediately and extracorporeal membrane oxygenation was inserted. left main truck obstruction was found and treatment was performed. according to the doctor, a blood clot attached to impella could not be denied. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25531 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12284 | 1220648-2025-25532 | 21145622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145622&pc=OZD | 12/27/2024 | 02/13/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574827 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was inserted on the patient with cardiomyopathy. while the impella was on support flow saturated and pump stop occurred. unsuccessful attempts were made to restart the 5.5. the 5.5 was pulled back across the aortic valve and the patient was cannulated for extra corporeal membrane oxygenation. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop and saturated flow has been completed. product was returned and evaluated. the pump passed all electrical testing. there was blood found in the purge line. high purge pressure reproduced and was resolved once cutting the catheter at ~70 cm, where the blood had reached in the purge line. upon initial inspection, there was no biomaterial found around the purge gap, impeller, or cannula. the motor housing was torn down and biomaterial and blood was found in between the rotor and stator as well as just under the purge gap. data logs were reviewed. the first pump stop occurred on the night of 12/27 with the final pump stop on 12/28. log review showed a gradual increase in purge pressure (only reaching ~8000 mmhg) from 12/24 to 12/26. when the logs begin again on 12/27, the purge pressure was ~1750 mmhg with a purge flow of 0. the pump was running at this point until multiple motor current spikes and "impelled stopped - motor current high" alarms. the pump was able to be restarted for about another hour before the mc spikes and alarms occurred again early morning on 12/28 and could not be restarted. flows were on the lower side from the time the purge pressure started gradually increasing on 12/24. when logs begin again on 12/27, the pump flow is completely saturated until the end of the case. there are no mc spikes or abnormalities seen prior to 12/27. the high purge pressure was due to biomaterial that led to blood ingress and the pump stop. the root cause of the pump stop was biomaterial based on evidence in the logs and biomaterial found inside the motor and just under the purge gap. the root cause of the saturated flow was biomaterial based on evidence in the logs and biomaterial found inside the motor and just under the purge gap.
12285 | 1220648-2025-25533 | 21145625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145625&pc=OZD | 12/27/2024 | 02/04/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564548 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient had intermittent palpable/dopplerable pulses in bilateral lower extremity but looked like mottling was present. bedside nurse reported dopplerable-palpable pulses when patient was given a 500ml bolus of albumin. disseminated intravascular coagulation panel ordered due to signs of clotting and platelet count of 8 after two units of blood were given. additionally, the patient developed sepsis and the impella was removed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components . ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the limb ischemia, sepsis, and thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia, sepsis, and thrombocytopenia could not be determined.
12286 | 1220648-2025-25534 | 21145632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145632&pc=OZD | 12/27/2024 | 02/13/2025 |  | 01/13/2025 |  |  | 12/27/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547308 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and it was aborted due to patient anatomy. no further details were provided. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy.
12287 | 1220648-2025-25535 | 21145636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145636&pc=OZD | 12/28/2024 | 02/21/2025 |  | 01/13/2025 |  |  | 12/28/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569977 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella 5.5 support and during support, there was high purge pressure. tissue plasma activator was added to the purge. it was recommended to replace the pump, however staff monitored and support continued. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. neither the impella device nor the data logs were returned for evaluation. additionally, clinical details provided were limited to determine the root cause. therefore, the root cause of the high purge pressure could not be determined. added h6 impact code 4644
12288 | 1220648-2025-25536 | 21145637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145637&pc=OZD | 12/28/2024 | 02/10/2025 |  | 01/13/2025 |  |  | 12/28/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and the patient ambulated and there were impella in aorta alarms. the impella was confirmed to be in the aorta. the patient was already being weaned from the impella and it was removed. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. product was not returned for review. the cause of the positioning issue was most likely use related per clinical information. d4 model number was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25536.e4 should have been left blank on manufacturer device report 1220648-2025-25536.
12289 | 1220648-2025-25537 | 21145639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145639&pc=OZD | 12/28/2024 | 02/04/2025 |  | 01/13/2025 |  |  | 12/28/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584979 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient with a history of ventricular tachycardia (vt) kept having vt while on impella 5.5 support. on day eight of impella support while the patient was walking, two vt episodes occurred, leading to two shocks from the implantable cardioverter defibrillator. under echocardiography guidance the impella was pulled back. the team will continue monitoring. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
12290 | 1220648-2025-25538 | 21145640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21145640&pc=OZD | 12/29/2024 | 02/27/2025 |  | 01/13/2025 |  |  | 12/29/2024 | 01/13/2025 |  |  |  |  | 01/13/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516539 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, purge pressure increased. there were no kinks in the purge lines. no improvement after replacing the purge cassette. the automated impella controller (aic) was replaced, but there was no improvement, so the pump was replaced. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. there were no data logs returned. the product was returned and there were no abnormalities noted. the purge gap was clean. the pump was purged and high purge pressure did not reproduced. purge fluid exited the purge gap and pressure was stable around 550 mmhg. the root cause of the high purge pressure could not be determined because data logs were not returned and high purge pressure was not reproduced on the returned product. d.9 device return date/information was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25538 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12291 | 1220648-2025-25469 | 21123390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21123390&pc=OZD | 12/23/2024 | 01/31/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493850 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had discolored right foot. post cp removal the right leg and foot appeared to have strong pulse. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12292 | 1220648-2025-25470 | 21123607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21123607&pc=OZD | 12/24/2024 | 02/21/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534621 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that the impella 5.5 had high motor current and left ventricle waveform were abnormal. the impella stopped and when it was removed a blood clot was found to have adhered. the patient is stable and survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop and thrombus has been completed. the impella device was not returned for evaluation. due to lack of product or data logs returned, the root cause of the pump stop cannot be confirmed. as the necessary information was not provided, the root cause of the thrombus was not determined g2 report source; foreign was missing from manufacturer device report 1220648-2025-25470.
12293 | 1220648-2025-25471 | 21123925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21123925&pc=OZD | 12/26/2024 | 02/12/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569059 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, impella had low pump flows. impella was weaned and explanted following low pump flows. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. the returned complaint pump was visually inspected. a tiny amount of biomaterial was observed around the purge gap. no broken blade/cannula kink was observed. inlet cage damage was observed. the returned pump was connected to automated impella controller to reproduce the reported low pump flow issue. the pump run for 10 minutes on p-9. no low pump flow was reproduced. data log review showed many suction alarms occurred. motor current spike was observed followed by dropping in flow at p-2. no positioning issue alarms were observed. clinical stated that biomaterial was noted in inlet cage. the root cause of the low pump flow issue most likely was biomaterial ingestion. it looked like the inlet cage damage occurred during removal or return of the device for analysis since no detail stated regarding this damage in the clinical description.
12294 | 1220648-2025-25473 | 21124172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124172&pc=OZD | 12/26/2024 | 02/04/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579072 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient rejected for an orthotopic heart transplant was implanted with an impella 5.5 device for mechanical circulatory support. approximately six days after support the patient reportedly experienced a stroke. it was noted the stroke was successfully managed without any surgical intervention. the patient was reported to be stable following the event. as per the plan of care update, the patient was successfully weaned and the impella 5.5 device was explanted seven days later with an outcome of the patient having survived after explant. no further information was noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
12295 | 1220648-2025-25474 | 21124291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124291&pc=OZD | 12/13/2024 | 02/03/2025 |  | 01/10/2025 |  |  | 12/13/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. when the pump was removed, clot was noted on the pump at the inlet and outlet. prior to removal of the 14 fr peel away sheath, the physician (md) could not aspirate back on it through the side arm with the 3-way stop cock on it. md took the dilator with a syringe of saline connected and pulled negative while inserting the dilator through the valve to aspirate the sheath. this was done 2 times. no clot was aspirated through this portion of the sheath. the md proceeded to close the groin site as he would normally over a wire and hemostasis was achieved. once the sheath was removed, the side arm was cleared of small clot by flushing forward onto a sterile towel on the field and outside of the patient¿s body. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 should have been adverse event only on manufacturer device report 1220648-2025-25474. e4 should have been left blank on manufacturer device report 1220648-2025-25474. h6 medical device problem code 1528 was reported incorrectly on manufacturer device report 1220648-2025-25474. new code was added to medical device problem code
12296 | 1220648-2025-25475 | 21124298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124298&pc=OZD | 10/21/2024 | 01/27/2025 |  | 01/10/2025 |  |  | 12/17/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right femoral artery thrombosis with a definite relationship to the impella device used: intra-operative ultrasound of right femoral artery with evidence of thrombosis within the artery with limited flow. pre-closed perclose sutures are in place but not tied down with visible knot at the skin. femoral artery cutdown performed, found to have a thrombus bulging from the arteriotomy. proximal and distal control obtained followed by thrombectomy with thrombus removed proximally. great inflow and back bleeding at completion. femoral artery was repaired primarily in a transverse fashion. hemostasis achieved. the patient recovered with no sequelae. the patient survived the event. | [5 empty] | The investigation of thrombus and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25475. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12297 | 1220648-2025-25472 | 21124362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124362&pc=OZD | 12/26/2024 | 02/20/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163557 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller had an impella defective alarm. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the pump not detected has been completed. console logs from the reported day of event show initially an interruption of data streaming, which seemingly recovered after successful handshake, and files appear to be transferred from aic to rlm, however data transfer stops, without icb process detecting connection issues. corresponding rlm journals show that rl00752 unexpectedly rebooted few times, but also kept resetting usb stack every 1m50s (1586 times throughout the case) due to perceived loss of usb communication from aic to rlm. there was also evidence of microsd card corruption, e.g. the rssi for the wifi network was satisfactory, and did not cause disconnections from the network. console logs did not contain any information relevant to the reported alarm sound issues, but revealed rfid communication errors which affected purge cassette detection, perceived as purge cassette being disconnected. despite aic falsely not detecting purge cassette, staff was still able to change purge cassettes, although with alarms #420 ¿complete the procedure¿ (as evidenced by ltlogs and rtlogs). hemodynamic support was unaffected by purge issues. console was returned for evaluation. upon functional testing, it was unable to reproduce unexpected rlm reboots or interruptions of data streaming, and the unit successfully connected to engineering wifi and data was sent from aic to rlm (we could not confirm if data was successfully sent to rlfs2). it is most likely that intermittent connection to impellaconnect portal was caused by interruptions of data streaming due to corrupt microsd card causing unexpected rlm restarts. no sound issues were observed - there were no audible anomalies of alarm sounds or pbm/battery buzzer, and sound levels were measured to be within specification (> 80db). root cause was undetermined, but in case alarm sound issue is intermittent, it is recommended to replace speaker assembly. purge cassette detection issues were confirmed, as the console was intermittently unable to detect a known-good purge cassette, due to malfunction of the rfid circuitry the root cause of the purge issue was a defective component. the root cause of the buzzer alarm inaudible could not be determined. added d9 device return date corrections to the initial mfr report: h4 should have been no. h8-should have been reuse.
12298 | 1220648-2025-25477 | 21124372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124372&pc=OZD | 12/25/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/25/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560163 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient was coughing and sat up which pulled the impella back into the aorta. echocardiography confirmed the pump was in the aorta and the decision was made to explant the impella. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. the root cause of the positioning issue most likely was due to patient movement when the patient started to cough and attempted to sit up which pulled the impella back into the aorta. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25477 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12299 | 1220648-2025-25476 | 21124458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124458&pc=OZD | 12/25/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/25/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572151 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient developed a large hematoma at their impella site. manual pressure and femstop were utilized to achieve hemostasis, however the patient's urine had become brown in color. due to the risk of hemolysis, the pump was explanted and the patient was escalated to impella 5.5 support. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The returned compliant cp device visually inspected. dried blood observed inside cannula. no biomaterial observed. no broken bladed or sharp edge found. clinical stated: patient had multiple system organ failure and coagulopathy conditions. also he the patient had a large hematoma to impella site and hematoma to right radial site. echo showed depth of 6+cm, impella was pulled back to a final position 4.9cm, d/t ectopy and the catheter not appearing to be stable decision was made not to pull back any further for fear of it popping out of the lv. the root cause of the hemolysis issue most likely was patient condition related since the patient had multiple system organ failure. e4 should have been left blank on manufacturer device report 1220648-2025-25476.
12300 | 1220648-2025-25478 | 21124469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124469&pc=OZD | 12/26/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553068 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient became hypertensive with suction alarms and blood products given. sterile sleeve advanced and touhy boost valve tightened and t lock secure. placement verified with trans-esophageal echocardiography (tee). after blood products were given, the impella increased back to p9. additionally, urine was noted to be pink. placement verified per echocardiography. the urine was red. no plasma free hgb drawn. no hemolyzed labs. suspected hemolysis. patient was made do not resuscitate, and due to failure to thrive decision was made by family to withdraw care. the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Additional information provided in h6 and h10. the investigation into the reported issues has been completed. device history review confirmed pump passed all post sterile inspection checks. data logs confirmed the failure mode and clinical narrative. logs showed low flow occurred during the case that triggered auto suction response & multiple suction alarms. clinical information states pump was in proper position during the incidents. no motor current spikes outside p level changes were observed. clinical information also states that after giving blood able to turn pump back to p-9. no other issues were found on the logs. the device was not returned for evaluation. based on clinical description and data review, the root cause of the reported issue of low flow was most likely patient condition. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12301 | 1220648-2025-25479 | 21124507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124507&pc=OZD | 12/20/2024 | 01/10/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  | 05/31/2025 |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345488 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella 5.5 support. after the pump was explanted, the patient decompensated, and aortic insufficiency was observed. the patient was emergently taken to the hybrid operating room to receive a transcatheter aortic valve replacement (tavr) procedure. patient stabilized after receiving a tavr. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
12302 | 1220648-2025-25480 | 21124569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124569&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/21/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had hemolysis, which was confirmed via labs. however, the patient passed away the same day. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h5 was inadvertently omitted from manufacturer device report 1220648-2025-25480. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12303 | 1220648-2025-25481 | 21124679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124679&pc=OZD | 12/28/2024 | 10/01/2025 |  | 01/10/2025 |  |  | 12/28/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569293 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The customer reported that the device exhibited an optical signal issue. | The user facility reported a patient with an extensive cardiac history including non-ischemic cardiomyopathy with an ejection fraction of 10% was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced ventricular fibrillation followed by low pump flow. approximately ten days after start of support, the patient went into ventricular fibrillation (vf) that morning; the patient's implantable cardioverter defibrillator was unable to shock out of the vf rhythm. the patient was therefore then defibrillated and converted. at this time, the patient was reported intubated, and the mcs level remained at p-8 while the patient recovered from the prolonged vf. the following day, it was noted there were no further arrhythmias since the previous day cardioversion. approximately six days later, suction alarm issues were noted. limited echocardiogram confirmed good positioning. the patient was given two liters of fluid and remains on impella mcs while awaiting a heart transplant. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the investigation for optical signal issue has been completed. data logs were reviewed and no hard failures of the optical sensor were observed. ¿impella position unknown¿ alarms occurring at the same time as suction alarms at time of reported issue. pump flows were slightly out of range for the p-level. ¿impella position unknown¿ alarms were resolved when pump was turned back up to p-8. clinical details noted that ao was adjusted at time of reported optical issues. additionally, patient had received 2l of fluid due to suction alarms occurring at the same time and pump position was confirmed via echo. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections to the mfr report: h6 med dev prob code 2917 added. | The impella device was not received from the customer as the device remains implanted and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow was most likely patient condition related as additional volume resolved the suction alarms. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6b added.
12304 | 1220648-2025-25482 | 21124907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21124907&pc=OZD | 12/22/2024 | 02/13/2025 |  | 01/10/2025 |  |  | 12/22/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578934 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and it was noted that post repositioning the impella, the patient had less ectopy. it was also reported that there was multiple impella motor stop alarms detected. attempts to restart were unsuccessful. the impella cp was replaced with an impella 5.5 and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the ectopy and the pump stop has been completed. data logs were reviewed and as reported, the pump stopped. leading up to the event, the only abnormality was a slight trend up in the purge pressure, but not to the level that would trigger alarms. then, there was a sudden spike in the motor current to 1400 ma that remained elevated, and high motor current pump stop alarm was triggered. the pump attempted to restart a number of times, but was not able. the pump was on the 10th day of use, and per design trace matrix, the impella cp's design life is eight days. the returned product was investigated and a purge gap was large and filled with biomaterial. considering there were no signs of biomaterial ingestion in the logs and the pump was used beyond the design life, the purge gap was most likely growing due to bearing wear, which caused the sudden spike in motor current. as the purge gap grew, it allowed for biomaterial to accumulate as a result. based on data logs and returned product the root cause of the pump stop was determined to most likely be bearing wear. the root cause of the ectopy could not be determined without additional information. d.9 date device returned/information was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk: percutaneous coronary intervention: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk: percutaneous coronary intervention: section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12305 | 1220648-2025-25483 | 21125103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125103&pc=OZD | 12/28/2024 | 02/12/2025 |  | 01/10/2025 |  |  | 12/28/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and went into ventricular tachycardia and subsequently expired. the impella cp device was successfully implanted and following the implant the pump "dove" too far into the left ventricle causing the patient to go into ventricular tachycardia. the patient was shocked one time to regain a stabile rhythm. percutaneous coronary intervention with four drug-eluting stents to the left anterior descending and left circumflex arteries was completed. support was continued for one more day. the patient reportedly expired the following day. it is unknown if the ventricular tachycardia event contributed to the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12306 | 1220648-2025-25484 | 21125131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125131&pc=OZD | 12/26/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577608 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preparation of an impella cp pump for mechanical circulatory support, the staff had issues with priming the device. it was stated that the device did not properly primed and that purge fluid was not seen exiting the catheter. shortly after, the pump was inadvertently started while outside of the patient with the red lumen still in place. the motor current spiked and the pump stopped. the staff was advised to replace the pump, but as no spare could be found, an intra-aortic balloon pump was subsequently placed for ongoing support. there was no patient harm. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported priming issue and pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided noted the following: pump stop / did not start: clinical details indicate that the pump was started outside the patient's body before they had removed the red lumen. based on the clinical details provided, the root cause of the pump did not start was determined to most likely be a red lumen issue. priming issue: since the product was not returned and data logs could not confirm purge flow, the root cause of the priming issue could not be determined.
12307 | 1220648-2025-25485 | 21125185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125185&pc=OZD | 12/22/2024 | 02/21/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493712 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had hemolysis. the hemolysis did resolve but the intervention was unknown. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. data log was not available to review. the cause of the hemolysis issue could not be determined since the complaint device was not returned for analysis and due to lack of clinical details.
12308 | 1220648-2025-25486 | 21125293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125293&pc=OZD | 12/28/2024 | 02/14/2025 |  | 01/10/2025 |  |  | 12/28/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567116 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock was implanted with an impella cp device was mechanical circulatory support (mcs). the patient was planned for weaning off the impella mcs and explant, pending diagnostic results for potential percutaneous coronary intervention. when the physician was floating the swan, the automated impella controller (aic) was knocked over and fell. the purge tubing had a hole, and a change out was advised. the aic had a superficial crack on the side but appeared to be properly functioning. the patient was stable following the event. approximately three days later the patient underwent balloon angioplasty to a circumflex artery. the following day the patient was successfully weaned and the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for automated impella control serial number (b)(6) with date of event 28-dec-2024 and identifier ending in (b)(6). | The investigation of product damage (bond failure) has been completed. the product was not returned for investigation. without the product or the pictures of the damage, it is not possible to investigate the failure. therefore, the root cause of the purge leak issue was not determined since product was not returned for investigation. e4 should have been left blank on manufacturer device report 1220648-2025-25486.
12309 | 1220648-2025-25487 | 21125346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125346&pc=OZD | 12/28/2024 | 12/01/2025 |  | 01/10/2025 |  |  | 12/28/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock was implanted with an impella cp device was mechanical circulatory support (mcs). the patient was planned for weaning off the impella mcs and explant, pending diagnostic results for potential percutaneous coronary intervention. when the physician was floating the swan, the automated impella controller (aic) was knocked over and fell. the purge tubing had a hole, and a change out was advised. the aic had a superficial crack on the side but appeared to be properly functioning. the patient was stable following the event. approximately three days later the patient underwent balloon angioplasty to a circumflex artery. the following day the patient was successfully weaned and the impella cp device was explanted. | [5 empty] | The investigation for the reported issue has been completed. the cause of the aic - damage before therapy could not be determined since product was not returned. events of this type will be monitored for trends. h6 investigation type, findings and conclusions. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for the impella cp serial number (b)(6) with date of event 28-dec-2024 and identifier ending in (b)(6).
12310 | 1220648-2025-25488 | 21125390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125390&pc=OZD | 08/31/2024 | 01/31/2025 |  | 01/10/2025 |  |  | 12/20/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527278 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular fibrillation with a "possible" relationship to the impella device used. during impella supported percutaneous coronary artery procedure, the patient developed ventricular fibrillation and was shocked two times and received pne round of cardiopulmonary resuscitation before return to sinus rhythm. amiodarone drip was started, and the patient was taken to intentsive care unit post procedure on support. the patient recovered with no sequelae. additional information was obtained by the physician. post stent placement, once the impella was already in place, is when the patient went into ventricular fibrillation and required a direct current cardioversions. the impella was properly placed and appropriately running. the physician did not feel the ventricular fibrillation was related to the impella as it was already in place and did not move. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12311 | 1220648-2025-25489 | 21125514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125514&pc=OZD | 12/13/2024 | 02/25/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547037 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was not able to be placed on a patient with acute myocardial infarction / cardiogenic shock who already had ascending replacement. due to the curvature of the aorta, the catheter could not be delivered from the anastomosis area. the cp insertion was aborted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition related due to the angulation of the aorta the impella could not be advanced from the anastomosis area and the shaft was bent. d.4 revised unique identifier (udi) # as it was reported incorrectly in the initial manufacturer device report number 1220648-2025-25489.
12312 | 1220648-2025-25490 | 21125673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125673&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579678 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. delivery issues were noted due to patient anatomy. the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided noted that the root cause of the delivery issue was most likely patient condition related, since the impella could not pass through the vasculature as per the clinical description provided.
12313 | 1220648-2025-25491 | 21125723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125723&pc=OZD | 12/23/2024 | 02/25/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  | 02/28/2026 |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492672 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a low purge pressure alarm appeared during impella cp support. the purge cassette was replaced during troubleshooting, but the issue remained. a visual inspection of the red plug and side arm connections was performed and a light connection was discovered. upon securing the connection, the issue resolved. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak-pump has been completed. impella cp was returned for evaluation. pump was returned cut at catheter. purge sidearm was returned taped at purge filter. a large crack in in the purge filter was found on returned pump. pump was purged and leak was reproduced and coming from the cracks in the filter and filter holes. clinical details noted that clinical team reached out to csc regarding "low purge pressure" alarms occurring. dextrose concentration had been increased and purge cassette was changed. purge leak eventually resolved. data logs were reviewed and lt log of case shows no ¿low purge pressure¿ alarms being triggered. purge pressure and purge flow remain stable throughout support. the root cause of the purge leak - pump was most likely due to a damaged sidearm filter.
12314 | 1220648-2025-25492 | 21125796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125796&pc=OZD | 12/23/2024 | 01/28/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568686 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was suspected to have community acquired sepsis. the patient was started on continuous rental replacement therapy. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25492. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25492 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25492 was submitted in accordance with updated procedures.
12315 | 1220648-2025-25493 | 21125937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21125937&pc=OZD | 12/26/2024 | 02/12/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574088 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Obstruction/Occlusion | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was noted that the patient had a calcified anatomy that prevented successful advancement of the device. the implant was aborted. the patient passed away in the procedural area subsequently. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature as per the clinical description provided.
12316 | 1220648-2025-25494 | 21126076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126076&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Other | [5 empty] | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Vascular Dissection | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Johnson and johnson posted on facebook concerning an impella product and there was a comment on the post that stated, "an impella was used on me in 2020 it slipped and nick an artery causing me to have a stroke,paralyzed my entire left side. still partially paralysed." no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 address is unknown. h.4 device manufacture date is unknown. | The investigation for the reported stroke and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke and vascular damage could not be determined.
12317 | 1220648-2025-25495 | 21126184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126184&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/23/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561715 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant and during the implant, the patient began to decompensate. cardiopulmonary resuscitation was performed and multiple attempts were made to place the impella. however, the impella 5.5 would not advance due to patient anatomy and vessel size. the implant was aborted and venoarterial extracorporeal membrane oxygenation was inserted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient condition because they were unable to further advance the pump due to patient anatomy and vessel size. g.1 revised reporting contact fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25495. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12318 | 1220648-2025-25496 | 21126471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126471&pc=OZD | 12/25/2024 | 05/06/2025 |  | 01/10/2025 |  |  | 12/25/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573699 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis four days into impella cp support. urine colour changed from amber to brown. the patient was escalated to impella 5.5 support in response to the condition. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the returned complaint cp device visually inspected. no biomaterial, no broken or sharp edge observed. data log reviewed: no motor current (mc) spike observed. quick resolved position alarms occurred in ventricle and aorta. no placement signal trend observed. hemolysis test conducted with result of mih=12.75<20, so no hemolysis related to impella can be confirmed after testing. no patient condition related confirmed. the cause of the hemolysis issue cannot be determined due to lack of clinical details and impella passed the hemolysis testing.
12319 | 1220648-2025-25497 | 21126520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126520&pc=OZD | 12/21/2024 | 02/25/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528386 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the following day the patient was reportedly bleeding from the access site because the sheath was inserted through the same blood vessel that was used for extracorporeal membrane oxygenation cannulation. consequently, site was sutured, and blood products (type and number of units unknown) were administered. the patient was reported to be "stable and nimble" following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. d.4 revised serial number as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25497. e.1 revised first name as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25497.
12320 | 1220648-2025-25498 | 21126752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126752&pc=OZD | 12/25/2024 | 02/06/2025 |  | 01/10/2025 |  |  | 12/25/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590750 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had hemolysis and loss of distal pulses. volume resuscitation and optimal impella placement optimization was attempted. however, the team elected for an impella explant and an intra-aortic balloon pump was inserted. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of hemolysis was most likely due to pump was not in the proper position. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
12321 | 1220648-2025-25499 | 21126762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126762&pc=OZD | 12/24/2024 | 02/04/2025 |  | 01/10/2025 |  |  | 12/24/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585420 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the right femoral venous sheath noted to be through and through vessel on computer tomography (ct) scan, with bleeding at site. the sheath was removed. hemostasis was achieved with manual pressure and through administration of blood products. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage issue was most likely use related.
12322 | 1220648-2025-25500 | 21126937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21126937&pc=OZD | 12/24/2024 | 02/07/2025 |  | 01/10/2025 |  |  | 12/24/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Insufficient Information | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. upon attempting to pass the impella cp through iliacs through the aorta, it was not possible to advance the impella cp due to patient anatomy and the procedure was aborted. it was reported that the patient decompensated further and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy.
12323 | 1220648-2025-25501 | 21127169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21127169&pc=OZD | 12/26/2024 | 02/12/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529595 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a hematoma was noted at the groin site shortly after arrival to the intensive care unit from impella cp placement. manual pressure was held and heparin drip was stopped and on hold. the next day, the patient had slight oozing at the site without hematoma. it was further reported that on day six of impella support, the patient had a percutaneous coronary intervention procedure and was complaining of severe abdominal pain and became less responsive. the family decided to make the patient do-not-resuscitate and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is determined to be patient movement since the hematoma appeared upon transferring to the intensive care unit. b.2 revised as required intervention to prevent permanent impairment/damage was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25501. h.6 code 1888 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25501.
12324 | 1220648-2025-25502 | 21127181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21127181&pc=PYX | 12/26/2024 | 02/04/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  | 01/31/2026 |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025462446 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced oozing at the insertion site post peel away exchange with sheath during impella rp placement. oozing resolved with extra mattress suture added by wrapping around sheath. the patient¿s family decided on comfort care upon arrival to the intensive care unit and patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp with smart assist system with automated impella controller: section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since oozing was resolved by adding extra mattress suture.
12325 | 1220648-2025-25503 | 21128034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21128034&pc=OZD | 12/25/2024 | 09/30/2025 |  | 01/10/2025 |  |  | 12/25/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that a temporary pump stop occurred and multiple unsuccessful restart attempts were encountered. the patient was hemodynamically stable and continued on extracorporeal membrane oxygenation support. | The complainant also reported that there was a placement signal not reliable. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | The investigation for the pump stop has been completed. the product was not returned for investigation. the data log shows a sudden pump stop on the 21st day, with an observed motor current spike and an alarm indicating "impella stop motor current high." the pump was not able to restart and was explanted. the "temporary pump stop" failure mode was updated to "pump stop" after reviewing the investigation data logs. the cause of the pump stop issue was not determined since product was not returned. the pump usage exceeded the design life of 8 days for the cp pump which is classified under misuse of the device. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed optical signal issue. the product was not returned for investigation. the product was not returned for investigation. the data log shows the placement signal and motor current trend during the reported false positioning aorta alarm. no abnormalities were noted in the 5s optical parameters. all placement-related calculations were made from the optical sensor on the cp pump, so the reported failure mode of "placement signal issue" was changed to "optical signal issue." the cause of the optical signal issue was most likely patient condition related, based on data log review. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding.
12326 | 1220648-2025-25504 | 21128065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21128065&pc=OZD | 12/27/2024 | 02/25/2025 |  | 01/10/2025 |  |  | 12/27/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535202 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, leaking was noted at the purge system directly below the pressure reservoir. impella cp was replaced. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the data logs were reviewed and the purge leak was confirmed. roughly two hours into support, there is a four minute drop in purge pressure down to 100 mmhg, resulting in a purge pressure low alarm. the alarm resolved within seconds, but the purge pressure remained at that level for the rest of the case. the product was not returned for investigation. clinical details reported that the leak was coming from directly below the pressure reservoir, and that there was damage to the sidearm at the yellow luer. it is unknown if excessive force was applied. based on data log review and the details provided, the root cause of the purge leak was determined to most likely be a damaged sidearm. the exact location of the leak could not be determined without returned product.
12327 | 1220648-2025-25505 | 21129057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21129057&pc=OZD | 12/13/2024 | 01/31/2025 |  | 01/10/2025 |  |  | 12/14/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580263 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had escalated care from an impella cp to an impella 5.5 pump. the acute myocardial infarction/cardiogenic shock patient suffered ventricular tachycardia and atrial fibrillation while on impella 5.5 support. the patient was able to be converted to normal sinus rhythm and remained on support for a total of 12 days until care was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
12328 | 1220648-2025-25506 | 21129631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21129631&pc=OZD | 12/24/2024 | 01/10/2025 |  | 01/10/2025 |  |  | 12/24/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579611 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a study patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding noted by the oozing at the right femoral site. the pressure bandage was removed and applied again, and one stitch was placed. the outcome to this issue is unknown; however, it was noted that the patient survived the procedure. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12329 | 1220648-2025-25507 | 21130011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21130011&pc=OZD | 12/26/2024 | 02/06/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1703855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support when the automated impella controller (aic) gave a blank screen for 30 seconds. no alarms were noted and it returned back to the placement screen. the aic was exchanged. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) display issue has been completed. the data logs were reviewed which confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. upon functional inspection, the loss of optical data failure mode was not reproduced. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
12330 | 1220648-2025-25508 | 21130030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21130030&pc=OZD | 12/26/2024 | 02/27/2025 |  | 01/10/2025 |  |  | 12/26/2024 | 01/10/2025 |  |  |  |  | 01/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584664 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. following the initial insertion via the left axillary artery, the pump was unable to advance completely out of the graft. the pump was removed and an angiogram of the left subclavian showed a dissection. a balloon was placed to fix the dissection and the pump was successfully inserted via right axillary access. following the delivery, oozing was noted from the impella site. manual pressure and a pressure dressing were applied to manage the condition and one unit of packed red blood cells was given. support continued following the intervention. the patient survived the event. | [5 empty] | The investigation for the delivery issues, access site bleed, and vascular damage has been completed. pump was not returned for investigation. as per the clinical information provided, patient had extreme stenosis in the peripheral vessel narrowing the vessel which let to a dissection when the pump was passed through the graft. therefore, the root cause of the vascular damage issue was patient condition related to small/ diseased vessels. as per the clinical information provided, the delivery through lsa was aborted due to stenosis and dissection and rsa was used for pump delivery. therefore, the root cause of the delivery issue was patient condition related to diseased vessels. after the delivery into the rsa, there was oozing noted from the impella site where manual pressure and pressure dressings, and blood was given to manage. however, no other information was provided. the root cause of the access site bleeding - major cannot be determine. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12331 | 1220648-2025-25441 | 21111338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21111338&pc=OZD | 12/21/2024 | 03/25/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548820 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Complainant in japan reported the patient had an attempted impella cp implant and it was aborted. there was an issue during the insertion of the 16fr sheath. the insertion was aborted, and an impella 5.5 was placed. the patient eventually expired (refer to 24-40053, ci-39969). |  |  |  |  | Corrections to the initial mfr report: b1-selected adverse event. b2-selected death and added date of death. b5-mso review statement. h1-changed to death. h6-added impact code 1802. | The investigation for the reported delivery issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the delivery issue could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12332 | 1220648-2025-25442 | 21111501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21111501&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580236 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 insertion was unsuccessful. during placement of 5.5 via graft on the axillary artery, the physician noted that there was difficulty passing cannula fully through the axillary introducer and graft into the artery. a soft jaw clamp remained at the anastomosis and the physician noted this possibly caused the issue. while there was no obvious sign of damage, the physician felt that a kink occurred where the motor meets the catheter while trying to insert the pump. when the clamp was taken off, the pump was passed through the graft without issue, however, while trying to place the pump in the left ventricle (lv), buckling of the wire and impella cannula occurred. the impella was removed and a new 5.5 was inserted into the lv without issue. the patient survived the event. | [5 empty] | The investigation for the reported product damage (cannula kink) issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details noted that pump was unable to pass through the axillary graft with soft jaw clamp on the anastomosis. once the clamp was removed, the pump was able to be inserted without issue. the root cause of the delivery issues was most likely use related as a clamp was present on the graft when attempting to insert the pump. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12333 | 1220648-2025-25443 | 21111761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21111761&pc=OZD | 12/20/2024 | 02/06/2025 |  | 01/09/2025 |  |  | 12/20/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577774 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had limb ischemia. prior to leaving the catheterization lab, a femoral bypass technique was performed due to limited blood flow on the impella insertion side of the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
12334 | 1220648-2025-25444 | 21111899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21111899&pc=OZD | 12/22/2024 | 02/13/2025 |  | 01/09/2025 |  |  | 12/22/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570731 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had limb ischemia. a femorofemoral bypass was discussed but due to the patient's condition it was not implemented and support was continued. very high doses of inotropes/vasopressors were needed due to a severe hemodynamic instability. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12335 | 1220648-2025-25445 | 21112012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112012&pc=OZD | 12/22/2024 | 02/04/2025 |  | 01/09/2025 |  |  | 12/22/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565518 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient had severe right iliac artery tortuosity. it is unknown when the dissection occurred but ultimately led to pump removal and intra-aortic balloon pump placement. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage - peripheral vessel issue was patient condition related to diseased vessels.
12336 | 1220648-2025-25446 | 21112222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112222&pc=OZD | 12/23/2024 | 02/03/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  | 07/31/2025 |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371695 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device for optimization prior to coronary artery bypass graft(cabg) and approximately three days after start of support, day of cabg, the positioning became compromised. the surgeon was unable to advance the pump to an ideal position and subsequently, the pump left three centimeters below the valve. support with the impella 5.5 pump was uninterrupted and continues despite the suboptimal positioning. there was no report of harm to the patient as a result of this event. | [5 empty] | The investigation of positioning issues has been completed. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. d6b explant date added. | The impella device was not received from the customer as the device remains implanted and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
12337 | 1220648-2025-25447 | 21112371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112371&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559344 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was leaking from the red impella plug. it was not a continuous bleed but it was a drip. the staff used tape and pressure dressing to control the leaking blood. the pump was replaced due to the leak. the patient survived the event. | [5 empty] | The investigation for the product damage has been completed. pump was returned for investigation. data log review was not required for this case. a hole in the catheter was found near the 80cm mark, which allows blood to pass through the catheter and leak from the red handle. no other physical abnormalities were reported on returned product. the et tube bleeding was unrelated to the impella. the cause of the product damage issue was use related. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12338 | 1220648-2025-25448 | 21112396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112396&pc=OZD | 11/15/2024 | 09/22/2025 |  | 01/09/2025 |  |  | 12/20/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561702 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that the pump was placed but lost reliable placement signal. the pump remained on for support despite the lost placement signal. | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right axillary artery access site graft thrombus with a "probable" relationship to the impella 5.5 device used: ¿thrombus was removed from the graft. the vessel was allowed to back bleed in a retrograde fashion and then in an antegrade fashion and finally in a retrograde fashion again to ensure no thrombus was present.¿ additionally, ¿hemolysis (b)(6)2024 requiring transfusion most recently overnight (b)(6) - likely in the setting of hemolysis (b)(6), removed (b)(6)¿. the patient recovered with no sequelae. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were reviewed, and it was confirmed that the placement signal not reliable (psnr) alarm triggered on 11/3. at the time, the signal-to-noise ratio (snr) dropped to zero, contrast oscillated between +3200/-3200, lamp maxed out, and light/gain both dropped. the complication resolved as reported later that day and there were no further issues throughout the case. product was not returned for investigation. upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. corrections that were made from the initial manufacturer device report number. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of thrombus and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source; study was missing from manufacturer device report 1220648-2025-25448
12339 | 1220648-2025-25450 | 21112447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112447&pc=OZD | 12/28/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/28/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586070 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) in post cardiotomy cardiogenic shock was brought to the cardiovascular lab for an emergent impella cp device placement for mechanical circulatory support. immediately following successful placement of the pump, the patient began moving and positioning in the left ventricle was lost. the physician was unable to re-advance the pump back into the left ventricle, and the decision was made to explant and replace the device which was successfully completed with no resulting harm. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely patient movement. e4 should have been left blank on manufacturer device report 1220648-2025-25450. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12340 | 1220648-2025-25456 | 21112507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112507&pc=OZD | 12/25/2024 | 02/04/2025 |  | 01/09/2025 |  |  | 12/25/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and developed hemolysis. the patient presented to the catheterization lab for impella cp device placement for left ventricle unloading with extracorporeal membrane oxygenation (ecmo). the impella was successfully implanted and approximately three after start of support, the patient developed hemolysis. the impella was reassessed prior showing in good position. echocardiogram showed mid-left ventricle deep at 4.2cm, however, the device was not repositioned. the patient remained on support and approximately nine days later the patient was escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
12341 | 1220648-2025-25451 | 21112531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112531&pc=OZD | 12/23/2024 | 02/04/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582911 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant. however, the staff was unable to make the dive down the axillary-aorta junction. the staff attempted the buddy wire technique. however, it was unsuccessful. all attempts were unsuccessful. the impella 5.5 implant was aborted. the patient remains stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact fax number was inadvertently omitted from manufacturer device report 1220648-2025-25451. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12342 | 1220648-2025-25449 | 21112633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112633&pc=OZD | 12/23/2024 | 02/05/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572815 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted successfully and started in auto. percutaneous coronary intervention was then attempted when stent dislodged unemployed into the left main artery. multiple strategies used to attempt to snare/retrieve the stent. suddenly, the automated impella controller alarmed impella stopped, motor current high, and flows dropped. the patient became unstable and the impella was exchanged for a new cp. the patient survived the event. | [5 empty] | The investigation of pump stop has been completed. the data log showed the pump stop with high motor current alarm. the pump restarted and ran for another 10 minutes with low pump flow. no abnormalities related to improper position or optical signal issues were reported. after reviewing the data log, the reported pump stop failure mode was updated to temporary pump stop. the cause of the temporary pump stop issue was not determined since the product was not returned and sufficient clinical information was provided. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.4 should have been left blank on the initial manufacturer device report number: 1220648-2025-25449 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4868 was previously entered incorrectly on the initial manufacturer device report number: 1220648-2025-25449 and a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number: 1220648-2025-25449. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12343 | 1220648-2025-25453 | 21112871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112871&pc=OZD | 12/24/2024 | 02/10/2025 |  | 01/09/2025 |  |  | 12/24/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with "severe" aortic insufficiency decompensated prior to transcatheter aortic valve replacement (tavr) in the intensive care unit and was brought to the catheterization lab to place an impella cp device for stabilization as a bridge to tavr. after the impella cp was successfully implanted through the pre-existing corevalve, the peel away sheath was removed. prior to prepping for the patient's transfer back to the unit, it was noted in auto, flows were unable to achieve greater than 2.6l/min. fluoroscopy revealed a kink in the catheter. multiple maneuvers were completed to change the position; however, the kink remained. the physician used the re-access port to remove the kinked impella catheter to restart with a new impella cp implant, which was successfully completed. the patient was reported to be stable following the event. | [5 empty] | The investigation of product damage (cannula kink) has been completed. failure mode: product damage (cannula kink) changed to failure mode: low pump flow because the device successfully implanted through the pre-existing corevalve, but they were unable to achieve flows greater than 2.6l/min on auto. fluoroscopy revealed a kink in the catheter. the cause of the low pump flow was a cannula kink which was observed prior to prepping for the icu. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12344 | 1220648-2025-25452 | 21112934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21112934&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581675 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient was taken back to the catheterization lab to try to reposition. the impella was pulled across the valve and was unable to recross. the doctor decided to explant the impella. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issues most likely was use related due improper technique as the impella pulled out across the valve and it was unable to recross it. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12345 | 1220648-2025-25454 | 21113139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113139&pc=OZD | 12/23/2024 | 02/10/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578404 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an emergent impella 5.5 implant due to having a pericardial effusion. the impella would not pass the bend into the aorta. multiple wires were attempted. the patient began decompensating and the effusion was worsening. the patient arrested most likely from the tamponade. the decision was made to stop cardiopulmonary resuscitation and the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump was not able to pass bend into aorta due to patient's anatomy. h6 code 2993 was reported incorrectly on manufacturer device report 1220648-2025-25454. new code was added to medical device problem code.
12346 | 1220648-2025-25455 | 21113209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113209&pc=OZD | 12/24/2024 | 02/12/2025 |  | 01/09/2025 |  |  | 12/24/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371823 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that prior to insertion, the zeroed parameters of an impella 5.5 pump fell outside of the normal limits. the pump was replaced in an attempt to resolve the issue, but the issue persisted. the automated impella controller was subsequently replaced and the issue was said to have resolved utilizing the new pump. there was no patient harm. | [5 empty] | The investigation of optical signal issue has been completed. according to the data and device analysis the root cause for the low snr and calibration issue was most likely due to improper connection of the white plug catheter to the front panel optical connector or the defective front panel optical connector. d9 date device returned to manufacturer added. h5 was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25455. h8 was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25455. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12347 | 1220648-2025-25457 | 21113434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113434&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583496 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump. following insertion, the pump was unable to traverse the patient's anatomy with notable difficulty in directing the cannula toward the apex when entering the aorta. after multiple failed attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because of the difficult patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12348 | 1220648-2025-25458 | 21113445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113445&pc=OZD | 12/25/2024 | 02/11/2025 |  | 01/09/2025 |  |  | 12/25/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524204 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support and some hours after the pump was implanted, the patient experienced continued hemolysis and ectopy. consequently, the decision was made to explant the impella cp device. the patient's status was monitored over night with a plan to implant an impella 5.5 device if needed. the patient was reported to be stable following the event. | [5 empty] | The investigation of the reported issues has been completed. device history record review confirmed that the pump passed all post sterile inspection checks. the impella device was not received for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12349 | 1220648-2025-25460 | 21113464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113464&pc=OZD | 12/24/2024 | 01/28/2025 |  | 01/09/2025 |  |  | 12/24/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563665 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted to support a patient with acute myocardial infarction / cardiogenic shock. the patient was bleeding around the sheath of the impella. the bleeding was not able to be controlled, and the cp was explanted. additionally, care was withdrawn and patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25460. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25460 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12350 | 1220648-2025-25461 | 21113853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113853&pc=OZD | 12/10/2024 | 02/24/2025 |  | 01/09/2025 |  |  | 12/24/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564603 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp smart assist support for 732.63 hours before expiring. the cause of death was hemorrhagic shock. the location of bleeding was unknown, and the cause of the hemorrhagic shock was unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. as per the clinical information provided, the bleeding location is unknown and there is no other information available as to how the bleeding occurred. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided.
12351 | 1220648-2025-25459 | 21113941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21113941&pc=OZD | 12/17/2024 | 02/07/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537558 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and experience a cerebral hemorrhage with a demised outcome. the patient was on both extracorporeal membrane oxygenation (ecmo) and impella mcs, and during support it was noted the patient had experienced a large cerebral hemorrhage that caused a deviation. it was stated the patient was mainly managed with ecmo, and the activated clotting time was approximately 220. the patient's heart started to recover; however, the patient expired approximately three days after start of support from the cerebral hemorrhage. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post inspection checks. as insufficient clinical information was provided, the root cause of the stroke/cva could not be determined.
12352 | 1220648-2025-25462 | 21114158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21114158&pc=OZD | 12/24/2024 | 02/10/2025 |  | 01/09/2025 |  |  | 12/24/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578509 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that prior to insertion, the zeroed parameters of an impella 5.5 pump fell outside of the normal limits. the pump was replaced in an attempt to resolve the issue, but the issue persisted. the auotmated impella controller was subsequently replaced and the issue was said to have resolved utilizing the new pump. there was no patient harm. | [5 empty] | The pump was not returned for investigation. according to the clinical team, the placement signal was out of normal limits during two attempts on console ic6495. a new pump also exhibited the same issue on console ic6495, which was resolved after switching to a different console. the cause of the optical signal failure was most likely air gap from between the aic and the patient cable plug or within the middle of the red handle ferrule, based aic investigation results, data log review and clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12353 | 1220648-2025-25463 | 21114471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21114471&pc=OZD | 12/20/2024 | 01/30/2025 |  | 01/09/2025 |  |  | 12/20/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573374 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intracranial Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a brain bleed. carotid artery was used for impella access, so it is possible the access site was connected to the brain bleed, but it is difficult to know for sure. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
12354 | 1220648-2025-25464 | 21117792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21117792&pc=OZD | 12/23/2024 | 01/28/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient developed a hematoma, oozing was noted from site. the physician attempted to express hematoma and hold pressure, decided to take back to cath lab and explant. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was most likely patient movement since the hematoma appeared after the arrival of patient into intensive care unit from the catheterization lab. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12355 | 1220648-2025-25465 | 21120728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21120728&pc=OZD | 12/16/2024 | 02/14/2025 |  | 01/09/2025 |  |  | 12/25/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583668 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient suffered a cerebral hemorrhage. the patient ultimately expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
12356 | 1220648-2025-25466 | 21120729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21120729&pc=OZD | 12/23/2024 | 01/28/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578360 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following insertion of an impella cp pump for mechanical circulatory support, a hematoma began to develop at the groin site. after being unable to achieve hemostasis via manual pressure, the decision was made to remove the pump. the patient was then sent to the vascular operating room to repair the vessel. two units of blood were given to the patient to counteract the condition. the physician was unsure as to the cause of the bleed from the left femoral artery. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25466 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 health effect - impact code 4624 was inadvertently omitted from manufacturer device report number 1220648-2025-25466 and was added. h.6 health effect - clinical code 1888 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25466. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12357 | 1220648-2025-25467 | 21120740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21120740&pc=OZD | 12/23/2024 | 02/18/2025 |  | 01/09/2025 |  |  | 12/23/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571390 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that roughly 19 days into support, an impella cp pump stopped unexpectedly. attempts to restart the device were unsuccessful and the decision was made to explant the device. support with extracorporeal membrane oxygenation remained ongoing despite the absence of an impella device. there was no reported patient harm. | [5 empty] | The investigation of pump stop has been completed. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. the pump was used beyond its design life of 8 days, which is classified as misuse of the pump beyond the indicated usage guidelines per the instructions for use. e4 should have been left blank on manufacturer device report 1220648-2025-25467. g2 report foreign; study was missing from manufacturer device report 1220648-2025-25467. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12358 | 1220648-2025-25468 | 21120747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21120747&pc=OZD | 12/16/2024 | 02/27/2025 |  | 01/09/2025 |  |  | 12/26/2024 | 01/09/2025 |  |  |  |  | 01/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504546 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the pump stopped and the device was removed. the patient's heart function had recovered. the patient survived the event. | [5 empty] | The investigation of pump stop has been completed. the pump was returned for evaluation. upon visual inspection, a large purge gap was observed on the returned pump. no data logs were returned for evaluation. clinical details noted that a pump stop occurred on day 12 of support. at time of pump stop, a sudden rise in motor waveform was observed. no purge related issues were noted at time of pump stop. the root cause of the pump stop was due to bearing wear beyond indicated design life. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25468. d.4 revied unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25468. e.4 should have been left blank on the initial manufacturer device report number 11220648-2025-25468 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12359 | 1220648-2025-25408 | 21101317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21101317&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  | 12/31/2025 |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450118 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and during the implant experienced access site bleeding followed by a hemorrhagic stroke during support. prior to the impella implant, the patient underwent a cardiac catheterization that revealed normal coronary disease. post catheterization chest x-ray revealed worsening pulmonary edema. three days later the patient underwent a modified bentall procedure with 23cm aortic valve replacement with intra-aortic balloon pump (iabp) support. postoperative complication encountered and venoarterial extracorporeal membrane oxygenation (va ecmo) was placed. ejection fraction was less than 10% with a non-ischemic cardiomyopathy and class iii-iv acute on chronic systolic heart failure. following, the patient was escalated to an impella 5.5 device to right axillary artery and removal of the iabp. during the impella 5.5 implantation through the 8mm graft, the peel away sheath backed out of graft while advancing the pump. excessive bleeding was observed. only one impella graft lock was in use and challenging anatomy requiring excessive manipulation to advance pump. blood products were given (types with number of units were not provided). the graft was clamped down with a fogarty until the physician was able to reinsert 22fr. peel away sheath back into place with two graft locks and big silks. the patient was transported to the unit intubated and sedated. approximately seven days after start of support, hemorrhagic stroke was confirmed. anticoagulation hold continued. continuous renal replacement therapy was in use; urine minimal dark brown/amber. the plan over next three days was to watch for brain swelling. long- term plans versus palliative care were discussed with family. following the cerebral vascular accident, it was noted the patient had generalized weakness, moved bilateral upper extremity to commands and spontaneously responded appropriately. the patient remains on impella mcs, and latest point of care update (day 15 of support), tracheostomy and percutaneous endoscopic gastrostomy placements planned; patient remains intubated, not sedated and ventilator dependent. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the reported issue has been completed. no device was received for evaluation. the device history review confirmed that the pump passed all post sterile inspection checks. the root cause of the access site bleeding issue was most likely patient condition related, as it was noted in the clinical description that the patient had challenging anatomy requiring excessive manipulation to advance pump.
12360 | 1220648-2025-25409 | 21101705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21101705&pc=OZD | 12/18/2024 | 02/04/2025 |  | 01/08/2025 |  |  | 12/18/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569114 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and subsequently, developed a clot. the patient was initially supported with intra-aortic balloon pump (iabp). the impella 5.5 was inserted to the right axillary artery. the iabp was turned off and removed over wire in exchange for an 8 frnch sheath to right femoral artery (rfa). the impella 5.5 pump position was verified on transesophageal echocardiogram and the site was secured. the patient was transferred to the unit on impella mcs. update post arrival to the unit noted access site was intact and jackson-pratt (jp) drain was stable with plans to remove the 8 french sheath from the rfa when activated clotting time decreases. approximately 14 days later, the unit update noted there was swelling at the insertion site since the jp drain was removed due to a clot "the other day." the actual day the clot and the relationship between the clot and the impella are unknown. five days later, the patient was noted to be doing well, the swelling had gone down, and the patient received a new jp drain. the patient continued on impella mcs until successfully weaned and explanted without issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
12361 | 1220648-2025-25410 | 21101841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21101841&pc=OZD | 12/13/2024 | 04/02/2025 |  | 01/08/2025 |  |  | 12/13/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574829 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the patient had gastrointestinal (gi) bleed. heparin was on hold due to gi bleed. bleeding was due to an open chest, not impella related. the patient was given two units of red blood cells (rbcs) the patient endured arrhythmias. supraventricular tachycardia (svt) causing suction events and very low flows. cardioversion was planned. the patient was given one unit of rbcs. the patient remains on impella support. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death. thus, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | B.2 added death and date of death due to additional information received. b.5 additional information was received from the clinical site regarding the patient outcome of death. d.6b explant date was added. h.1 type of reportable event was revised due to new additional information received. h.6 health effect - impact code was added due to additional information received. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25410. h.6 code 1721 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25410.
12362 | 1220648-2025-25411 | 21101935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21101935&pc=OZD | 12/20/2024 | 02/06/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  | 04/30/2025 |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336760 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support developed bleeding at the graft anastomosis. additional sutures were applied and two units of blood were administered. bleeding was controlled. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, additional sutures applied and bleeding controlled as per the clinical description provided.
12363 | 1220648-2025-25412 | 21102020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102020&pc=OZD | 12/20/2024 | 01/23/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592693 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered positioning issues as the pump was pulled into the aorta while attempting to reposition it. as a result, a new pump and pump sheath were used instead. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12364 | 1220648-2025-25413 | 21102163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102163&pc=OZD | 12/23/2024 | 02/25/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582561 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and upon start up, the impella cp pump experienced high motor current followed by a pump stop. both the impella cp pump and automated impella controller were replaced to resolve the issue and support the patient through the high-risk pci. it was reported there was no harm to the patient as a result of this event. | [5 empty] | The investigation of pump did not start has been completed. the data logs were investigated and the complication was confirmed. the instant the user attempts to start the pump the first time, the motor current spiked to 1487 ma and a high motor current pump stop occurs. the pump was able to restart a few minutes later, however, the pump flows were immediately saturated. after five minutes of running, the p-level was reduced, and another motor current spike occurred, followed by low pump flows. these signatures indicate that during the first start up attempt, biomaterial was most likely ingested into the purge gap prior to or at the same instant, causing the subsequent saturated flows. then at the second motor current spike, it may have got kicked out into the cannula, causing the low pump flows. the product was not returned for investigation. based on data log review, the pump did not start was determined to most likely be due to biomaterial. e.2 revised address line 2 as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25413. h.6 code 1503 was reported incorrectly on teh initial manufacturer device report number 1220648-2025-25413. a new code was added to medical device problem codes to reflect that the pump did not start. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for the automated impella controller serial number (b)(6) with date of event 23-dec-2024 and identifier ending in (b)(6).
12365 | 1220648-2025-25414 | 21102379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102379&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558681 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support was bleeding from multiple sites including impella access site. two units of blood products were administered. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition since the patient was bleeding from multiple sites. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25414 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12366 | 1220648-2025-25415 | 21102388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102388&pc=OZD | 12/20/2024 | 03/21/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525605 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Positioning Problem | [5 empty] | Cardiac Arrest | Fever | Hematoma | Hemorrhage/Blood Loss/Bleeding | Sepsis | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Additionally, the patient had a fever that was suspected to be from sepsis. antibiotics were started. the cause of the sepsis is unknown. | A complainant reported the patient was on impella cp support. while on support, the patient had cardiac arrest after impella out from left ventricle, extracorporeal membrane oxygenation was required. after 24 days running the impella cp blocked, physicians couldn't restart it after three attempts. the impella was explanted and the procedural outcome was the patient survived surgery. | It was reported that there was bleeding/hematoma that was managed by an unplanned intervention. |  |  |  | As noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.5 revised to include an additional information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25415. h.6 revised to include additional health effect - clinical codes, health effect - impact code, and medical device problem code that were inadvertently omitted from manufacturer device report number 1220648-2025-25415. h.10 instructions for use for the additional adverse events that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25415 were added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. the pump was not returned for investigation. data logs were not returned as well. without the data logs and pump, the failure could not be further investigated. therefore, the root cause of the pump stop issue was not determined since the product was not returned and data logs were not available as well. the pump was used for 24 days which is over the design life of the impella cp pump. therefore, this is misuse of the pump. clinical details are very limited about the access site bleeding issue. bleeding/hematoma was managed by unplanned intervention. therefore, the root cause of the access site bleeding-major issue was not determined since the product was not returned and insufficient clinical information was provided. the root cause of the cardiac arrest was unable to be determined due to insufficient clinical details. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.5 revised to include an additional adverse event that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25415. e.1 revised address line 1 and 2 as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25415. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25415 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 revised to include an additional health effect - clinical code and health effect - impact code that were inadvertently omitted from manufacturer device report number 1220648-2025-25415. h.10 instructions for use for the additional adverse event that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25415 were added.
12367 | 1220648-2025-25416 | 21102424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102424&pc=OZD | 12/20/2024 | 02/04/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540154 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support right distal pulse was no longer dopplerable. a right lower extremity fasciotomy was done. post procedure, distal pulses were confirmed by physician. additionally, the patient had a gastrointestinal bleed and heparin was held. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and gi bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and gi bleed could not be determined as the device was not returned nor sufficient clinical details.
12368 | 1220648-2025-25417 | 21102898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102898&pc=OZD |  | 02/13/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 63 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Disability | Life Threatening | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication was received titled clinical and pathological characteristics of immune checkpoint inhibitor associated fulminant myocarditis," and indicated a patient (unknown race and ethnicity) undergoing treatment for lung cancer was admitted due to fever and chest pain. echocardiography revealed diffuse severe left ventricle hypokinesis with left ventricle ejection fraction of 10%. coronary angiography revealed no steno-occlusive lesions. due to hemodynamic instability, venoarterial extracorporeal membrane oxygenation (va-ecmo) and intra-aortic balloon pump (iabp) were introduced. the patient was transferred to an outside hospital on the third day and an impella cp device was inserted to replace the iabp. the next day, the patient's pupils were observed, and a head computed tomography (ct) scan confirmed "extensive" cerebral hemorrhage in the left frontoparietal lobes associated with ventricular perforation and uncal herniation. the patient's coma was deemed irreversible, so surgery was not an option. although cardiac function gradually improved, the patient expired of intracerebral hemorrhage with uncinate herniation on the 13th day of support. | [5 empty] | At the time of this mdr writing some mandatory information is unknown: date of onset and date of death are unknown; respective fields left blank. device-specific information other than the brand name is unavailable, and therefore respective sections (d4, d6 and h4) are reflected as "unknown" or field left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ryo izumi, et al. (2024). clinical and pathological characteristics of immune checkpoint inhibitor associated fulminant myocarditis. cardio-oncology (2024) 10:82. | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12369 | 1220648-2025-25418 | 21102978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21102978&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574348 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Septic Shock | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, oozing at the right axillary site was noted around the site. heparin was held for transport. the patient continued to bleed from his mouth. additionally, mixed cardiogenic shock -septic shock was being treated, requiring increase in norepinephrine to .3, antibiotics started. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding and infection/sepsis has been completed. the impella device was not returned for evaluation. the root cause of theissues was not determined since the product was not returned and insufficient clinical details were provided.
12370 | 1220648-2025-25419 | 21103122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103122&pc=OZD | 12/22/2024 | 01/31/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient¿s impella access leg was noted to be cooler in comparison to the other leg but a dopplerable pulse was present. after removal, the patient¿s leg remained cooler than the opposite leg. the right foot was also noted to be markedly paler. pulses present by doppler in bilateral lower extremities (ble) after explant. vascular surgery was consulted but at this time there is no plan for vascular intervention. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: h6 impact code 4627 changed to 4641.
12371 | 1220648-2025-25420 | 21103138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103138&pc=OZD | 12/20/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560994 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was implanted to support a patient with acute myocardial infarction/cardiogenic shock. during support, the patient experienced oozing from the impella access site. the perclose was tightened, suture was placed at the site, dressing was changed, angle was matched, and manual pressure was held. two units of packed red blood cells were administered in the last 24 hours. it was also reported that the patient experienced loss of right peripheral leg pulses. the doctor came to the bedside and reestablished passive flows using a conduit from the contralateral right femoral artery to the left superficial femoral artery. pedal pulses are now palpable. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and the limb ischemia has been completed. the device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because sutures were tightened and additionally two sutures were placed to achieve hemostasis. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
12372 | 1220648-2025-25421 | 21103492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103492&pc=OZD | 12/21/2024 | 02/05/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569838 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the staff noticed a large hematoma forming around the impella site. despite applying manual pressure, the hematoma continued to grow. critical care ordered to start massive transfusion protocol. norepinephrine had to be started, and the console began alarming suction. the doctor came to the bedside and the plan was to return to the catheterization lab to close the right groin and get access in the left groin and place a new device. after the plan was formulated, the family came to the bedside and decided to withdraw care and the patient expired. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, patient was rather disoriented and began thrashing about in bed and attempting multiple times to sit up as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12373 | 1220648-2025-25422 | 21103621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103621&pc=OZD | 12/22/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had some oozing at the impella access site post agitation. heparin was paused and one unit of packed red blood cells was given. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement. impella site oozing over night due to agitation as per the clinical description provided.
12374 | 1220648-2025-25423 | 21103622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103622&pc=OZD | 12/21/2024 | 02/10/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569159 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon arrival to the intensive care unit from impella cp placement, the patient had oozing from the right femoral artery. a quick clot was applied, dressing was changed, and bival was to be started at night. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
12375 | 1220648-2025-25424 | 21103847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103847&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had right upper extremity swelling and a nonocclusive deep vein thrombosis noted in subclavian vein. the next day, the impella was removed. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. to determine the true root cause of the thrombus, the clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
12376 | 1220648-2025-25425 | 21103885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103885&pc=OZD | 12/23/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560866 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella 5.5. while accessing left ventricle (lv), the doctor experienced some difficulty traversing wire and catheter toward atrioventricular (av). continuously directed toward descending aorta once past innominate. finally, was able to get wire into lv but also experienced the same difficulty with cannula. attempted with multiple .018 wires, buddy wires & constant manipulation attempts. unable to torque toward av and the doctor felt it would be easier to place impella cp so 5.5 was aborted. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because 5.5 insertion was aborted due to complex vascular anatomy of the patient. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12377 | 1220648-2025-25426 | 21103943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21103943&pc=OZD | 12/22/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551602 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient developed hemolysis evident by pink/red tinged foley/urine. the impella cp was explanted and an impella 5.5 was used. hemolysis cleared as a result. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the cause of the hemolysis issue cannot be determined since the complaint device not returned for analyzing and insufficient clinical data provided.
12378 | 1220648-2025-25427 | 21104398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21104398&pc=OZD | 12/23/2024 | 02/05/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  | 06/30/2025 |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the staff was unable to advance an impella cp pump through the patient's tortuous vasculature during attempted delivery for mechanical circulatory support. after multiple failed attempts, the implant was aborted. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella could not pass through the vasculature as per the clinical description provided. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12379 | 1220648-2025-25428 | 21104904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21104904&pc=OZD | 12/20/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571341 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, bleeding was noted to impella site. manual pressure was held. fluid and 2 units of packed red blood cells were administered for volume. the patient was noted to be in a-fib, synchronized cardioversion administered twice. impella cp was repositioned under fluoro. the patient was moved to the intensive care unit and femstop was placed. distal pulses were not dopplerable. the physician was aware but deemed the patient to be dependent on impella support and distal pulses were monitored. care was withdrawn and the patient expired. reportable due to patient harm and potential relationship between the cause of death and the impella. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issues has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post sterile inspection checks. to determine the true root cause of the reported issues, sufficient clinical information would need to be assessed in conjunction with evidence from the returned device. as all necessary information was not provided, the root cause of the reported issues was unable to be determined.
12380 | 1220648-2025-25429 | 21105191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21105191&pc=OZD | 12/20/2024 | 02/07/2025 |  | 01/08/2025 |  |  | 12/20/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561194 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was prepped and placed in the patient but unable to get it to cross into the lv due to the patient¿s anatomy. the 5.5 was removed and replaced with a cp. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2025-25429 in accordance with updated procedures. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis". ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12381 | 1220648-2025-25430 | 21105351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21105351&pc=OZD | 12/21/2024 | 02/10/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590688 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and there was impella defective alarms. three different aics were tried and all three gave the same alarm. the pump was replaced. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the pump not detected issue has been completed. impella cp returned for evaluation. upon visual inspection, dried blood was present on the impeller blade. no biomaterial was present on the returned pump. no blood present in the red handle. pump was purged and run at p-9 for approximately 6 minutes and no pump stop or alarms occurred. data logs were reviewed and lt log from field shows that no "impella defective" alarms were triggered. clinical details noted that pump was experiencing "impella defective" alarms while pump was running and troubleshooting with multiple aics was attempted. the root cause of the pump not detected cannot be definitively determined. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email.
12382 | 1220648-2025-25431 | 21105530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21105530&pc=OZD | 12/21/2024 | 01/28/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  | 06/30/2025 |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379199 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support the patient had hemolysis. the impella cp was removed and replaced with an intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25431 was submitted in accordance with updated procedures. g.1 revised manufacturing site name and address section as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25431. h.6 health effect - impact code 4629 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25431. new codes were added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2025-25431 in accordance with updated procedures.
12383 | 1220648-2025-25432 | 21105821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21105821&pc=OZD | 12/20/2024 | 02/19/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575957 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp moved towards the aortic side and was unable to cross back into the left ventricle. the cp was exchange for a new one. there was no harm to the patient. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12384 | 1220648-2025-25433 | 21107361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107361&pc=OZD | 12/21/2024 | 01/30/2025 |  | 01/08/2025 |  |  | 12/21/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, distal pulses on the impella leg were lost and the patient developed a cold limb. the decision was made to wean down/off impella per doctor¿s order and place an intra-aortic balloon pump (iabp). the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25433.
12385 | 1220648-2025-25434 | 21107364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107364&pc=PYX | 12/22/2024 | 02/03/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025528998 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella rp flex for mechanical circulatory support. following insertion, it was documented that the physician observed a kink in the catheter. the pump was subsequently removed and replaced to resolve the issue. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since the pump kinked and the second was implanted successfully. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: rp flex with smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12386 | 1220648-2025-25435 | 21107370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107370&pc=OZD | 12/22/2024 | 02/06/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  | 10/31/2025 |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445187 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted to support a patient with acute myocardial infarction / cardiogenic shock. the patient on support had a change in urine color, output, and labs were hemolyzed. position was confirmed by trans-thoracic echocardiography. the impella was running on p9 due to the patient being on cardiogenic shock (cgs). it was further reported that the patient had episodes of sustained runs of ventricular tachycardia. the patient was cardioverted. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined d6b explant date was inadvertently omitted from manufacturer device report 1220648-2025-25435. e4 should have been left blank on manufacturer device report 1220648-2025-25435. g1 reporting contact fax number was inadvertently omitted from manufacturer device report 1220648-2025-25435.
12387 | 1220648-2025-25436 | 21107375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107375&pc=OZD | 12/22/2024 | 02/10/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530568 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella cp. upon advancing impella, physician was unable to get pump through the mid descending aorta. after multiple minutes of trying, the physician aborted and the pump was removed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. corrections to the initial mfr report: d4 updated model number. g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12388 | 1220648-2025-25437 | 21107379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107379&pc=OZD | 12/22/2024 | 02/06/2025 |  | 01/08/2025 |  |  | 12/22/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573497 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp was implanted there was oozing at the impella access site. manual pressure was applied. forward tension was pulled and blue suture hub pushed more into arteriotomy, and bleeding stopped. additionally, on day four of impella cp support the patient was escalated to impella 5.5 due to thrombus. the patient had a large descending aortic thrombus that required vascular surgery to remove prior to explanting cp. clot formed around the catheter and not inside the device itself. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25437. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25437 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2025-25437 was submitted in accordance with updated procedures. h.6 code 4642 was added since the initial manufacturer device report number 1220648-2025-25437 was submitted as the patient was escalated to an impella 5.5. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12389 | 1220648-2025-25438 | 21107392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107392&pc=OZD | 12/23/2024 | 02/06/2025 |  | 01/08/2025 |  |  | 12/23/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560809 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during removal of an implanted impella cp pump, minimal flow was seen down the left femoral artery. an angiogram was performed and a thrombus was seen from the left iliac to the left superficial femoral artery. a balloon angioplasty and penumbra lightening were used for thrombus extraction and flow returned to the area. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25438 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12390 | 1220648-2025-25439 | 21107411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107411&pc=OZD | 12/24/2024 | 02/04/2025 |  | 01/08/2025 |  |  | 12/24/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590428 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and the following day after start of support the patient developed hemolysis coinciding with low flows on support level p-9. the patient's urine was dark red, and the patient was hypertensive. it was noted the position was not check at this time under transesophageal echocardiogram. however, the noted plan was to wean the impella to support level p-4 overnight, start bicarbonate in the purge and systemic heparin drip, and complete an echocardiogram the next morning to reevaluate the need for impella mcs. the patient was reported to be stable following the event. the device was eventually explanted that next morning; the patient's outcome post explant was noted as unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and low pump flow has been completed. the product and data were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided. the failure mode low flow was removed as flow 3.4 l/min on p-9 was in normal range 3.3 - 3.7 l/min. e4 should have been left blank on manufacturer device report 1220648-2025-25439.
12391 | 1220648-2025-25440 | 21107414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21107414&pc=OZD | 12/22/2024 | 02/04/2025 |  | 01/08/2025 |  |  | 12/26/2024 | 01/08/2025 |  |  |  |  | 01/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the automated impella controller (aic) had a yellow error alarm. there was no change in the auxiliary flow rate, and the alarm resolved naturally after a few minutes. no similar alarms were observed until the device was removed. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console log confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. real time logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)¿ was determined to be a firmware issue. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25440. d.2 revised common device name since the initial manufacturer device report number 1220648-2025-25440 was submitted in accordance with updated procedures. d4 revised serial number and model number as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25440. e.1 revised name prefix/title, first name, middle name, last name and address line 1 as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25440. e.3 revised occupation as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25440. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25440 was submitted in accordance with updated procedures. h.6 revised component code as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25440. | The impella device was received from the customer on (b)(6) 2025, therefore, an investigation of the device was underway. should any new information be received, a supplemental mdr will be filed.
12392 | 1220648-2025-25356 | 21091245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091245&pc=OZD | 12/13/2024 | 02/04/2025 |  | 01/07/2025 |  |  | 12/13/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577243 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Obstruction/Occlusion | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. upon explant, it was noted that the superficial femoral artery was occluded. the decision was made to bring the patient to the vascular lab for open repair. the procedural outcome was the patient survived surgery. | It was further reported that biomaterial was found on the device. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b.5 added information that was inadvertently omitted from the initial manufacturer device report number. h.6 code added due to new information that was added to b.5 h.10 additional instructions for use were added due to information that was added to b.5.
12393 | 1220648-2025-25357 | 21091303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091303&pc=OZD | 12/12/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/18/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516456 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient experienced a loss of vision in the right eye. a computed tomography (ct) scan showed high grade right internal carotid artery stenosis but no ischemia bleeding or clots. symptoms resolved by return from ct. the patient received a transplant and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. the returned product was visually inspected and biomaterial was found on the returned product around the catheter/motor housing. the source of the stroke clot is unknown. as this clinical information was not provided, the true root cause of the stroke/cerebrovascular accident could not be determined. d.9 device returned date/information was added.
12394 | 1220648-2025-25358 | 21091448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091448&pc=OZD | 12/16/2024 | 06/27/2025 |  | 01/07/2025 |  |  | 12/16/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589733 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced ventricular tachycardia during the implant procedure with an access site hematoma following. the right axillary surgical pocket was created; the vessel was noted to be small but adequately sized for impella insertion. non-beveled graft anastomosed to artery without issue, and the pump was inserted, successfully placed following arm manipulation to traverse severe tortuosity at anastomosis. persistent ventricular tachycardia developed during pump placement requiring shock one time at startup. an intra-aortic balloon pump was placed on standby and was explanted once impella mcs was started. the patient was extubated and transported to unit in stable condition. two days following start of support a hematoma developed with pain at the site. heparin was discontinued and site monitored. the following day noted there were no acute events overnight, impella at support level p-6 with appropriate flows and the hematoma showed improvement. heparin was restarted. forward two days a hematoma was noted and again heparin was discontinued. next day, the hematoma was still present with oozing. heparin remained discontinued. now day seven, the patient remained stable on minimal inotropic support with impella support level at p-4. however, the purge side arm was noted to be leaking. purge flows and pressures were within defined limits. purge pressure and motor current were monitored. heparin still remained discontinued due access site hematoma. however, the decision was made the following day to prophylactically remove impella mcs for purge sidearm leaks and risk of pump stop. the patient was noted to be stable following the explant. however, the outcome as it relates to the hematoma is unknown. | [5 empty] | The investigation of access site bleeding, vf/vt/af/at, and purge leak - pump has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of the access site bleeding could not be definitively determined as limited clinical details were provided and no product was returned for evaluation. purge leak - pump: the root cause of the purge leak - pump was most likely due to damaged sidearm based on clinical details, however, the exact location of the leak could not be definitively determined as no product was returned for evaluation. vf/vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
12395 | 1220648-2025-25359 | 21091513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091513&pc=OZD | 12/19/2024 | 10/03/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575222 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | Failure to Sense | Device Sensing Problem |  |  |  | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported that during coronary artery bypass graft with impella cp support, a high purge pressure alarm occurred on the impella cp, followed by low flows and placement signal not reliable alarms. the impella cp was replaced urgently. a blood clot was found in the impella outflow cage after removal. the patient is stable and survived the event. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | Additional information the patient expired. |  |  |  | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Additional information provided in h3, h6 and h10. the investigation into the reported issues has been completed. device history record review confirmed that the pump passed all post-sterile inspection checks. high purge pressure: the product was returned and evaluated. there was biomaterial found on the underside of the impeller and wrapped around the purge gap. the reported issue was reproduced and resolved when the catheter was cut just above the motor housing. there was no evidence of blood ingress in the purge line. data logs show a motor current spike at ~1500 with an "impella stopped - motor current high" alarm shortly after the pump was turned back on from surgical mode. at this point, the purge pressure began to rise quickly from 400 to >1000 mmhg. the pump was able to be restarted according to the logs. the root cause of the high purge pressure was most likely biomaterial ingestion as biomaterial was found wrapped around the purge gap of the returned pump and logs show a motor current spike before a quick rise in purge pressure. the field also observed a clot in the outlet upon pump removal. saturated flow. as stated above, there was biomaterial found on the underside of the impeller and wrapped around the purge gap. hpp did reproduce in the lab and the pump was not run. there was no evidence of blood ingress. as stated in the clinical details, the average pump flow was 2.0 l/min at p-1. log review showed saturated flow corresponding with the motor current spike and rise in purge pressure. the root cause of the saturated flow was most likely biomaterial as biomaterial was found in the outlet and the data logs show evidence of biomaterial in flow, purge pressure, and motor current. optical signal issue. no abnormalities were found with the 5s optical parameters of the returned pump. the pump passed laser continuity and no kinks or bond issues were found. the clinical details stated the patient had an on-pump arrest before going into the cabg procedure. log review showed a decline in placement signal leading up to putting the pump in surgical mode. psnr alarms occurred at this point but quickly resolved. there were also mulitple impella position unknown alarms. the placement signal recovered by the end of the case data logs. the root cause of the optical signal issue was most likely patient condition related as the alarms were not reproduced in the lab, the returned pump had no abnormalities, and the clinical details suggest deteriorating patient condition. temporary pump stop. the failure "temporary pump stop" was added due to the discovery of an "impella stopped - motor current high" alarm with a pump stop in the data logs. as mentioned above, biomaterial was found on the underside of the impeller and around the purge gap on the returned pump. no other abnormalities were found. log review showed the motor current spike at the time of the alarm and pump stop. the pump was able to be restarted. the root cause of the temporary pump stop was most likely biomaterial as biomaterial was found in the outlet of the returned pump and there was a motor current spike ~1500 ma at the time of the alarm. | B.5 additional information was received, and abiomed's medical safety officer review was completed. h.6 code 3331 was reported incorrectly on previous manufacturer device report number 1220648-2025-25359. a new code has been added. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no abnormalities were found with the 5s optical parameters of the returned pump. the pump passed laser continuity and no kinks or bond issues were found. log review showed a decline in placement signal leading up to putting the pump in surgical mode. the placement signal recovered by the end of the case data logs. the root cause of the optical signal issue was most likely patient condition related as the alarms were not reproduced in the lab, the returned pump had no abnormalities, and the clinical details suggest deteriorating patient condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b2 death and death date updated. b5 updated to include death description. h1 type of reportable event. h6 updated to include optical signal coding.
12396 | 1220648-2025-25360 | 21091594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091594&pc=OZD | 12/18/2024 | 01/31/2025 |  | 01/07/2025 |  |  | 12/18/2024 | 01/07/2025 |  |  | 01/31/2026 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461233 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion, the pump was unable to advance past the innominate. after multiple failed attempts, the implant was aborted and an impella cp pump was placed for patient support. there was no resulting patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella backloaded and advanced easily through graft and unable to advance past innominate as per the clinical description provided. corrections to the initial mfr report: d4 updated model number. g1 mdr reporting contact email. h5 should have been yes. h6 impact code changed to 4629. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12397 | 1220648-2025-25361 | 21091596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091596&pc=OZD | 12/18/2024 | 02/04/2025 |  | 01/07/2025 |  |  | 12/18/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the pump kit was opened and access obtained placing 14fr introducer into right femoral artery. during backloading of impella, the wire broke in the sterile field. the physician made the decision to open a new device. a new device was successfully implanted. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use related, during backloading of impella, wire broke the sterile field as per the clinical description provided.
12398 | 1220648-2025-25362 | 21091604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091604&pc=OZD | 12/20/2024 | 01/24/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs). approximately six days after start of support, the automated impella controller (aic) had persistent "purge disc not detected" alarms overnight and into the next morning. the alarm was unresolved despite replacing the purge tubing. however, after exchanging to another aic, the "purge disc not detected" alarm resolved. there was no report or indication of interruption of support; impella mcs continued until the patient was successfully bridged to transplant approximately nine days later. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. according to device and data analysis the root cause of this issue was a broken purge disc retainer e4 should have been left blank on manufacturer device report 1220648-2025-25362.
12399 | 1220648-2025-25363 | 21091616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091616&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/23/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577237 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Hospitalization | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received from the protect iv study regarding an unknown age patient that had low filling pressures at the beginning of the impella implant procedure. the pressure was noted to be worse after the procedure and the following day. there was no major bleeding; however, there was a small drop in hemoglobin. the hypertension was felt to be caused by a mix of hypovolemia and myocardial thinning. the outcome of the event is reflected as ongoing. | [5 empty] | The investigation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. additional information that was received and the following fields were updated since the initial manufacturer device report number 1220648-2025-25363 was submitted. a.1, a.2, a.3, a.4 and a.5. b.7. d.1 brand name. d.4 d.6a and d.6b. e.1 initial reporter name and address. h.4 device manufacture date b.2 revised selection as the previously selection was incorrect on the initial manufacturer device report number 1220648-2024-12650. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-12650 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 codes 4641 and 4607 were reported incorrectly on the initial manufacturer device report number 1220648-2024-12650 and a new code was added. | Patient- and device-specific information, initial reporter and implant/explant dates are unavailable at the time of this medical device report writing and have been reflected as "unknown" or field left blank accordingly. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12400 | 1220648-2025-25364 | 21091706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091706&pc=OZD | 12/22/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/22/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565518 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and following a right iliac dissection was noted and prompted removal of the impella cp. it was noted that the patient has an "extremely" tortuous anatomy, and it is unknown when the dissection occurred. the information indicated an intra-aortic balloon pump was implanted to support the patient. the patient was reported to be stable following the event. | [5 empty] | Patient information and implant/explant dates are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. as per the clinical information provided, the root cause of the vascular damage issue was most likely patient condition related due to tortuous vessels.
12401 | 1220648-2025-25365 | 21091709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091709&pc=OZD | 12/16/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/16/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576672 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and during support experienced thrombocytopenia and ventricular tachycardia. the patient presented to the emergency department for st-elevation myocardial infarction. left heart catheterization revealed 100% left anterior descending artery occlusion. the patient was stented, and next day returned to the catheterization lab for an intra-aortic balloon pump secondary to biventricular heart failure. the right ventricle was medically managed, and the impella 5.5 was successfully implanted with transplant workup to follow. approximately three days into impella mcs, heparin was discontinued with a switched to argatroban due to heparin-induced thrombocytopenia (hit). hemoglobin noted to drop four points since impella insertion. however, no blood unit was given at the time of the note. next day, swelling in the right upper extremity was noted, and the following day, an ultrasound revealed deep vein thrombosis. treatment for this event is unknown. five days following, the patient hit positive again and argatroban started. the next day, thereafter, the patient had a 45-beats run of ventricular tachycardia. echocardiogram was completed and revealed the inlet was interacting with the apex. the device was pulled back one centimeter, pump free flowing in the left ventricle. latest update, patient now approximately 55 days on support, noted the patient is still hit positive and not a candidate for advanced options at this time. the patient ambulates daily and remains on impella mcs. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the purge cassette leak, ventricular tachycardia (vt), and thrombocytopenia has been completed. the returned product was inspected and there were no physical abnormalities. the pump was primed for about 30 minutes and there was no purge leak detected. the root cause of the purge leak - pump was most likely due to a use-related issue. the cause of the thrombocytopenia and the vt was not determined due to insufficient clinical details. b.1 added product problem due to inivestigation results. d.6b explant date was added. d.9 device returned date was added. h.6 code 1250 was added due to investigation results. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25365. h.6 code 440 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25365.
12402 | 1220648-2025-25366 | 21091848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091848&pc=PYX | 12/21/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/21/2024 | 01/07/2025 |  |  | 08/31/2025 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024396597 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the attempted implant of an impella rp device for mechanical circulatory support (mcs) to a patient with cardiogenic shock from an acute myocardial infarction was unsuccessful as the impella rp was unable to cross the valve. amplatz and lunderquist wires were added without success as the physician was still unable to cross the valve with pigtail of impella rp. two kinks developed in shaft decreasing push ability. the decision was made to hold the impella rp implant, manage the patient's condition with medications, and reassess. the patient was eventually supported with an impella rp flex, which was successfully implanted the following day. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) has been completed. the root cause of the delivery issue was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the kinked cannula is the result of delivery issue.
12403 | 1220648-2025-25367 | 21091860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091860&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/18/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was moderate arterial bleed at the impella site. angle matching and manual pressure was performed to no avail. a computed tomography scan was performed and the patient was found to have an extravasation and hematoma extending up to retroperitoneal space. vascular surgery was consulted and patient was taken to operating room for femoral artery cutdown and impella removal. the patient remained stable throughout the procedure and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage was most likely use-related as not all best practices were utilized.
12404 | 1220648-2025-25368 | 21091985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21091985&pc=OZD | 12/23/2024 | 02/04/2025 |  | 01/07/2025 |  |  | 12/23/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591366 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and approximately three days after start of support, the sterile sleeve of an implanted impella cp pump detached from the touhy-borst connection. the staff was advised on how to maintain sterility should repositioning be required. support with the implanted pump continued uninterrupted. there was no report of patient harm as a result of the event. the device was successfully weaned and explanted the following day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. the pump was not returned for investigation. a data log review was not required for this case run. according to the clinical report, the sterile sleeve was found loose from the tuohy-borst connection. the cause of the product damage issue was not determined since the product was not returned and sufficient clinical information was not provided.
12405 | 1220648-2025-25369 | 21092090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092090&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  | 12/31/2024 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292785 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 placement was aborted due to patient¿s anatomy. the patient is stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, impella couldn't pass through the vasculature as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12406 | 1220648-2025-25370 | 21092165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092165&pc=OZD | 11/14/2024 | 02/18/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528316 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was documented via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured non-sustained ventricular tachycardia with a probable relationship to the implanted impella 5.5 device. beta blockers were increased in response to the condition. it was additionally reported that the patient experienced recurring ventricular contractions with a possible relationship to the impella device during support. support continued with the implanted pump despite the conditions. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause was not determined. | G.2 added an additional report source selection as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25370.
12407 | 1220648-2025-25371 | 21092227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092227&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/21/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559897 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted to support a patient with acute myocardial infarction and cardiogenic shock. while on support the impella had a few impella stopped alarms. the team was able to successfully restart the cp. heparin had been stopped by the doctor because of gastrointestinal bleed. it was further reported that the patient was placed in comfort care as the patient was not improving. the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the pump stop and gastrointestinal bleed (gi) has been completed. pump was not returned for investigation. data logs were investigated and "pump stop" alarms were confirmed. there was a sudden motor current spike and an auto p-level change response along we flows dropping to zero. purge pressure was seen to rise promptly as well. these are symptom seen in a sudden clot ingestion. clinical mentions that heparin was stopped due to internal bleeding which further supports clot formation in the body. therefore, the root cause of the pump stop issue was most likely biomaterial. as per the clinical information provided, there was a gi bleed but no additional information was provided. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12408 | 1220648-2025-25372 | 21092239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092239&pc=OZD | 11/29/2024 | 02/03/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485780 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) impella registry, that a patient suffered from acute limb ischemia during impella cp support. the patient developed evidence of arterial occlusion (cool extremity with decreased pulse). an arterial duplex confirmed right lower extremity ischemia due to an occluded right common femoral artery from the impella cp sheath and arterial plaque. the pump was subsequently removed and the patient was escalated to impella 5.5 support. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12409 | 1220648-2025-25373 | 21092440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092440&pc=OZD | 12/19/2024 | 07/03/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584465 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, oozing was noted from the impella insertion site and open chest site. two units of red blood cells were given. heparin was discontinued. the device was explanted weeks later and the patient survived the explant. | Additionally, it was reported that after 18 days of support the impella cp experienced immediate stoppage and the impella was removed. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and pump stop has been completed. the impella device was not returned for evaluation. pump stop: clinical details noted that pump stop occurred after 18 days on support. data logs were reviewed and real time (rt) log at time of pump stop showed a motor current spike with a pump stop alarm 13. pump was attempted to be restarted multiple times without success. the root cause of the pump stop could not be definitively determined as no product was returned for evaluation and limited clinical details at time of pump stop were provided. access site bleeding - major: clinical details noted that both impella access site and open chest were oozing and two units of blood were administered and manual pressure was applied to manage bleeding. heparin was paused and activated clotting time (act) was at 189 at time of bleeding. the root cause of the access site bleeding was most likely patient condition related as patient was bleeding from multiple access sites. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25373. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25373. h6 health effect - impact code 4627 and medical device problem code 1503 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-25373. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-25373.
12410 | 1220648-2025-25374 | 21092441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092441&pc=OZD | 12/19/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  | 09/30/2025 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397132 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant was placing a 5.5 pump for the support of a patient admitted in acute myocardial infarction / cardiogenic shock. the pump became damaged with the insertion attempts. the kinked pump was removed and replaced by a cp. the patient remains stable. | [5 empty] | The investigation of product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient condition as the patient had tortuous vessels. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12411 | 1220648-2025-25375 | 21092519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092519&pc=OZD | 12/20/2024 | 02/18/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025584897 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that five days into patient support, an impella 5.5 pump began to experience high purge pressure coinciding with a drop in purge flow rate. lysis therapy was initiated and the issue resolved. there was no resulting patient harm. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12412 | 1220648-2025-25376 | 21092641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092641&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547575 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had dark urine with confirmed hemolysis. the pump was repositioned and the patient's lactate dehydrogenase was trending down. it is unknown if blood products were given to the patient. | [5 empty] | The investigation for the hemolysis has been completed. complaint device not returned for analysis. data logs were reviewed finding no motor current (mc) spike observed at time of the reported issue no continuous suction alarms occurred. no placement signal not reliable alarms observed. no other placement signal alarms seen at the time of the issue. clinical stated dark urine was observed. confirmed hemolysis with trended lactate dehydrogenase (ldh). pump position was deep, followed by impella repositioning. ldh down to 1500 after repositioning of impella and issue solved. the root cause of the hemolysis issue most likely was due to improper positioning. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12413 | 1220648-2025-25377 | 21092787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092787&pc=OZD | 12/18/2024 | 02/20/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553437 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. during insertion, the catheter shaft completely kinked and bent, making it impossible to place in the patient. the patient had an infection, which made it difficult to retrieve the catheter. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. based on clinical description, the root cause of delivery issue was most likely due to patient condition. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12414 | 1220648-2025-25378 | 21092809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092809&pc=OZD | 12/18/2024 | 02/19/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570683 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Renal Failure | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, low pump flow, high purge pressure, and persisting purge pressure alarm occurred. the team used available tissue plasminogen activator (tpa) protocol. additionally, the patient endured renal failure, and reduced liver function. heparin-induced thrombocytopenia (hiit) was confirmed. heparin was switched to argatroban. pump in ventricular malposition. the doctor tried to reposition pump at bedside, impossible to reposition pump, decision to explant pump surgically. upon explant massive thrombus in 8mm graft was visible, pump catheter was stuck in thrombus. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure, positioning issues, renal failure, thrombocytopenia, and thrombus has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. the impeller and purge gap were clear of biomaterial. the pump was primed and ran successfully without noted purge-related abnormalities. the data log shows a 25-minute gradual rise in purge pressure, followed by a high purge pressure alarm. this condition was resolved after 20 hours, with a decreasing trend noted after administering tissue plasminogen activator (tpa). no other abnormalities were reported related to the positioning issue. all optical parameters were within specification range, and the only trend noted was in the placement signal correlating with motor current. according to the clinical findings, the pump was found deep in the ventricle at the end of the case, and the doctor decided to replace the pump. additionally, heparin was switched to argatroban after the patient was diagnosed with heparin-induced thrombocytopenia (hit). renal failure continued even after pump explant. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and returned pump investigation. the cause of the positioning issue was not determined since sufficient clinical's information was not provided. the cause of the renal failure issue was not determined due to insufficient clinical details. the cause of the thrombocytopenia and thrombus issues were not determined due to insufficient clinical details. d9 date device returned to manufacturer added. b3 date of event was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25378.
12415 | 1220648-2025-25379 | 21092971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21092971&pc=OZD | 12/20/2024 | 01/24/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1503160 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller was in use when the controller cart leg broke and the console fell to the floor. further details are unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - cart issue issue and aic - damage during therapy has been completed. some minor markings were found on some parts of the front and rear housing but the console still functions normally, there were no issues found during the case. the device functioned/operated to specification. the aic cart was not returned and the cause of the issue was not determined since product was not returned.
12416 | 1220648-2025-25380 | 21093001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21093001&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552631 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp there were suction events overnight and cp was repositioned under echo. subsequently, the patient developed red tinged urine and hemoglobin dropped. the next day hemolysis continued. bedside echo showed collapsed left ventricle with deep impella position. impella was pulled back. suction resolved. small pericardial effusion noted. patient started experiencing increased intermittent ventricular tachycardia (vt) run requiring direct current cardioversion x 1. more suction events occurred overnight, and intermittent vt. echo in the morning showed questionable position toward mitral valve. care team monitored closely until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis, positioning and the ventricular tachycardia has been completed. the pump was not returned for investigation. the data log shows the pump position aorta alarm with a trend in placement signal and several suction alarms on 12/19. there were no issues reported with the optical signal or motor current abnormalities. according to the clinical report, the pump was found shallow during the hemolysis event and was corrected. there was an increase in intermittent vt while the pump was found slightly deep. hemolysis started resolving after this point. additionally, a small hematoma was noted during the pump implant. the cause of the hemolysis issue was improper positioning of the device, as the corrected pump position resolved the hemolysis, and this was recorded in the data log. the root cause of the positioning issue was as a result of the hemolysis issue. the root cause of the vt could not be determined without additional details.
12417 | 1220648-2025-25381 | 21093049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21093049&pc=OZD | 12/19/2024 | 01/30/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573797 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the doctor noted extravasation under fluoroscopy at the sheath site. the doctor felt there may be a tear at the insertion site due to proximity to bifurcation. the decision was made to explant the current impella cp over an 0.035 wire. a 16 french dry seal sheath was placed and did prevent extravasation per fluoroscopy. a new impella cp was then placed. the patient is stable and the procedure was successful with another device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. due to the lack of product returned and insufficient clinical details, the root cause of the vascular damage - peripheral vessel could not be determined.
12418 | 1220648-2025-25382 | 21093175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21093175&pc=PYX | 12/20/2024 | 01/28/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561563 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support. while on support, a large hematoma was noted at the impella site. no heparin was given and femostop was applied. care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since the hemostasis was achieved after applying femstop and discontinuing heparin. b.2 revised selections as required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25382.
12419 | 1220648-2025-25383 | 21093211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21093211&pc=OZD | 12/19/2024 | 01/14/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577827 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed without any issues. when percutaneous coronary intervention was being performed with cp support the physician placed the 14f peel away sheath, but the left half of the orange circle came off the sheath when dilator was removed. the doctor replaced the orange piece and elected to move forward. single access with 7f companion sheath utilized. sheath started leaking near the end of the case. companion sheath was removed. peel away sheath removed and replaced with repositioning sheath. blood loss was minimal, and no transfusion needed. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. no product or images were returned for analysis. clinical reports all best practices were followed and mention calcification of vessels. the root cause of introducer damage - mechanical was not determined as no product or images were returned for analysis. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
12420 | 1220648-2025-25385 | 21093744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21093744&pc=OZD | 12/19/2024 | 01/29/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  | 08/31/2025 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398309 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was inserted to support a patient with ami/cgs. the patient was transferred to another facility on impella support. upon arrival it was noted that the patient had dark red urine. the cp was explanted, and the patient was escalated to impella 5.5 due to hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2025-25385 in accordance with updated procedures. h6 code 1884 and 4627 were reported incorrectly on manufacturer device report 1220648-2025-25385. new codes were added to health effect - clinical code and health effect - impact code.
12421 | 1220648-2025-25386 | 21094387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21094387&pc=OZD | 12/20/2024 | 04/02/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508319 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that two days into support, an implanted impella cp pump experienced a sudden suction problem coinciding with a self-limited ventricular tachycardia. the motor current peaked intermittently with persistent suction issues for a "few" hours. the staff was advised to explant the pump, but the decision was made to maintain support. | Despite that decision to keep the patient on support, the pump was removed later that morning and the patient expired. there is not enough information to disassociate the use of the impella with the patient's passing. |  |  |  |  | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | Additional information provided in d9, h3, h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post-sterile inspection checks. the pump was returned for investigation. no physical abnormalities were reported on the returned product. some small clear biomaterial was noted in the purge gap, but it was also cleared out during the test run in a bucket of water. the pump operated as per the specifications without any noted low pump flow. the data log shows that the pump flow reduced after the first motor current spike during the case, and a few more motor current spikes were reported in the 2 hours before the flow returned to the normal range. the cause of the low pump flow issue was most likely biomaterial, based on data log review and return product investigation. cardiac arrythmia was reported during impella support. to make a determination as to the cause of the cardiac arrythmia, clinical information must be considered in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. | B.2 revised from the initial manufacturer device report number 1220648-2025-25386 to reflect the patients death and date of death. required intervention was reported incorrectly on the initial manufacturer device report number 1220648-2025-25386. b.5 added patient outcome information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25386. e.1 revised facility name as it was previously reported incorrectly on the initial manufacturer device report number 1220648-2025-25386. g.1 manufacturing site name and address was correct on initial manufacturer device report number 1220648-2025-25386. this section was reported incorrectly on the follow up 1 manufacturer device report number 1220648-2025-25386 and was revised back to what was reflected on the initial manufacturer device report number 1220648-2025-25386 h.1 type of reportable event was revised from the initial manufacturer device report number 1220648-2025-25386 to reflect the patients outcome. h.6 codes 1721, 4611, 3331 were all reported incorrectly on the initial manufacturer device report number 1220648-2025-25386. codes 2095, 4110 and 1802 were added as they were inadvertently omitted from previously submitted manufacturer device report number 1220648-2025-25386. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-25386 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
12422 | 1220648-2025-25387 | 21095315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095315&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502435 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that several attempts were made to deliver the impella 5.5 across the aortic valve but were unsuccessful due to stenosis. impella 5.5 placement was aborted and impella cp was inserted successfully. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation.the root cause of delivery issue is determined to be patient condition because the impella 5.5 was unable to cross the stenosis of valve but the implantation of the impella cp was successful. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12423 | 1220648-2025-25388 | 21095470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095470&pc=OZD | 12/19/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527352 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had access site bleeding. intervention is unknown but heparin was removed from the purge after the bleeding. it is not confirmed if this was due to the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
12424 | 1220648-2025-25389 | 21095643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095643&pc=OZD | 12/20/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552604 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and the patient went into atrial fibrillation. the patient was cardioverted and a bedside echocardiogram was performed. the pump showed a little deep so staff pulled it back some. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | It was further reported that the pump encountered purge pressure high alarms followed by purge system blocked alarm. all troubleshooting was performed and the cassette was changed with no resolution. tissue plasminogen activator was initiated for purge fluid and there was intermittent resolution of alarms throughout the morning. full resolution occurred and the purge was switched back to bicarbonate. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure and the atrial fibrillation has been completed. the data logs were evaluated and "high purge pressure" and "purge system blocked" alarms were confirmed. the logs showed no motor current spikes outside p-level changes. the logs showed a gradual rise in purge pressure till high purge pressure alarm was triggered. clinical information mentioned that tissue plasminogen activator (tpa) was administered and the logs aligned and showed that purge pressure returned to normal values. the returned product was inspected and there were no physical abnormalities. the pump was run in the lab and high purge pressure did not reproduce. the root cause of the high purge pressure was determined to be biomaterial. the root cause of the atrial fibrillation could not be determined without additional information. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b.1 product problem was added since the initial manufacturer device report number 1220648-2025-25389 was submitted due to additional information received. b.5 additional information was received. d.9 device return date/information was added. h.6 additional codes were added due to additional information received. h.10 additional instruction for use were added due to additional information received.
12425 | 1220648-2025-25390 | 21095650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095650&pc=OZD | 09/08/2024 | 02/05/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  | 12/31/2025 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435723 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia (vt) with a "possible" relationship to the impella device used: after leaving the operating room, the patient had multiple vt events. the patient was started on an amiodarone bolus. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of ventricular tachycardia has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the ventricular tachycardia was not determined. h.5 revised as this field was inadvertently left blank on the initial manufacturer device report number 1220648-2025-25390.
12426 | 1220648-2025-25391 | 21095828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095828&pc=OZD | 12/20/2024 | 02/03/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, hematoma and bleeding was noted after the peel away sheath was removed and the reposition sheath was inserted. a femostop was needed. the doctor perforated the coronary artery and gave thrombin which then clotted off another vessel which led to the doctor giving 20,000 units of heparin. perforation required a covered stent and/or balloon tamponade. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - great vessel has been completed. the pump was not returned for investigation. as per the clinical information provided, while inserting the pump doctor perforated the coronary artery which led to perforation as well as bleeding at the access site and hematoma. patient had occluded vessels. therefore, the root cause of the vascular damage issue was most likely patient condition related diseased vessels. d6b explant date was omitted from manufacturer device report 1220648-2025-25391
12427 | 1220648-2025-25392 | 21095946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21095946&pc=OZD | 12/20/2024 | 02/05/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556187 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient developed a gastrointestinal bleed (gib). heparin was reduced and staff was concerned for clots. the pump was explanted the following day. the patient survived the event. | [5 empty] | The investigation for the gastrointestinal (gi) bleed has been completed. pump was not returned for investigation. as per the clinical information, the patient experienced a gi bleed which is not related to impella. therefore, the root cause of the vascular damage issue was not determined since relation between impella and gi bleed is unknown. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12428 | 1220648-2025-25393 | 21097519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097519&pc=OZD | 12/18/2024 | 02/03/2025 |  | 01/07/2025 |  |  | 12/18/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583388 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a hematoma at the access site and a possible gastrointestinal bleed (gib). the hematoma was resolved with manual pressure and managing the patient's activated clotting time with medication adjustment. heparin was used in the purge and may have contributed to the gib. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the gastrointestinal bleed (gi) and access site bleed has been completed. the pump was not returned for investigation. the data log shows a temporary pump stop during the case with saturated pump flow and a motor current spike to 1400 on 12th day of pump use. the pump was successfully able to start and returned to normal operation. pump ran another 3 days after temporary pump stop. also, several motor current spikes were reported after the temporary pump stop. as per the clinical, right groin has small hematoma, pressure held and hematoma resolving. also, multiple site was bleeding at the end of case run and heparin was held. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not returned. the cause of the vascular damage issue was not determined since product was not returned ad sufficient clinical information was not provided. the relation between the gi bleed and the impella was unknown. the cause of the temporary issue was not determined since product was not returned. the product use exceeded the 8 days of design life.
12429 | 1220648-2025-25395 | 21097530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097530&pc=OZD | 12/19/2024 | 01/24/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572823 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Hematoma | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support while on support, the nurse called in stating the purge fluid is leaking out of the filter/purge reservoir area. additionally, a hematoma was noted at the axillary access site. no intervention was taken. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the product was not returned. due to limited clinical details and no product returned, the root cause of the purge leak - pump cannot be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2025-25395.
12430 | 1220648-2025-25396 | 21097644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097644&pc=OZD | 12/16/2024 | 02/18/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547021 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp suddenly stopped on day 14 of support. several unsuccessful attempts were made to restart. the impella was removed, and the patient was placed on intra-aortic balloon pump with extracorporeal membrane oxygenation. a red blood clot was found inside the cannula and at the outflow cage, along with linear biological material. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Additional information provided in h6 and h10. the investigation of pump stop and thrombus has been completed. the data logs were reviewed, and on the 14th day of support there was a sudden trigger of alarm 13 - high motor current pump stop. leading up to the event, purge pressure/flow were stable, but there were several distinct steps up in the motor current. then, a final spike in the motor current above 1400 ma triggered the first pump stop alarm. the pump was able to be restarted for short periods (minutes) several times, but in the end, a final pump stop upon trying to restart led to the pump being disconnected. clinical details report that thrombus was noted to be in the outflow cage upon explant. per design trace matrix, impella cp's design life is 8 days, and the pump was on the 14th day of support. however, without returned product, it cannot be confirmed what the cause of the motor current spike and subsequent pump stop was. as the necessary information was not provided, the root cause of the thrombus was not determined. corrected information provided in e4. note: e4 should have been left blank on the initial manufacturer device report.
12431 | 1220648-2025-25397 | 21097657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097657&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541821 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an impella cp implant and after the implant, the patient was shown deep and the pump was kept as is. the patient developed red urine and the p level was lowered. the patient's urine became more clear with the lowering of the p level and creatine levels were trending down. no blood products were given and support continued. | [5 empty] | The investigation of positioning issues has been completed. data logs showed the pump ran without issue during support. the product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number. d.6b explant date was inadvertently omitted from the initial manufacturer device report number and was added. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number was submitted in accordance with updated procedures. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12432 | 1220648-2025-25400 | 21097661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097661&pc=OZD | 12/19/2024 | 02/05/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579754 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had an apical thrombus visible on transthoracic echocardiogram. the thrombus was not moving. support continued and patient remained stable. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25400 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12433 | 1220648-2025-25398 | 21097668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097668&pc=OZD | 12/19/2024 | 02/03/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583410 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. impella support was aborted due to poor vascular access and unable to advance impella. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient had poor vascular access. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12434 | 1220648-2025-25399 | 21097669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097669&pc=OZD | 12/19/2024 | 02/05/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580922 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella 5.5. after multiple attempts to pass impella 5.5 pass the brachiocephalic artery into the ascending aorta, the case was aborted due to tortuous anatomy. no patient harm has been reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had tortuous anatomy. e4 should have been left blank on manufacturer device report 1220648-2025-25399. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12435 | 1220648-2025-25401 | 21097670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097670&pc=OZD | 12/19/2024 | 02/05/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  | 02/28/2026 |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474077 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient endured right axillary bleeding impella 5.5 access site with a "definite" relationship to the impella device. the patient developed some bleeding at the impella 5.5 access site at the right axillary artery. dressing changed, heparin held, blood products transfused. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: d4 updated model number, g1 mdr reporting contact email.
12436 | 1220648-2025-25402 | 21097673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097673&pc=OZD | 12/20/2024 | 01/29/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the plastic holder of the (automated impella controller) aic purge pressure disc is defective. the aic was exchanged. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. corrections to the initial mfr report: d4 updated model number g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12437 | 1220648-2025-25404 | 21097677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097677&pc=OZD | 12/23/2024 | 02/10/2025 |  | 01/07/2025 |  |  | 12/23/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580253 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported delivery issues with an impella cp to a patient with cardiogenic shock from an acute myocardial infarction. the patient presented for transcatheter aortic valve replacement (tavr) procedure. after valve deployment, the patient decompensated. it was suspected the valve leaflet likely occluded right coronary artery (rca). the patient received six shocks. impella was cp prepped, access gained; however, the pump insertion was aborted due to severe right ventricle (rv) dysfunction. a temporary pacemaker was placed, and ultimately, the patient was transferred for venoarterial extracorporeal membrane oxygenation mechanical circulatory support. there was no report of adverse effects to the patient as a result of the aborted impella implant. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump insertion was cancelled due to severe rv dysfunction. the instructions for use for the impella cp with smartassist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. corrections to the initial mfr report: added d6a/d6b. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12438 | 1220648-2025-25405 | 21097680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097680&pc=OZD | 12/23/2024 | 02/13/2025 |  | 01/07/2025 |  |  | 12/23/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1291424 |  | 0042-0000-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and upon start up, the impella cp pump experienced high motor current followed by a pump stop. both the impella cp pump and automated impella controller were replaced to resolve the issue and support the patient through the high-risk pci. it was reported there was no harm to the patient as a result of this event. | [5 empty] | The investigation of automated impella controller (aic) - pump stop has been completed. on 23-dec-2024, the pump was plugged into the console but was immediately unplugged a minute later. the pump was then plugged back in but unplugged 14 seconds later by the user. finally, the pump was plugged back in and allowed to fully prime and calibrate and the pump was started but immediately got a motor current high alarm causing a pump stop. this was then followed by a retrograde flow alarm due to the pump stop. the console restarted the pump and it was able to run for six minutes before the pump was removed from the console. the real time logs confirmed the initial pump stop but did not add any context for what caused it. the console was not returned for investigation. since the complaint date, the console has been used and has not reproduced the failure mode. the cause of the pump stop was not determined since the product was not returned and logs were not edifying. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25405 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 revised health effect - impact code as the previously entered code on the initial manufacturer device report number 1220648-2025-25405 was incorrect. | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for the impella cp serial number (b)(6) with date of event 23-dec-2024 and identifier ending in (b)(6).
12439 | 1220648-2025-25406 | 21097686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097686&pc=OZD | 12/21/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/21/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient newly diagnosed with "severe" cardiomyopathy was implanted with an impella cp for mechanical circulatory support and post implant experienced access site bleeding. the impella cp was successfully implanted via the right femoral artery. the peel away sheath was removed, the repositioning sheath was advanced, secured and sent to the unit in stable condition. on the unit approximately couple hours later, oozing at the access site was observed and addressed with angle match and femostop placement. the touhy was locked at 90cm and a knee immobilizer placed. it is unknown if anticoagulation was held given the bleeding. however, the patient was maintained until transfer to outside hospital for heart transplant work up which occurred the following day. | [5 empty] | The impella device was not received from the customer as it appears the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation into the reported issue has been completed. no device was received for evaluation. device history review confirmed the pumped passed all post sterile inspection checks. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
12440 | 1220648-2025-25407 | 21097689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21097689&pc=OZD | 12/20/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574176 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and after the implant, the patient had atrial fibrillation (af) with sustained suction events. volume was given with one unit of packed red blood cells. additionally, while on support, there was access site bleeding. the jackson-pratt drain output was approximately one liter and the patient had a drop in hemoglobin. two units of packed red blood cells, one unit of fresh frozen plasma and albumin were administered. support continued. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleed, low pump flow, and atrial fibrillation has been completed. complaint device not returned for analysis. data logs were reviewed finding no motor current (mc) spike observed. many unknown position alarms occurred. suction alarms occurred after case initiated. the root cause of the low pump flow was patient related as the suction and low pump flow resolved after volume was given. the root cause of the access site bleed and atrial fibrillation could not be determined without additional information. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock stated: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
12441 | 1220648-2025-25394 | 21098090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21098090&pc=OZD | 12/19/2024 | 02/12/2025 |  | 01/07/2025 |  |  | 12/19/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported during booting process the automated impella controller restarted several times and showed at the end the message that there occurred a failure during booting. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. the console logs from the reported day of event are consistent with complaint and show that console presented with system self-check failure, due to permanent loss of communication to one of its internal components. the console was inspected and corrosion on power battery manager (pbm) affecting pbm1 serial communication, was found near supercaps c69, c70. the cause of the aic issue was contamination. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2025-25394. h5 was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25394. h8 was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25394. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12442 | 1220648-2025-25403 | 21098113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21098113&pc=OZD | 12/20/2024 | 02/07/2025 |  | 01/07/2025 |  |  | 12/20/2024 | 01/07/2025 |  |  |  |  | 01/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520517 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the 7fr companion sheath was contaminated when the patient kicked and dropped sterile equipment of the bed. no patient harm has been reported. | [5 empty] | The investigation for the introducer damage has been completed. data logs are not relevant to the complication and product was not returned for investigation. clinical details provide that the companion sheath became decontaminated during case start of the single-access pci procedure. there was no harm to the patient and the procedure carried on without issue. based on clinical details provided, the root cause of the cosmetic defect was determined to be a use issue. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12443 | 1220648-2025-25321 | 21078539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078539&pc=OZD | 12/18/2024 | 02/04/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531446 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | Hemolysis | Ischemia | [10 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp device was placed without issues. however, since the patient had a small left ventricle, the patient experienced bouts of ectopy without causing any issues. an echocardiogram was completed and showed the impella cp pump angled towards the mitral valve. however, the physician ultimately determined the current positioning was acceptable. later the same day, the patient's lab showed hemolyzed despite urine output color and volume being ideal. additionally, it was noted there was no flow in the leg with impella access. there was no pulse in the foot and leg was cold. a bed-side echocardiogram was completed and noted the position looked optimal no longer interfering with the mitral valve. an evaluation for extracorporeal membrane oxygenation was planned. following, the impella was successfully weaned and explanted after 22 hours of support. it was reported the patient was stable following the event. | [5 empty] | The investigation has been completed. the pump was not returned for investigation. the total pump run time was 22 hours, and data logs were provided for the first 7.5 hours of the case run. there were no issues reported regarding improper position, suction, or placement signal on data log. according to the clinical information, labs showed hemolysis despite ideal urine output. there was no flow in the impella leg, and the foot was cold with no pulse. the impella appeared shallow on extracorporeal membrane oxygenation. there were also bouts of ectopy. the cause of the positioning issue was most likely use related, based on given clinical details. the cause of the hemolysis, limb ischemia, and cardiac arrythmia was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12444 | 1220648-2025-25322 | 21078654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078654&pc=OZD | 12/16/2024 | 01/27/2025 |  | 01/06/2025 |  |  | 12/16/2024 | 01/06/2025 |  |  | 09/30/2025 |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445467 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported while on support, they were unable to doppler pulses on admission to the intensive care unit. the patient experienced arrhythmia which was managed by temporary pacer. the patient¿s right leg was ischemic. the impella cp was removed due to limb ischemia. | [5 empty] | The investigation for the reported limb ischemia and arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and arrhythmia could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12445 | 1220648-2025-25323 | 21078674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078674&pc=OZD | 12/17/2024 | 01/23/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578578 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support placed for a patient admitted in with a non-st segment elevation myocardial infarction and ventricular tachycardia (vt). the 5.5 pump and extracorporeal membrane oxygenation (ecmo) was placed as escalation from a impella cp. the patient suffered had a single vt episode overnight that require a single shock. rhythm review and same morphology as admitting rhythm that patient was scheduled for vt ablation. during ablation, case was aborted due to hemodynamic instability requiring intra-aortic balloon pump and ultimately veno-arterial ecmo and impella. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
12446 | 1220648-2025-25324 | 21078748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078748&pc=OZD | 12/16/2024 | 02/05/2025 |  | 01/06/2025 |  |  | 12/16/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573005 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when a patient arrived at the unit from impella cp insertion, oozing was noted around the insertion site. two units of packed red blood cells were given. the site was redressed and surgicell was added. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
12447 | 1220648-2025-25325 | 21078843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078843&pc=OZD | 11/19/2024 | 02/26/2025 |  | 01/06/2025 |  |  | 12/19/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542847 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. approximately six days after start of support, an automated impella controller (aic) failure occurred resulting in a pump stop. the aic was swapped, and support was successfully reinitiated with the implanted impella 5.5 pump. the outcome of the patient as a result of the event was unknown. however, there was no indication or report of harm resulting from the failure. | [5 empty] | Added related manufacturer report number to h(10). | The investigation into the reported issue has been completed. no product was returned for evaluation. data logs were reviewed and real time logs at time of pump stop showed pump p-level was turned down and then ¿impella stopped¿ alarm was triggered. there were no abnormalities seen with the placement signal or purge pressure. multiple attempts to restart the pump were attempted without success. clinical details noted that pump was being turned down before pump stop occurred. additionally, the automated impella controller was changed and the pump was able to be restarted. the root cause of the pump stop could not be definitively determined as no product was returned for evaluation and limited clinical details were provided. | Following initial inspection of the automated impella controller (aic), approximately 30 days following the event, it could not be determined if the issue was solely linked to the aic although the pump restarted after exchange of the aic. therefore, both the aic and pump will be investigated. the impella 5.5 device was not received from the customer as it appears the device remains implanted, supporting the patient at the time of this medwatch report writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-24566 for the associated automated impella controller. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12448 | 1220648-2025-25326 | 21078973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21078973&pc=OZD | 12/20/2024 | 02/10/2025 |  | 01/06/2025 |  |  | 12/20/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575903 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 38-week pregnant patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and one day following start of impella mcs, high purge pressure was experienced. the patient presented to the emergency department one day prior with palpitations and chest pain electrocardiogram showed st-elevation myocardial infarction versus myocarditis. catheterization was completed and ruled out blockages. the following day, the patient was implanted with an impella 5.5 device and is now on venoarterial extracorporeal membrane oxygenation and impella mcs (ecpella). the patient was noted on continuous renal replacement therapy as well. the following day in the morning, tissue plasminogen activator (tpa) was started due to elevated purge pressure above 1100 and purge flow around 2.4. with tpa flows were noted to be improving. the patient continued on ecpella support with work up for orthotopic heart transplantation with hope of recovery. 12 days on support, the nurse reported automated impella controller (aic) had a battery failure. the battery was at 50%, and it was plugged back in, but the alarm continued. it was indicated the aic was replaced. the patient was eventually decannulated and the impella 5.5 device was successfully weaned and explanted in the cardiovascular operating room after a total of 14 days on support. the outcome of the patient following the event, post explant of the device is unknown. however, there was no report of resulting harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this is one of two medical device reports related to the patient implant experience. refer to initial medical device report for the automated impella controller serial number (b)(6) with date of event 30-dec-2024 and identifier ending in (b)(6). | The investigation for the high purge pressure has been completed. log review showed purge pressure rose to above 1000 mmhg and purge flow decrease within 15 minutes. at the beginning of the increase, there was a small motor current spike and shift, a small shift in pump flow, and a speed deviation. the data logs showed that the purge pressure returned to normal levels after approximately 15 hours post-tissue plasminogen activator (tpa). the root cause of the high purge pressure was most likely biomaterial as the logs indicate biomaterial and the high purge pressure was resolved with tpa.
12449 | 1220648-2025-25327 | 21079997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21079997&pc=OZD | 12/17/2024 | 01/31/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578928 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, they had an embolic stroke. additionally, a clot was found on the cp. patient was scheduled to go for 5.5 escalation but this has been deferred and the cp was removed due to discovery of evolving hemorrhagic strokes to areas of previous infarct. the patient remained on extracorporeal membrane oxygenation support and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details.
12450 | 1220648-2025-25328 | 21080357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080357&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025549767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fracture | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. during the procedure, the guidewire fractured and lower impella flows were noted. soon after, there was a hard pump stop. attempted multiple restarts, and new automated impella controller (aic) with no success. patient remained hemodynamically stable, and physician proceeded to pull cp into descending aorta for the remainder of the case. impella was removed due to pump stop. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and product damage (guidewire damage - great vessel) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12451 | 1220648-2025-25329 | 21080379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080379&pc=OZD | 12/18/2024 | 02/04/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  | 01/31/2026 |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456922 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient when after just under nine hours, there was an observation made of access site bleeding. there was an allegation of a retroperitoneal bleed. this prompted the explant of the impella cp and placement of the axillary impella 5.5 and extracorporeal membrane oxygenation. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
12452 | 1220648-2025-25330 | 21080417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080417&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545605 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and it was aborted due to an occluded iliac. attempt was made to pass sheath across right iliac lesion and was unsuccessful. the doctor opted not to treat peripheral disease and proceeded without impella support. the patient remained stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12453 | 1220648-2025-25331 | 21080476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080476&pc=OZD | 12/18/2024 | 01/29/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532207 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. they were unable to pass the impella 5.5 or impella cp through left axillary artery cutdown site. the doctor met difficulty with passing the catheter through the graft anastomosis. they observed a sharp angle entering into the aortic arch, and vessel size was smaller than originally measured. multiple attempts were made with each catheter before aborting. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to the size of patient's vessel. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12454 | 1220648-2025-25332 | 21080693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080693&pc=OZD | 12/18/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580274 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implant, the patient developed a hematoma at their incision site. manual pressure and systemic anticoagulation were held for several hours to manage the condition. support with the implanted pump continued following the intervention. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12455 | 1220648-2025-25333 | 21080811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080811&pc=OZD | 12/18/2024 | 02/05/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520004 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | Infusion or Flow Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. following implant, the pump was being repositioned when pump saturation was noted. a purge system blocked alarm appeared and the waveforms on the aic flipped. the waveforms eventually corrected themselves, but the purge system blocked alarm persisted and tissue plasminogen activator was administered. the following day, the patient was hypercoagulable and the impella experienced low flows. a transesophageal echocardiogram showed there was no movement of the heart coinciding with large clot formations in the left atrium and aorta. it was decided that no further action was possible and the pump was removed. the patient passed away shortly after. it is unknown if complications with the pump contributed to the patient's passing. | [5 empty] | The investigation for the saturated flow, low pump flow, high purge pressure, and thrombus has been completed. no product was returned. when the pump was started, flows were variable at p-4 and the motor current was flat. there were "impella position unknown" and "suction" alarms. about an hour into the case, there was a small shift in the motor current and increase in flows. about two hours later, there was a spike in motor current followed by a rapid increase of purge pressure to 1200 mmhg over the span of five minutes. "purge system blocked" and "high purge pressure" alarms were seen. saturated pump flows then occurred reaching almost 4.5 liters per minute at p-4. the high purge pressure lasted for about two hours and decreased along with the pump flows. the pump flows lowered and then remained low throughout the case. the patient was on extracorporeal membrane oxygenation (ecmo) and also receiving low flows on ecmo. there was a trend in the placement signal. once the pump was removed, there was a large clot formation in the left atrium and aorta. the root cause of the low pump flows is most likely due to the patients condition based on the logs and the lack of volume of the patient. the root cause of the high purge pressure is most likely due to ingested biomaterial based on the log signature and that tissue plasminogen activator resolved the higher pressures. the root cause of the saturated flow is most likely due to ingested biomaterial based on the log signatures. as the necessary information was not provided, the root cause of the thrombus was not determined. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25333 was submitted in accordance with updated procedures. h.6 additional codes were added to health effect - impact code section as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-25333. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12456 | 1220648-2025-25334 | 21080854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21080854&pc=OZD | 12/18/2024 | 01/31/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566833 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a thrombus in their left ventricle during impella cp support. in response to the noted thrombus, the decision was made to begin weaning the pump for explant. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
12457 | 1220648-2025-25335 | 21081060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21081060&pc=OZD | 12/18/2024 | 09/22/2025 |  | 01/06/2025 |  |  | 12/19/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574053 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, a thrombus was discovered in left ventricle (lv). per echocardiography report, there is a small mobile echo density consistent with thrombus attached to the ventricular portion of the impella 5.5 and there is thrombus in the left atrial appendage. the doctor stated the lv thrombus measured 3 mm. the recent trans-esophageal echo (tee) made no mention of thrombus. impella was repositioned. systemic heparin infusion was turned off at 0700 on (b)(6) due to excessive output from mediastinal chest tubes. impella was switched from bicarbonate purge to heparin purge. the patient will not be going back on systemic due to retroperitoneal (rp) bleeding and bleeding problems around inflow cannula of the extra corporeal membrane oxygenation circuit. the patient survived the event. | Additional information indicated that, during patient support, the pump lost its placement signal overnight. audio alarms had been silenced. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. upon review of data logs, it is apparent that the console is the potential cause of the issue and no problem was found with the pump. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem b5 updated with additional information. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-25335. | D.6b explant date was added. the investigation of vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information was not provided. the relation between mediastinal and impella was unknown. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ h.6 thrmobus code was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25335 and was added. h.10 revised instructions for use since the initial manufacturer device report number 1220648-2025-25335 was submitted to reflect thrombus.
12458 | 1220648-2025-25336 | 21081105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21081105&pc=OZD | 12/18/2024 | 01/31/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509850 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while placing the pump across the anastomosis, the impella 5.5 met resistance. the physician discussed the potential of vessel damage and then decided to go to an impella cp for support. the chest was already opened. however, the cardiovascular system did not have enough space to place the graft to implant. the impella 5.5 was removed and the impella cp was placed without issue. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since resistance was met due to patient's anatomy during implant. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12459 | 1220648-2025-25337 | 21081361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21081361&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579128 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Priming Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during use of an impella 5.5 pump for mechanical circulatory support. it was documented that during set up, high purge pressure alarms appeared despite no issues being noted with the setup or components. the console was replaced. however, purge alarms continued to appear. in response to the ongoing alarms, a de-air process was completed, but a purge open alarm subsequently appeared. the pump was implanted despite the noted issue and was shortly followed by a purge system blocked alarm. the pump shut off to p0 and the decision was made to replace the device. there was no patient harm resulting from the failure. | [5 empty] | The investigation of high purge pressure, pump stop, and priming issue has been completed. the impella 5.5 and purge cassette were returned for evaluation. upon visual inspection, no abnormalities were noted. the pump was purge in a bucket of water for approximately 15 minutes and high purge pressure did not reproduce. the pump was turned onto p-9 and was run for five minutes and high purge pressure and pump stop did not reproduce. purge values remained steady during pump run in the lab. real time log from first console used showed high purge pressure occurred immediately after pump was initiated. high purge pressure seems to have resolved after ¿purge system open¿ alarms occur. no logs were able to be obtained from the second console used when the pump stop occurred as the console number was not provided. clinical details noted that high purge pressure occurred immediately after initiating pump. additionally, "purge system open" and "air in purge" alarms were noted during case. the root cause of the priming issue could not be definitively determined as the returned pump was able to purge and run while in lab. the root cause of the high purge pressure could not be definitively determined. the root cause of the pump stop could not be definitively determined as pump stop was not reproduced in the lab and no data logs were provided at time of pump stop. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12460 | 1220648-2025-25338 | 21081595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21081595&pc=OZD | 12/18/2024 | 02/04/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1579730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported controller error on the automated impella controller (aic) during prep. there was no reported patient involvement. | [5 empty] | The impella device was received from the customer on (b)(6) 2024 therefore, an investigation of the device is underway, a supplemental mdr will be filed once the investigation is completed. | The investigation for the automated impella controller (aic) controller failure has been completed. data logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure as well as multiple "algs suspended opm crc error" messages. real time logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. the reported complaint of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)¿ was determined to be a firmware issue. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. a.1 revised as information was inadvertently omitted from initial manufacturer device report number 1220648-2025-25338. d.4 revised serial number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25338. e.1 revised name prefix/title, first name and last name as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25338. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25338 was submitted in accordance with updated procedures.
12461 | 1220648-2025-25339 | 21081873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21081873&pc=OZD | 12/18/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  | 10/31/2025 |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411160 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. following insertion, it was noted that the device had difficulty advancing through the graft anastomosis. it was additionally documented that there was a sharp angle entering into the aortic arch and the vessel size was smaller than originally measured. this resulted in the cannula kinking during the delivery attempt. after multiple failed attempts, the decision was made to abort the implant. there was no resulting patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to the size of patient's vessel. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist f section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 information was inadvertently omitted from initial manufacturer device report number and was added. g.1 revised reporting contact email since initial manufacturer device report number was submitted in accordance with updated procedures. h.6 code 4629 was previously entered incorrectly on initial manufacturer device report number. a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number in accordance with updated procedures. | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12462 | 1220648-2025-25340 | 21083576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21083576&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580267 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during an attempted implant of an impella 5.5 device for mechanical circulatory support. following insertion, it was documented that the patient's anatomy prevented the device from being placed successfully. the device was removed and axillary access was obtained on the opposite side, but again, the pump was unable to advance due to the patient's anatomy. the cannula subsequently kinked and the decision was made to abort the implant. an impella cp device was then placed for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
12463 | 1220648-2025-25341 | 21084609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21084609&pc=OZD | 12/14/2024 | 02/18/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571590 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | A complainant reported the patient was on impella cp support. while on support, the patient had limb ischemia. the pump was explanted. issue was resolved by the vascular surgery. care was withdrawn and the patient expired. reportable due to patient harm and potential relationship between the cause of death and the impella. | Additional information from the clinical site was received indicating that the patient did not expire. the impella was explanted due to acute bronchial bleeding following resuscitation and emergency intubation. anticoagulation was therefore no longer possible. the patient is hemodynamically stable, and the leg was documented to be observed, but no ischemia. the patient survived the event. |  |  |  |  | Date of death is unkonwn. unknown udi # and manufacturing date. the impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | As noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.2 and a.3 revised as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25341. b.2 added date of death as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25341. d.4 revised model number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25341. d.4 added expiration date and unique identifier (udi) # as they were inadvertently omitted from initial manufacturer device report number 1220648-2025-25341. e.1 revised address line 1 as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25341. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2025-25341 was submitted in accordance with updated procedures. g.1 revised manufacturing site name and address (all fields) as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25341. h.4 added device manufacture date as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25341. h.6 added health effect - impact code 4624 as it was inadvertently omitted from initial manufacturer device report number 1220648-2025-25341. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2025-25341 in accordance with updated procedures. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | B.2 death selection and date of death were previously entered incorrectly on manufacturer device report number 1220648-2025-25341 and should of been left blank. b.5 corrected information from the clinical site was added since the initial and first follow up manufacturer device reports number 1220648-2025-25341 were submitted. h.1 due to new corrected information received the type of report was revised since the initial and first follow up manufacturer device reports number 1220648-2025-25341 were submitted. h.6 health effect - clinical code 1942 and health effect - impact code 1802 and 4624 were entered in error on previous manufacturer device reports number 1220648-2025-25341. new health effect - clinical code and health effect - impact code were added.
12464 | 1220648-2025-25342 | 21085203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21085203&pc=OZD | 12/18/2024 | 02/06/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542524 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the md place impella cp for support as the patient had a spiral dissection from the lm-lad. the impella was inserted without difficulty and started in auto. md elected to fix the lad. repo sheath inserted and sutured to the patient. removing the drape, staff noted that there was a hematoma on the left groin where the impella and 6fr central line was. manual pressure was held on the groin. echo was performed by the bedside which showed that the impella was too deep. md adjusted the impella but the sterile sleeve was torn in the process. md attempted to reattach the sleeve with sterile tegaderm, and the patient was prepped and transferred by careflite to another facility. the patient coded overnight. rn staff stated that patient was on rapid infusions and that patient possibly had coagulopathy disorder. rn reported that the patient was bleeding from all access line points, eyes, and nose. no further information or details were given. the patient expired after 11 hours of impella support. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation has been completed. as per the clinical information provided, the patient was transferred after impella insertion to another hospital and hematoma was noticed by the medical doctor while undraping. activated clotting time was said to be around 300, but the timing of the transfer is unclear. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient information was provided. clinical information mentioned that after the patient was moved to a different facility through helicopter, the medical doctor stated that the pump was seated deep which was discovered during echocardiogram. clinical information mentioned that while repositioning the pump, the sterile sleeve of the pump was accidently torn by the doctor which was later controlled by tegaderm. there is a lack of information on positioning of device prior to transfer and the exact timing of the transfer/discovery of improper positioning, as well as any patient conditions. the cause of the positioning issue was not determined due to lack of clinical details and no product was returned. the sterile sleeve damage occurred during the repositioning and thus is captured under positioning. b.2 required intervention was selected incorrectly on the initial manufacturer device repor number 1220648-2025-25342. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2025-25342 was submitted in accordance with updated procedures.
12465 | 1220648-2025-25343 | 21085717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21085717&pc=OZD | 12/19/2024 | 02/06/2025 |  | 01/06/2025 |  |  | 12/19/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582852 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Great Vessel Perforation | Intraoperative Cardiac Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 placement the left ventricle access was obtained but unable to pass the 5.5 due to the patient having a small anatomy. once the patient was in cardiothoracic intensive care unit following impella 5.5 attempt, echocardiography performed at bedside and tamponade noted. a decision was made to emergently bring patient to cardiovascular operating room (cvor) for tamponade and direct aortic approach of impella 5.5. once in cvor and sternotomy done, the patient was found to have wire injury in apex of the heart. repaired and direct 5.5 placed. the patient received multiple units of packed red blood cells, fresh frozen plasma and platelets throughout the procedure. the patient is stable and survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported vascular damage and delivery issues has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the root cause of the delivery issues was most likely patient condition as the patient had small vessels including a small aortic arch.
12466 | 1220648-2025-25344 | 21086717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21086717&pc=OZD | 12/17/2024 | 02/06/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578619 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. as impella explant was attempted, resistance was met at the 14f cook sheath. physician attempted to pull pump through the sheath at which point the motor housing snapped within the sheath. the physician stated that the patient had limited vascular options and that tortuosity of the axillary was likely cause. additionally, the patient bled from insertion site throughout access and multiple sheath exchanges. hematocrit was drawn and results were 14 leading to 2 units of packed red blood cells given. the physician removed the sheath entirely and exchanged for a short 14f sheath. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the access site bleed and the pigtail detachment of peripheral vessel has been completed. the impella cp and cook sheath were returned for evaluation. no abiomed 14 french x 25 centimeter (cm) introducer was returned for evaluation. upon visual inspection, the pump cannula was torn from the motor housing. the cannula and nitinol coil were stretched on the end of the cannula. the delo joint on the motor housing was fully intact. additionally, the pump catheter was torn at the 30 cm mark. the cook sheath was returned and had multiple kinks on the sheath body. clinical details noted that patient had limited vascular options and vessel tortuosity. additionally, it was noted that excessive force was applied when attempting to explant the pump through the cook sheath. cannula/pigtail detachment - peripheral vessel was investigated under product damage (detached inflow/outflow) - peripheral vessel as the pump cannula and inflow cage were found to be detached from the motor housing. the root cause of the product damage (detached inflow/outflow) was use related as the incorrect sheath was used and excessive force was applied. the root cause of the access site bleeding - major was patient condition related as patient had tortuous vasculature at access site. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25344 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12467 | 1220648-2025-25345 | 21086751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21086751&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578927 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support and peripheral veno-arterial extracorporeal membrane oxygenation (va ecmo) went to the cardiovascular operating room (cvor) for axillary impella 5.5 placement. the cp and peripheral va ecmo was going to be removed. upon arrival to cvor, a transesophageal echocardiogram was performed, and significant thrombus formation was noted in the left ventricle cavity. a decision was made to abort 5.5 placement and remain on ecmo and cp. the patient was not a candidate for lytic therapy and was made comfort care. over night the patient had a gastrointestinal bleed therefore the medical doctors and family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
12468 | 1220648-2025-25346 | 21086822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21086822&pc=OZD | 12/17/2024 | 01/31/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542910 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being delivered to support a patient admitted in with aortic stenosis and non-st segment elevation myocardial infarction. the 5.5 delivery failed despite all troubleshooting measures and instead the team placed a impella cp. no harm or injury was noted to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the severe calcification in the left subclavian artery and the insertion was aborted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12469 | 1220648-2025-25347 | 21087005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087005&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/06/2025 |  |  | 12/17/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580272 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and 10 minutes after the implant, the automated impella controller alarmed. ¿high purge pressure¿ followed by intermittent periods of ¿purge system blocked¿. unable to resolve with the usual troubleshooting. staff monitored and support continued. additionally, a pump stop occurred. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the high purge pressure and pump stop has been completed. the product was returned for investigation. no physical damage was observed on the purge cassette or pump. the purge cassette was successfully primed and connected to the pump, but high purge pressure was noted. pump passed the electrical testing. the pump was unable to start when placed in a bucket of water, and a significant amount of biomaterial was observed on the impeller and purge gap. upon further investigation, a blockage was identified in the motor housing, and it was confirmed that purge fluid escaped from the catheter after cutting near the motor housing. the data log shows a rise in purge pressure and low pump flow, with an active high purge pressure alarm after pump insertion. the purge pressure gradually returned to normal after 4 hours of the case run, and flow improved. however, the purge pressure gradually rose again after 2 days, followed by the pump stopping with a high motor current alarm and saturated pump flow. the cause of the high purge pressure issue was biomaterial, based on returned product and data log review. the cause of the pump stop issue was biomaterial, based on returned product and data log review.
12470 | 1220648-2025-25348 | 21087036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087036&pc=OZD | 11/25/2024 | 02/18/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533700 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, there was high purge pressure. tissue plasma activator was added to the purge. the issue resolved and support continued. the patient expired after 99.75 hours of impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the pump was returned for investigation with only the short-term log for the first 10 hours, which reported no abnormalities. high purge pressure issue reported on 4th day of pump use. the returned product showed no physical abnormalities, no biomaterial was observed on the impeller or purge gap, and the pump was primed successfully with no issues reproduced on purge pressure. as per the given clinical detail, tpa administration successfully resolved the high purge pressure. the cause of the high purge pressure issue was most likely biomaterial, based on return product investigation and provided clinical details. d9 date device returned to manufacturer added. b1 should have been both adverse event and product problem on manufacturer device report 1220648-2025-25348. e4 should have been left blank on manufacturer device report 1220648-2025-25348. h6 code 4582 was reported incorrectly on manufacturer device report 1220648-2025-25348. new code was added to health effect - clinical code.
12471 | 1220648-2025-25350 | 21087040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087040&pc=PYX | 12/18/2024 | 02/07/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025564022 |  | 1000323 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a rp flex placed to support a stemi patient. the pump stopped after just under 1 hour. the team attempted to restart and had concerns of clot ingestion. the pump was removed and replaced. | [5 empty] | The investigation for the pump stop has been completed. rp flex was returned for evaluation. upon visual inspection, no abnormalities were noted. no damage to pump observed. no biomaterial present on pump or impeller blade. pump passed electrical testing. pump was run in the lab for approximately 15 hours and pump stop did not reproduce. rt log of pump run shows multiple pump stops with alarm #13 for motor current high pump stops. pump is able to start back up immediately after pump stops without drop or loss of pump flow. purge parameters are not affected at time of pump stops. multiple pump stops occurred at the beginning of support then the pump was able to run for approximately 30 minutes before more pump stops occurred again. clinical details noted that patient's act was within range during support and pump was restarted multiple times during support. clinical team suspected clot ingestion, however, no specific signatures of clot ingestion were observed in case logs. as the pump was able to be restarted and run after a pump stop, the reported pump stops will be investigated under the failure mode of "temporary pump stop". the root cause of the temporary pump stops cannot be definitively determined as the pump stop did not reproduce when run in the lab and no specific ingestion signature was observed in the logs. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ corrections to the initial mfr report: d4 model number updated. d9-device available for eval changed to yes. g1 mdr reporting contact email. | The impella device was received from the customer on 30-dec-2024 therefore, an investigation of the device is underway, a supplemental mdr will be filed once the investigation is completed.
12472 | 1220648-2025-25349 | 21087041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087041&pc=OZD | 12/18/2024 | 01/29/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had frank hematuria. the pump was repositioned. the patient's hemoglobin and platelets were stable. the pump was explanted and the patient was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hematuria has been completed. the product was not returned for review. based on clinical description and data analysis, the root cause of hematuria was most likely due to pump was not in the proper position. the following have been reported incorrectly or missing from manufacturer device report.
12473 | 1220648-2025-25351 | 21087044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087044&pc=OZD | 12/18/2024 | 02/05/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553099 |  | 0048-0045 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. two hours after the patient arrived in intensive care unit, flows abruptly decreased on impella. waveforms were no longer correlated to aline and were uncoupled. intensivist performed bedside echo and visualized a large clot on inlet cage. pump was on p4 and getting constant suction alarm. pump remained on p4 and was eventually removed once act came down. no issues with removal, no clot visually seen on impella catheter. the doctor said it is possible there was a left ventricle thrombus he was unaware of prior to placing pump. | [5 empty] | The impella device was received from the customer on (b)(6) 2024, therefore, an investigation of the device is underway. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the returned compliant cp device visually inspected. no cannula kink or broken blade observed. no biomaterial observed. no low flow issue reproduced in investigation lab. data logs reviewed: no motor current spike observed. suction alarms occurred at the time of reported issue. low flow observed at p-4 right after suction alarm. the root cause of the low pump flow issue most likely was biomaterial ingest since clinical team stated on echocardiogram visualized a large clot on inlet cage. the following have been reported incorrectly on manufacturer device report 1220648-2025-25351 d4 model number d8 device available for evaluation ( the device was returned) e4 should have been left blank g1 reporting contact email.
12474 | 1220648-2025-25352 | 21087055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087055&pc=OZD | 12/18/2024 | 02/05/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515067 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support was taken to the catheterization lab for removal of impella as there was concern of clot on the inlet. trans-esophageal echocardiography did not reveal any thrombus. the patient is stable. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus has been completed. the pump was retuned for inspection and no abnormalities were noted upon investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. d.9 device return date/information was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25352 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2025-25352.
12475 | 1220648-2025-25353 | 21087211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087211&pc=OZD | 12/18/2024 | 01/27/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567184 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, an angiogram of leg showed occlusion of the common femoral artery. wire and balloon were introduced to the affected area restoring minimal flow to the distal extremity. vascular surgery was consulted for exploratory surgery of the right leg. fem/fem bypass occlusive right femoral artery. impella leg was cold purple without pulses. the right radial site had a hematoma and was also discolored without pulses. after readjusting perfusion sheaths the right leg had return of pulses. pump was explanted. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12476 | 1220648-2025-25354 | 21087219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087219&pc=OZD | 12/30/2024 | 01/31/2025 |  | 01/06/2025 |  |  | 12/18/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580445 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and after the pump was implanted, the pump was repositioned and after the patient went into ventricular tachycardia (vt). pump position was confirmed. support continued. additionally, the patient had a hematoma at the access site. staff monitored and anticoagulation was withheld. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported arrhythmia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: d4 updated model number.
12477 | 1220648-2025-25355 | 21087226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21087226&pc=OZD | 12/19/2024 | 01/30/2025 |  | 01/06/2025 |  |  | 12/19/2024 | 01/06/2025 |  |  |  |  | 01/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574877 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | Thrombosis/Thrombus | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported they were on impella 5.5 support and the patient had a right and left ventricle thrombus. when trying to wean extra corporeal membrane oxygenation, the patient would not wake up. the patient was taken for a computed tomography scan and the patient was found to have a hemorrhagic stroke. the family was made aware and care was withdrawn. the patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of stroke/cva and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25355. h6 codes 1770 and 4617 were reported incorrectly on manufacturer device report 1220648-2025-25355.
12478 | 1220648-2025-25291 | 21066222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21066222&pc=OZD | 12/12/2024 | 12/04/2025 |  | 01/03/2025 |  |  | 12/12/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684102 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review there is not enough evidence to exclude the automated impella controller (aic) as an associated factor in the patient's outcome. | Us complainant reported that a patient was transferred from (b)(6) hospital. the bedside team was changing over consoles to (b)(6) aic 10000. after transferring the alarm signaled for "air in purge line" when re-priming the purge cassette through the menu, the screen froze and did not advance. the impella was then transferred to a new console and was functioning appropriately. no patient harm has been reported. |  |  |  |  | B1 revised from the initial manufacturer device report number 1220648-2025-25291 to reflect both adverse event and product problem b2 revised from the initial manufacturer device report number 1220648-2025-25291 to reflect the patient outcome death and date of death added. b5 revised from the initial manufacturer device report number 1220648-2025-25291 to reflect the mso statement and patient outcome e4 should have been left blank on the initial manufacturer device report number 1220648-2025-25291 as it is unknown if the initial reporter also sent the report to the food and drug administration. h1 type of reportable event revised from the initial manufacturer device report number 1220648-2025-25291 to reflect the patient outcome death h5 selection was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-25291 h6 revised from the initial manufacturer device report number 1220648-2025-25291 to reflect the patient outcome death h6 codes 925 was reported incorrectly the initial manufacturer device report number 1220648-2025-25291. h8 usage of device selection was inadvertently entered incorrectly on the initial manufacturer device report number 1220648-2025-25291. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. data analysis: rlfs logs were acquired for (b)(6) and (b)(6). pump 501360 began on (b)(6). 2 days in, the pump was removed from the console and 90 seconds re-plugged in. at the same time, a purge pressure disc was acknowledged as inserted but there was no purge cassette read. one minute later the purge pressure disc was disconnected followed by an air in purge alarm. the pump was moved over to (b)(6) and then (b)(6) but each time, there was a purge disc detected but no purge cassette read. eventually, on (b)(6), the console is power cycled but then the purge disc and purge cassette are both read. device analysis: (b)(6) was returned as that was the console associated with the complaint in salesforce but performed as expected. root cause: there was no issue found with any of the consoles during the case. the air in purge alarm was triggered due to the purge disc already disconnected but the console still turning the spindle thus there would be no pressure differential. the same behavior across three consoles not reading the purge cassette indicates most likely not a console issue. the devices functioned/operated to specification. the investigation for the reported issue has been completed. h6 investigation type, findings and conclusions.
12479 | 1220648-2025-25292 | 21066330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21066330&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) had a broken purge clip. this issue occurred during use. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) purge disc/flag issues has been completed. console logs did not contain any data relevant to this complaint. the console was returned for evaluation. the reported broken pump connector cover was confirmed. the pump connector was also observed to be partially lodged inside the aic. aside from the reported damage to the pump connector cover and the observed partially lodged pump connector, no other issues were found with this console. the root cause of this issue was most likely use related. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12480 | 1220648-2025-25293 | 21066444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21066444&pc=OZD | 12/17/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 96 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump being placed to support a patient admitted in for a high risk percutaneous coronary intervention. the team placed the 14fr sheath, but prior to the delivery of the cp the team observed neurological issues / cerebrovascular accident. the cp placement was aborted. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12481 | 1220648-2025-25294 | 21066728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21066728&pc=OZD | 12/14/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/14/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp placement was aborted due to the patient¿s anatomy. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through the vasculature as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12482 | 1220648-2025-25295 | 21066899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21066899&pc=OZD | 12/14/2024 | 01/22/2025 |  | 01/03/2025 |  |  | 12/14/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571271 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant, the patient became agitated and was intubated. an echocardiogram was performed and a large mural clot was discovered in the patient's left ventricle. the patient was returned to the catheter lab and the impella cp was explanted in response to the condition. it was documented shortly after that the patient passed away. it is unknown if the complication contributed to the patient's passing. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12483 | 1220648-2025-25296 | 21067067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067067&pc=OZD | 12/15/2024 | 01/22/2025 |  | 01/03/2025 |  |  | 12/15/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534733 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. one unit of packed red blood cells was transfused and the doctor redid the stitch to the groin. the patient was on impella support for 202.15 hours before the patient expired. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2025-25296. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
12484 | 1220648-2025-25297 | 21067098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067098&pc=OZD | 12/15/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/15/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564790 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient lost distal pulses in their left side. as the patient was not a viable candidate for an external fem-fem bypass, the decision was made to explant the device once the planned percutaneous coronary intervention was completed. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12485 | 1220648-2025-25298 | 21067401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067401&pc=OZD | 12/10/2024 | 02/05/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546105 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Increase in Pressure |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. it was documented that during support, the purge flow suddenly decreased. a tissue plasminogen activator protocol was initiated to prevent a purge block. after six hours, the purge flow returned to normal. roughly six days later, the pump flow rate began to decline and hemolysis began to develop. attempts to reposition the pump were unsuccessful as placement into the mid left ventricle space proved difficult. due to the ongoing issue, the decision was made to explant the pump. following explant, a large clot was noted on the inlet. a new impella 5.5 pump was placed for ongoing support. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the hemolysis, low pump flow, positioning issue, and high purge pressure has been completed. the product was not returned. data logs are not available after the aic swap. review of the available logs showed no abnormalities that would aid in the investigation since the complaints. the root cause of the high purge pressure was most likely biomaterial as tpa was added to the purge and resolved the issue. clinical details mention pump positioning against the mitral apparatus on the day of the complaint as well as a clot on the explanted pump. it was noted that there were no anatomical abnormalities and no cannula kinks. the cause of the low pump flow was not determined since product was not returned and insufficient logs and clinical details were provided. the md attempted repositioning after the pump was found to be against mitral apparatus but the attempt failed. it was noted that there were no anatomical abnormalities and no cannula kinks. therefore, the root cause of the positioning issues was most likely use-related. clinical details mentioned pump positioning against the mitral apparatus on the day of the complaint as well as a clot on the explanted pump. however, the root cause of the hemolysis was not determined since product was not returned and insufficient logs and clinical details were provided.
12486 | 1220648-2025-25299 | 21067493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067493&pc=OZD | 12/14/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/14/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574184 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support and during support, the automatic impella controller (aic) had persistent "purge disc not detected" alarms. the purge tubing was replaced to no avail. the aic was replaced and support continued. | The complainant reported a patient had an impella 5.5 implant and during the implant, the patient had ventricular tachycardia (vt). when advancing the wire across the valve and into the lv, the patient had sustained vt requiring defibrillation at 200j once. the patient was prone to arrythmia. however, it is unknown if the impella implant contributed to the vt. |  |  |  |  | Udi# for the aic is unknown. device has been returned as of (b)(6) 2025, and evaluation is pending . once completed a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. b.1 revised selection to adverse event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. b.2 revised as a required intervention should of been selected on initial manufacturer device report number 1220648-2025-25299. b.5 revised describe event or problem as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. manufacturer device report number 1220648-2025-25362 was submitted to represent the automated impella controller event. d.1 revised brand name as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. d.4 revised section as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. d.9 revised as previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2025-25299 was submitted in accordance with updated procedures. h.1 revised selection to adverse event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. h.3 revised selections as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. h.5 revised as previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. h.6 codes were revised to reflect the correct information in b.5 as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. h.8 revised as previously entered incorrectly on initial manufacturer device report number 1220648-2025-25299. h.10. the pump was not returned. therefore, the additional manufacturer narrative on initial manufacturer device report number 1220648-2025-25299 was incorrect to say that the device had been returned.
12487 | 1220648-2025-25300 | 21067732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067732&pc=OZD | 12/15/2024 | 03/11/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581012 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan had a cp support ongoing for a patient admitted in with worsening heart failure and st elevated myocardial infarction. the limb became ischemic from the access site/pump location. the harm/injury prompted the team to explant the pump. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e1 initial reporter name prefix was inadvertently omitted on the initial manufacturer device report number 1220648-2025-25300.
12488 | 1220648-2025-25301 | 21067757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21067757&pc=OZD | 12/17/2024 | 01/22/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577951 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Bradycardia | Ventricular Fibrillation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient went into ventricular fibrillation. the patient was defibrillated once and returned to bradycardia with heart block. arrhythmia occurred an hour after repositioning the impella. it was noted that arrhythmia was unrelated to the impella cp. care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported atrial arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the atrial arrhythmia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12489 | 1220648-2025-25302 | 21068078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068078&pc=OZD | 12/16/2024 | 01/30/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Defective Device | [6 empty] | Hemolysis | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, labs were hemolyzed and red tinged urine was present in the foley bag despite impella positioning in left ventricle. continuous renal replacement therapy (crrt) was started. additionally, the plastic tightener on 14f dilator was removed from 14f dilator and left attached to 14f peel away sheath. plastic ring was secured using clamps then cut off from around the cp catheter without any complications. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and introducer damage has been completed. pump was not returned for investigation. data logs were investigated and at the time of noted hemolysis, there was no suction or motor current spike. placement signal was normal. therefore, the root cause of the hemolysis issue was not determined since product was not returned. clinical information mentions that the plastic ring of the dilator was stuck in the sheath but without the product back the failure cannot be investigated. therefore, the root cause of the introducer damage issue was not determined since product was not returned.
12490 | 1220648-2025-25303 | 21068583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068583&pc=OZD | 12/16/2024 | 02/06/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570155 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, the patient's blood pressure plummeted. the staff suspected a perforation, which was confirmed at the aortic root. the doctor reported that the rupture/perforation was possibly caused by impella as the aorta was very frail. the aortic root has been replaced surgically with graft and the patient was stable after complication management. pump was explanted. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12491 | 1220648-2025-25304 | 21068657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068657&pc=OZD | 12/16/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was being prepped for an impella cp implant and during priming, d5 was leaking out of the purge filter. a crack was noted in the filter. the staff decided to not use this impella and use a new pump instead. no patient harm has been reported. | [5 empty] | The investigation of purge leak ¿ pump has been completed. no product was returned for evaluation. clinical details noted that a leak was present at the purge filter during pump priming. the root cause of the purge leak was most likely a damaged sidearm but the exact location of the leak could not be confirmed. e4 should have been left blank on manufacturer device report 1220648-2025-25304. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿.
12492 | 1220648-2025-25305 | 21068664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068664&pc=OZD | 12/16/2024 | 02/12/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that on day two of impella cp insertion the patient¿s labs were coming back hemolyzed. the plan was to decrease the p-level to see if the hemolysis clears but the next day the patient decompensated and expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Complaint device was not received for analysis. device history record review confirmed the pump passed all post sterile inspection checks. data logs were reviewed: no motor current spike outside of p-level changes observed. suctions occurred. many impella unknown position alarms observed. no clear indications observed from reviewing the logs to confirm the root cause of the reported issue. the root cause of the reported issue could not be determined since the complaint device was not returned for analysis and insufficient clinical data was provided.
12493 | 1220648-2025-25306 | 21068672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068672&pc=PYX | 12/17/2024 | 01/03/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  | 01/31/2025 |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305910 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The us complainant had a rp placed to support a patient admitted in with right heart failure and pulmonary embolism. the pump was placed but there were issues, the first being purge pressure alarms that resolved with purge cassette replacement. additionally, there was bleeding observed which prompted the team to apply pressure and not restart any heparin in the infusion. the team also observed signs of hemolysis. the team unsure if the urine color change was from hemolysis or traumatic foley. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp with smart assist system with automated impella controller. section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
12494 | 1220648-2025-25307 | 21068852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21068852&pc=OZD | 12/16/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572841 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a cp pump to support a patient with history of coronary artery bypass graft and suffering from ventricular tachycardia in the hospital. the acute myocardial infarction / cardiogenic shock patient did have the cp explanted and escalated to the 5.5 pump when transferred to the tertiary center. the patient had signs of hemolysis and limb ischemia. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis and limb ischemia has been completed. data logs showed pump ran without issue during support. there were no alarms triggered during the case. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of the limb ischemia could not be determined without additional information.
12495 | 1220648-2025-25308 | 21069088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21069088&pc=OZD | 12/17/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577173 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support explanted due to access site bleeding. a large hematoma was noted on the impella side. the patient was taken to the operating room for removal but the patient crashed and placed on venoarterial extracorporeal membrane oxygenation support. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
12496 | 1220648-2025-25309 | 21069260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21069260&pc=OZD | 12/10/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/10/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505442 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and while on support, the patient was found to be heparin-induced thrombocytopenia positive. all heparin was stopped and 50 milliequivalent of sodium bicarbonate was run through the purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. h.5 revised as yes selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25309.
12497 | 1220648-2025-25310 | 21069494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21069494&pc=OZD | 12/15/2024 | 09/24/2025 |  | 01/03/2025 |  |  | 12/15/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578462 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella cp support. while on support, placement signal low alarm was reported. the fellow did an echocardiogram and agreed that the impella was too deep. to achieve better flows, the device was repositioned multiple times using echocardiogram guidance on the unit and in the catheterization lab. however, desired flows could not be achieved, even with optimal positioning. consequently, the decision was made to explant the device, reassess, and replace it with a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were investigated and "placement signal low" alarms were observed along with a trending placement signal. the 5s parameters were in specification. clinical updates mention that patient had volume issues and was given volume in the afternoon. therefore, the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-25310. d10 concomitant medical products and therapy dates updated. | The investigation has been completed. investigation of positioning issues was changed to low pump flow since the device was reported to be repositioned in order to improve flows. the pump was not returned for investigation. as per the clinical information provided, there were suction alarms and "impella position unknown" alarms while led to repositioning of pump multiple times, both with and without echo. the attempts were unsuccessful and therefore the root cause of the positioning issues which led to low pump flow was most likely patient condition related due to the trending placement signal and "placement signal low" alarms. h.6 medical device problem code was changed since the initial manufacturer device report number 1220648-2025-25310 was submitted due to investigation results.
12498 | 1220648-2025-25311 | 21069709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21069709&pc=OZD | 02/10/2022 | 01/03/2025 |  | 01/03/2025 |  |  | 02/10/2022 | 01/03/2025 |  |  | 04/30/2022 |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1475146 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was admitted to the hospital with complaints of shortness of breath, elevated troponin and in st elevation myocardial infarction. in the cardiac catheterization lab, the patient was found to have occlusions in both left anterior descending (lad) and right coronary (rca) arteries. the patient was urgently intubated due to shortness of breath and difficulty laying flat. an impella cp was implanted for mechanical circulatory support, and the patient had stents placed in their lad and rca. overnight, the patient decompensated and required an increase of pressor support. the medical team decided to escalate the patient support to an impella 5.0. during impella 5.0 implant, the medical team experienced difficulty advancing the impella 5.0 past the iliac. a peripheral balloon was used, along with viperglide to lubricate the sheath and graft. the impella 5.0 was then successfully delivered into the left ventricle. on (b)(6) 2022, it was reported a large amount of bleeding was found coming from the jackson-pratt (jp) drain at the impella insertion site. heparin was decreased, and six units of blood were transfused to the patient. hemostasis was unable to be achieved, and the patient was taken for a computed tomography scan to assess the source and the extent of the bleeding. a fem stop was placed at the access site, however oozing persisted and the patient required an additional 4 units of blood to be transfused. there was no report that hemostasis was ever achieved. it was additionally noted that the patient expired on support the following day. there is not enough evidence to exclude impella as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12499 | 1220648-2025-25312 | 21070956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21070956&pc=OZD | 12/15/2024 | 01/13/2025 |  | 01/03/2025 |  |  | 12/15/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp removal, the perclose got caught on the back wall and a cutdown was performed. the patient is stable and the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the vascular damage has been completed. no product was returned for analysis. clinical details do not mention any excessive force or patient anatomy issues. the root cause of the vascular damage - peripheral vessel was not determined as no product was returned and limited clinical information was provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12500 | 1220648-2025-25313 | 21071903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21071903&pc=OZD | 12/16/2024 | 02/04/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579154 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant, an impella 5.5 pump began to experience saturated flows. the pump was replaced with a second impella 5.5 pump for ongoing patient support. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of saturated flow has been completed. according to the data and device analysis the root cause of the saturated flow was due to biomaterial. d9 date device returned to manufacturer added. h6 code 2993 was reported incorrectly on manufacturer device report. new code was added to medical device problem code.
12501 | 1220648-2025-25314 | 21072034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21072034&pc=OZD | 12/16/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump for a patient with cardiomyopathy and was awaiting transplant. the pump had been supporting for 41 days when there was high purge pressure. the pump remained on for support despite the purge pressure/purge blockage with tissue plasimongen activator. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. based on data logs, the root cause of the high purge pressure was determined to most likely be a use issue. d6b was omitted from manufacturer device report 1220648-2025-25314. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12502 | 1220648-2025-25315 | 21073031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21073031&pc=OZD | 12/05/2024 | 06/23/2025 |  | 01/03/2025 |  |  | 12/05/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569582 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown ethnicity) with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and was escalated to an impella 5.5 device the following day. while on impella 5.5 mcs support, sterile sleeve separation was discovered. the patient had multiple episodes of projectile vomiting and when changing the axillary site dressing, day 16th of support, condensation and particles of emesis were noted inside the sterile sleeve. the sterile sleeve appeared to have broken off the distal touchy borst valve. no further details were provided. impella 5.5 device continues to support the patient, and there has been no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer as the device remains implanted, supporting the patient at the time of this medwatch writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (sleeve damage) has been completed. on day 16 of support, pt had multiple episodes of projectile vomiting. whilst changing his axillary site dressing, rn noticed there was condensation and particles of emesis inside the sterile sleeve. the sleeve appears to have broken off the distal tb valve and vomiting noted inside sterile sleeve of 5.5. the pump was returned where damage was noted to the sterile sleeve as it was separated from the anchoring ring used to secured the sleeve along the catheter. the cause of the product damage (sleeve damage) was not determined as not enough clinical details were provided to determine the mechanism in which the anchoring ring became disconnected from the sterile sleeve. d6b added. d9 device returned to manufacturer date added.
12503 | 1220648-2025-25316 | 21074076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21074076&pc=OZD | 12/13/2024 | 02/19/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536655 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump experienced an unexpected stoppage during patient support. an increase in motor current was noted leading up to the event. the pump was replaced due to the noted failure. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. there were no data logs returned for this case. the pump was looked at in the lab and there was tegaderm wrapped around the sidearm. there was no biomaterial seen or any abnormalities with the pump. the pump was soaked to remove the dried dextrose. it was noted that the purge gap was large, and the pump was unable to be run. the root cause of the pump stop was due to bearing wear as it was used for 20.79 days and the intended design life is 8 days according to the design trace matrix. d9 date device returned to manufacturer added. b2 should have been left blank on manufacturer device report. b3 date of event was inadvertently reported incorrectly. d4 model number was inadvertently reported incorrectly. e4 should have been left blank on manufacturer device report. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 code 4868 was reported incorrectly on manufacturer device report. new code was added to health effect - clinical code.
12504 | 1220648-2025-25317 | 21074079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21074079&pc=PYX | 12/16/2024 | 02/04/2025 |  | 01/03/2025 |  |  | 12/16/2024 | 01/03/2025 |  |  | 02/28/2026 |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025518053 |  | 1000323 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Sense | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a rp flex placed to support a patient admitted in for mvr/cabg and found to have rv dysfunction post procedure. the pump was explanted after 3 days due to signal loss and low flows. the pump was not replaced after explant. | [5 empty] | The impella device was received from the customer on 23-dec-2024 , therefore, an investigation of the device is underway, a supplemental mdr will be filed once completed. | The investigation of low pump flow and optical signal issue has been completed. the "optical signal issue" failure mode was changed to "placement signal issue" because the complaint was in relation to the dp sensor. the returned pump was inspected and there were no physical abnormalities. there was no biomaterial in the inflow cage. the dp sensor was imaged and had no abnormalities. the pump was run in the lab and the low pump flow issue did not reproduce. the data logs from the case confirm low flow for matching p-level with pump flows drifting down at p6 and placement signal drifting upwards. the pump was run in the flow loop test and there was dp sensor drift. the root cause of the placement signal issue was determined to be due to a sensor drift. the failure mode "low pump flow" was deleted as this was in reference to the unreliable calculated flow rate due to the sensor drift. the following have been reported incorrectly or missing from manufacturer device report.
12505 | 1220648-2025-25318 | 21074081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21074081&pc=OZD | 12/17/2024 | 02/03/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580047 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the uk had a cp pump placed to support a patient admitted in with an st elevated myocardial infarction. the patient's limb became ischemic but, due to the inability to remove the device due to patient condition, the pump remains on for support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added e4 should have been left blank on manufacturer device report 1220648-2025-25318.
12506 | 1220648-2025-25320 | 21074083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21074083&pc=OZD | 12/16/2024 | 01/31/2025 |  | 01/03/2025 |  |  | 12/18/2024 | 01/03/2025 |  |  |  |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569115 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a hematoma and experienced some numbness around their impella 5.5 access site. no intervention was required and support with the implanted pump was maintained. roughly 17 days into support, the pump began to experience suction issues. volume was given and the device was repositioned, but the issue persisted. as the patient was beginning to become overloaded, additional volume could not be added and the decision was made to explant the pump. upon removal, biomaterial was seen in the outlet of the pump. the patient survived the event. | [5 empty] | The investigation for the low pump flow has been completed. the returned product was inspected and biomaterial was observed around the impeller. the pump was run in the lab after clearing and the low pump flow issue did not reproduce. the data logs do not show any motor current spikes but show suction alarms as reported. the root cause of low pump flow was determined to be biomaterial. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12507 | 1220648-2025-25319 | 21074084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21074084&pc=PYX | 12/17/2024 | 01/23/2025 |  | 01/03/2025 |  |  | 12/17/2024 | 01/03/2025 |  |  | 03/31/2025 |  | 01/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024390035 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support. while on support, rp flex was noted to be in suboptimal position on daily chest x-ray, trans-esophageal echo (tee) demonstrated outflow in right ventricle. the patient was brought to the catheterization lab and device attempted to be repositioned. after multiple unsuccessful attempts the doctor decided to remove and replace rp flex impella. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12508 | MW5164077 | 21053739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21053739&pc=PYX | 12/07/2024 | 12/23/2024 |  | 01/02/2025 |  |  |  | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP SMART- ASSIST CATHETER |  |  | PYX |  |  | 1000323 |  |  | ABIOMED, INC. |  |  | US |  | 2025551163 |  |  |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 82 YR |  | Male | 77 KG | Non Hispanic | White | Death | No | Yes | [7 empty] | Insufficient Information | [12 empty] | Death | Other | Life Threatening | Required Intervention | [25 empty] | Initial submission | [9 empty] | Impella console display unable to receive placement signal from impella catheter due to faulty catheter.
12509 | 1220648-2025-25272 | 21054087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21054087&pc=OZD | 12/12/2024 | 02/04/2025 |  | 01/02/2025 |  |  | 12/12/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579077 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during placement, difficulty with the pump and guidewire was encountered, followed by a stroke with a demised outcome. the pump was unable to pass on .018 guidewire. the physician stated the guidewire was too lax and buckling under pressure of the pump advancement. the decision was made to use .027 wire from an impella rp flex kit, and the pump successfully advanced. next day after start of support, it was noted the patient experienced a bilateral acute middle cerebral artery (mca) stroke. prognosis and change in neurologic status were indicated incompatible with recovery. care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the delivery issues and the thrombus has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the guidewire was unable to pass through the vasculature. the root cause of the thrombus could not be determined without additional information.
12510 | 1220648-2025-25273 | 21054394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21054394&pc=OZD | 12/13/2024 | 02/05/2025 |  | 01/02/2025 |  |  | 12/13/2024 | 01/02/2025 |  |  | 10/31/2025 |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420157 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day three of impella cp support, the patient¿s leg was amputated as the peel-away sheath was unable to be removed. the patient remained stable and the procedural outcome was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
12511 | 1220648-2025-25274 | 21054530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21054530&pc=OZD | 03/06/2024 | 02/04/2025 |  | 01/02/2025 |  |  | 12/13/2024 | 01/02/2025 |  |  | 12/31/2025 |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436368 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured gastrointestinal (gi) bleed with a "possible" relationship to the impella 5.5 device used: ¿3/6. large blood bm x 2 with hypotension, heparin stopped, vasopressin initiated along with blood transfusions 3/7. additional gi bleed 7 units of prbcs and 1-unit ffp¿. the patient recovered with no sequelae. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. no product was returned. the cause of the vascular damage- peripheral vessel issue was not determined since the product was not returned, unknown relation to impella with gastrointestinal bleed being observed and due to insufficient information. g.2 revised as study was missing from the initial manufacturer device report number 1220648-2025-25274.
12512 | 1220648-2025-25276 | 21055334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21055334&pc=OZD | 12/14/2024 | 02/07/2025 |  | 01/02/2025 |  |  | 12/14/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565144 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that on day three of the patient being on impella cp support a low-pressure alarm appeared in the flushing system. the cassette was replaced but the alarm did not resolve. a leak from the side arm of the flushing line was observed. the impella was removed. the patient is stable. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The pump was returned for investigation, but the data log was not provided. no physical damage was observed on the returned product. when the pump was started and primed, purge leak was observed coming from a hole near the 90cm mark of the catheter, most likely caused by the suturing near the repo hub. no leak was observed from sidearm. the cause of the purge leak was damaged purge tubing with catheter punctured at 90 cm marking.
12513 | 1220648-2025-25277 | 21055340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21055340&pc=OZD | 12/15/2024 | 02/03/2025 |  | 01/02/2025 |  |  | 12/15/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient decompensated and needed emergent stenting of left main/ left anterior descending artery / circumflex artery. the surgeon recommended escalation to impella 5.5. on arrival to intensive care unit, the impella cp was noted to be pulled back from ventricle. severe bleed noted to femoral access site. impella was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. a.1 revised patient identifier as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-25277. b.6 should of been left blank on the initial manufacturer device report number 1220648-2025-25277 as it belongs in b.7. information was added to b.7 d.4 revised serial number, lot number, expiration date and unique identifier (udi) # were previously entered incorrectly on initial manufacturer device report number 1220648-2025-25277. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25277 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-25277.
12514 | 1220648-2025-25278 | 21055368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21055368&pc=OZD | 12/16/2024 | 02/05/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578372 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in acute myocardial infarction / cardiogenic shock. the pump was placed but there were persistent hemolysis concerns. the urine color had changed and despite repositioning the signs persisted. the pump remained on with the hemolysis until explant and transition to an intra-aortic balloon pump. the patient¿s outcome is unknown. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. the returned impella cp pump was visually inspected. biomaterial was observed around the impeller and the purge gap. no broken blade or sharp edge was observed. data logs were reviewed. no motoe current spike was observed outside of the p-level changes. many suctions occurred. quick resolved impella position in ventricle alarms was observed and hemolysis issue reported the next day. the pump was repositioned under echocardiogram and hemolysis improved. the patient is coagulopathic. it looks like the red urine is related to coagulopathic condition of the patient. biomaterial was ingested around impeller during the support. the root cause of the hemolysis issue most likely was improper positioning related since the hemolysis reportedly resolved after repositioning. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2025-25278.
12515 | 1220648-2025-25279 | 21055550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21055550&pc=OZD | 12/14/2024 | 01/31/2025 |  | 01/02/2025 |  |  | 12/14/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on the same day of impella cp insertion some oozing was noted at the patient¿s groin site. manual pressure was held, femstop, and suture was added. the bleeding was resolved. additionally, the patient following a coughing episode had positioning issues, followed by purge pressure high alarm with flows around 1.5ml/hr. during this time the patient also experienced stroke like symptoms. computed tomography scan was negative, and symptoms improved within an hour. it was further reported that care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and high purge pressure has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the high purge pressure issue was most likely biomaterial, based on data log review.
12516 | 1220648-2025-25280 | 21055717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21055717&pc=OZD | 12/15/2024 | 01/23/2025 |  | 01/02/2025 |  |  | 12/15/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572785 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was going to be taken to the operation room for a cut down and explant of cp due to left ventricle thrombus that was found after the impella was placed. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
12517 | 1220648-2025-25281 | 21056824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21056824&pc=OZD | 12/16/2024 | 02/14/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558954 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 92 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an impella cp implant and the pump kinked in the left ventricle. the pump was explanted and a new cp was implanted. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the product damage has been completed. returned complaint cp pump was visually inspected. cannula kink observed near outlet. no other abnormality found. the root cause of product damage (cannula kink) was use related due to improper technique. added h6 impact code 4629.
12518 | 1220648-2025-25282 | 21057251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21057251&pc=OZD | 12/12/2024 | 01/29/2025 |  | 01/02/2025 |  |  | 12/12/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577567 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient. after a day of support, which included and access site bleed, the team made choice to explant and escalate to a 5.5 pump. after explant and groin closure the team discovered a vessel damage that prompted the team to intervene and repair the vessel. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the vascular damage has been completed. the pump was not returned for investigation, and the data log was not required for this case. additionally, the left femoral artery was repaired after it was dissected during the removal of the pump. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: b2-selected death and added date of death. b5-patient outcome and mso review. g1 mdr reporting contact email. h1-changed to death. h6-impact code 1802.
12519 | 1220648-2025-25283 | 21057348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21057348&pc=OZD | 12/16/2024 | 01/29/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1701689 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller error and a battery failure alarm. the aic was removed from service. there was no patient involvement. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) battery failure red alarm and yellow alarm has been completed. console logs from the reported complaint date of 2024.12.16 are consistent with the complaint. the aic issued multiple battery failure (transport connection blown/missing) [event set #360] and controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) [event set #24] alarms on the complaint date. the console was returned for evaluation. both the controller error and the battery failure were able to be reproduced. the results of running the batttest utility on telnet revealed that the cell stacks in both batteries had voltages of less than 2.5v which resulted in battery lockout. a misaligned communication ribbon cable on the kbd was also discovered after inspecting the internal components of the console. the reported controller error was determined to be a loss of communication issue between the kbd and 4-in-1 due to misaligned communication ribbon cable on the kbd. the communication errors resolved after reseating the ribbon cable. the reported battery failure issue was determined to be caused by depleted batteries resulting in battery lockout. console logs from the reported complaint date of 2024.12.16 are consistent with the complaint. the aic issued multiple battery failure (transport connection blown/missing) [event set #360] and controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) [event set #24] alarms on the complaint date. the reported controller error was determined to be a loss of communication issue between the kbd and 4-in-1 due to misaligned communication ribbon cable on the kbd. the communication errors resolved after reseating the ribbon cable. the reported battery failure issue was determined to be caused by depleted batteries resulting in battery lockout. added d9 device return date corrections to the initial mfr report: h5 corrected to no. h8 corrected to reuse.
12520 | 1220648-2025-25284 | 21057363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21057363&pc=OZD | 12/09/2024 | 02/21/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  | 12/31/2025 |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025470423 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan reported that the patient on day one of impella cp support developed severe lower limb ischemia. the impella was removed. it was also reported that patient care was withdrawn and the patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.2 revised as required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25284. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25284 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25284. h.6 component code 925 was reported incorrectly on the initial manufacturer device report number 1220648-2025-25284 and a new code was added.
12521 | 1220648-2025-25286 | 21057368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21057368&pc=OZD | 12/04/2024 | 02/03/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025559525 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification via abiomed's long-term outcome and quality indicator (loqi) impella registry was received regarding a patient who developed a right chest hematoma with a definitive relationship to the impella device. it was documented that following implant, a hematoma developed over the right axillary pocket coinciding with complaints of pain. the patient was brought to the operating room for exploration and washout of the right axillary pocket and hematoma. a jackson-pratt drain was also placed to manage the hematoma and the condition resolved. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25286 and was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-25286 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 study selection was inadvertently omitted from the initial manufacturer device report number 1220648-2025-25286 and was added.
12522 | 1220648-2025-25285 | 21057369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21057369&pc=OZD | 12/16/2024 | 02/06/2025 |  | 01/02/2025 |  |  | 12/16/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565319 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient developed hemolysis. a decision was made to explant impella due to elevated creatinine. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction : ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12523 | 1220648-2025-25287 | 21059293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21059293&pc=OZD | 12/15/2024 | 02/11/2025 |  | 01/02/2025 |  |  | 12/15/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1334194 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and during support, the pump stopped with automated impella controller (aic) failure alarms. prior to the pump stop, the patient was in the process of being weaned off the impella 5.5 mcs and was down to support level p-3 on day 18 of support with plans for removal of the device later in the said week. however, the following morning, a pump stop alarm with aic failure alarms was triggered. the physician along with the heart failure team decided that there was time to obtain another aic and restart the pump after the pump stop troubleshooting. therefore, the impella 5.5 pump was pulled to the ascending aorta, and cardiothoracic surgery was called for removal of the impella 5.5 device. it was reported that the pump removal went well and the patient remained stable during the event, and the patient was noted to be stable, as well, following the event. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two medical device reports associated with this event. refer to manufacturing report number 1220648-2024-24878 for the associated impella 5.5 pump. | The investigation of automated impella controller (aic) - pump stop has been completed. the pump was being used continuously for 18 days, and then on the complaint date, the console displayed ¿impella stopped (mechanical/electrical issue) ¿ alarm¿. this confirmed the reported failure mode. the real time log showed that the pump flow dropped to zero, and the motor current spiked above 30000. though the user tried restarting the pump several times, it was unsuccessful. this console was tested and was unable to reproduce the reported failure mode upon running the optical test pump for four hours. no hardware issue related to the reported failure mode was found on the console.
12524 | 1220648-2025-25288 | 21060733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21060733&pc=OZD | 05/16/2022 | 01/02/2025 |  | 01/02/2025 |  |  | 05/17/2022 | 01/02/2025 |  |  | 01/31/2024 |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173660 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the impella stopped suddenly with high motor current on the 22nd day of support. the clinicians attempted to restart the pump three times, however their efforts were unsuccessful. resultantly, the pump was explanted and replaced with another impella cp. no patient harm was reported due to the malfunction. | [5 empty] | Data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. the pump was unable to be restarted, with each restart attempt resulting a motor current spike to 1500. an enlarged impeller purge gap was observed on the returned pump. no other external damage or abnormalities were found. in conclusion, it was determined that the cause of the pump stop was bearing wear after 22 days of use, which is beyond design life. this report is being filed as part of a retrospective review of historical records.
12525 | 1220648-2025-25275 | 21060760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21060760&pc=OZD | 01/28/2024 | 10/20/2025 |  | 01/02/2025 |  |  | 02/08/2024 | 01/02/2025 |  |  | 10/31/2025 |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420730 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Bradycardia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella cp support. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right leg ischemia. distal pulses were diminished on (b)(6) 2024. vascular ultrasound ((b)(6) 2024) concerning for proximal stenosis. the patient underwent vascular surgery on (b)(6) 2024 for embolectomy of the right external iliac artery. operative findings note that the perclose device captured the posterior wall causing occlusion of the iliac artery. distal pulses were recovered. additionally, the patient that endured symptomatic bradycardia with a "probable" relationship to the impella device. after arrival to the cardiac intensive care unit post-impella cp procedure, the patient developed worsening bradycardia. impella suck-down events to reduce power to p5. bedside echo reveals improper impella placement, which was repositioned. atropine was also given, which improved hemodynamics. amiodarone and lidocaine infusions were stopped. impella cp powered up to p7 after repositioning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Instructions for use as follows: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d6b was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25275. g3 date received by manufacturer was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25275, the correct date should have been 23dec2024. h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2025-25275. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned, and logs are not applicable. limb ischemia - reversible limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations thereof detailed description of the symptoms experienced by the patient physical examination findings, including pulse assessment and visual examination of the affected limb diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography laboratory test results, including blood tests for markers of inflammation or clotting disorders any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. vf/vt/af/at cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. optical signal issue the fm optical signal issue was removed since there were no indications of optical-related issues and the incorrect positioning was resolved with standard troubleshooting. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include optical signal coding. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
12526 | 1220648-2025-25289 | 21061304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21061304&pc=OZD | 12/14/2024 | 02/14/2025 |  | 01/02/2025 |  |  | 12/14/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573700 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had a cp pump support initiated for a patient admitted in acute myocardial infarction/cardiogenic shock. the pump was supporting when the team observed signs of hemolysis. the patient expired with care withdrawal prior to any intervention for the hemolysis. the support had been for 37 hours. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the patient presented signs of hemolysis. the patient had decreased urine production and the color was noted to be changed. the team discussed care escalation and swan placement. swan placement would have allowed for better hemodynamic monitoring. neither of these measures (escalate/swan) were taken. the pump was repositioned as the pump was very close to the mitral. no other interventions were noted for the hemolysis. it was noted that care was withdrawn and the patient expired after 37 hours of support and without intervention taken for the hemolysis. |  |  |  |  | The impella device was not received from the customer, therefore, an investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed the pump passed all post sterile inspection checks. the device was not received for evaluation. data logs were reviewed. no motor current spike was observed. impella position in ventricle alarmed at p-2 and end of the case. no clear evidence was observed from reviewing the logs. as clinical details stated that the device was very close to the mitral valve and left shallow in attempt to free pump from papillary muscle, the root cause of the hemolysis was most likely improper positioning due to improper technique. corrected information provided in b5 and h10. | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue most likely was improper positioning due to improper technique. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25289. g1 manufacturing site fax number was inadvertently reported incorrectly on manufacturer device report 1220648-2025-25289.
12527 | 1220648-2025-25290 | 21063177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21063177&pc=OZD | 12/15/2024 | 01/22/2025 |  | 01/02/2025 |  |  | 12/15/2024 | 01/02/2025 |  |  |  |  | 01/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523179 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and had low pump flow and would ultimately expire. the patient arrived at the emergency department and fell while ambulating in the parking lot. emergency room staff started cardiopulmonary resuscitation (cpr) and obtained return of spontaneous circulation. the patient coded again, st-elevation myocardial infarction was called, and the patient was taken to the catheterization lab for percutaneous coronary intervention which was performed to the right coronary artery. the impella cp device was then placed via the right femoral artery. a temporary transvenous pacer was placed as well. the patient was transported to the unit on mcs. it was noted the patient continued to code multiple times from arrival to through the icu admission. impella was turned to support level p-0 for approximately one minute as the team made the decision to pronounce the patient. the patient began spontaneously breathing, and the impella cp was restarted with cpr efforts. the nurse stated the was still being coded and suction alarms remained, possibly related to a clot. suspicion of clot was noted related to spike in motor current when impella was restarted. patient eventually succumbed to cardiovascular demise; care was withdrawn after approximately nine hours of mcs. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed low flow occurred after pump started that triggered auto suction response and suction alarms. next, pump was run with low flow for 6 hours and stopped manually. pump then was restarted with motor current (mc) spiked followed flow dropped to 0.5 l/min. this event remained to the rest of the case. product was not returned for review. patient was deteriorating the entire case with cpr required. data logs show flows are on the lower side the entire case and trending downwards while on p-6. suction alarms are continuous throughout. the mc spike occurs at the end of the case with an additional drop in flow. the overall cause of the low flows was most likely patient condition related. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12528 | 1220648-2024-25214 | 21046672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21046672&pc=OZD | 10/07/2022 | 12/30/2024 |  | 12/31/2024 |  |  | 10/07/2022 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023247447 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support during a ventricular tachycardia ablation. it was reported while on impella 5.5 support, the automated impella controller (aic) became damaged during a cassette change. it was reported a piece of the aic's air sensor was broken off, and the aic required replacement with the back up aic unit to resolve the issue. there was no patient harm reported. | [5 empty] | Data analysis: the automated impella controller (console) logs were looked but they were found to be not relevant with the failure mode. device analysis: the console was tested after arriving in field service. the purge disc retainer was confirmed to be broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
12529 | 1220648-2024-25213 | 21046677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21046677&pc=OZD | 08/16/2021 | 12/30/2024 |  | 12/31/2024 |  |  | 08/16/2021 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1467028 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that while on impella 5.5 support, the automated impella controller (aic) would not recognize the purge cassette after it was replaced. two cassettes were used to attempt to resolve the issue, however the issue was only resolved after the aic was replaced with another unit. there was no patient harm reported. | [5 empty] | Data analysis: automated impella controller (console) logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was tested after arriving in field service. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). root cause of the broken purge flag is due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
12530 | 1220648-2024-25217 | 21046934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21046934&pc=OZD | 12/11/2024 | 01/27/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570062 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed for over two weeks when a pump stop occurred. multiple attempts were made to start the cp at different p levels without success. motor current spike at start over 1400ma and the impella did not start. the cp was explanted. the patient is stable without impella support and survived the event. | [5 empty] | The root cause of the pump stop has been completed. no product was returned for analysis. the logs confirm pump was run longer than indicated. the pump was run for 16 days, long past the intended design life of 5 days for european cases. the root cause of the pump stop was most likely due to bearing wear. h.6 code 4115 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-25217. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-25217 stated that the device was discardd but the device was not returned. type of investigation code has been updated to reflect not returned. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12531 | 1220648-2024-25218 | 21046998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21046998&pc=OZD | 12/11/2024 | 02/12/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528382 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant in japan reported the patient was on impella cp smart assist support and during support, the patient had mitral regurgitation. staff stated that it is possible that the pigtail was entangled in the papillary muscle. the pump was removed. the patient survived the event. | [5 empty] | The investigation for the heart valve damage has been completed. the pump was not returned for investigation. data logs were were not returned as well. as per the clinical information provided, the pump was mispositioned where the valve was touching the tip of the impella and simultaneously sucked in the mitral valve worsening the regurgitation. therefore, the root cause of the heart valve damage issue was most likely improper positioning of the device due to improper technique. as noted in the impella instructions for use: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g.2 revised as foreign selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25218. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25218. h.10 some of the instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2024-25218 and were added. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
12532 | 1220648-2024-25219 | 21047031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047031&pc=OZD | 12/11/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, a yellow controller error displayed on the automated impella controller (aic). no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller error-yellow alarm has been completed. console logs from the reported event date align with the complaint, showing a ¿controller error (#1045: opm comm crc invalid)" was triggered during the clinical procedure. this error occurred before the console was shut down and replaced. rt log data revealed that the controller error resulted in all 5s signals from the optical bench (placement, snr, contrast, lamp, gain) showing as -16384, indicating a crc error. inspection of the console, the 5s measurements from the optical bench showed that all parameters were within specification. monitoring of the analog data from the optical bench¿s ccd array did not reveal any defects in the ¿fringe¿ pattern. the console was tested with a known good pump and purge system, with no observable anomalies. visual inspection of the internal circuitry of the console did not reveal any issues. in conclusion, the reported controller error was traced back to a communication protocol anomaly in the optical bench firmware (fw), which caused misalignment in the data communication packets received by the epc. the aic sw operated as designed.
12533 | 1220648-2024-25220 | 21047049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047049&pc=OZD | 11/23/2024 | 02/05/2025 |  | 12/31/2024 |  |  | 12/02/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025538063 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a purge system open alarm appeared during impella 5.5 support. troubleshooting was performed and a fluid leak was discovered at the yellow luer lock with liquid seen in the filter and reservoir. there was no noted damage to the purge cassette/tubing. plastic tape was utilized to partially resolve the leak, but the motor current suddenly increased the following day and the pump stopped. the pump was explanted with no known harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of pump stop and purge leak ¿ pump has been completed. the device and logs were received for investigation. the clinical details noted that a purge leak was present during support and motor current was observed to be rising before pump stop occurred. the root cause of the pump stop was most likely due to a purge leak at the yellow luer connector of the purge sidearm. the failure mode of "purge leak - pump" was removed and investigated under "pump stop. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-25220.
12534 | 1220648-2024-25221 | 21047059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047059&pc=OZD | 12/13/2024 | 05/09/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539709 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was implanted with an impella 5.5 for mechanical circulatory support as a bridge to heart transplant, and eight days into support, impella access site bleeding occurred. there were multiple dressing changes, and extra suture was applied. additionally, an arm sling was used to minimize the patient movement. next day there was no active oozing from access site only with patient movement. following this day, impella access site was still slightly saturated and managed with a dressing change. the donor was available, the patient was transported and quickly listed for heart transplant after arrival. the patient tolerated the impella 5.5 support and was bridged to heart transplant. | Additional information received via loqi (abiomed¿s long-term outcome and quality indicator impella registry) of acute thrombocytopenia. relationship to pump is listed as "possibly". loqi states: patient had impella/explant/heart transplant on (b)(6) 2024. overnight patient had bleeding and required massive transfusion (5 prbcs, 4 ffps, 2 platelets, and 10 cryo). on (b)(6) 2024, patient had chest exploration and washout, where minor bleeding was cauterized. platelet trend from (b)(6) 2024: 89-40-132-159-174-184. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. should additional information regarding thrombocytopenia, a supplemental report will be submitted. b5 revised with additional information received from the clinical site. h6 revised to reflect the additional information g2 report source; study was missing from manufacturer device report 1220648-2024-25221. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by adding additional suture. e4 should have been left blank on manufacturer device report 1220648-2024-25221.
12535 | 1220648-2024-25222 | 21047115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047115&pc=OZD | 12/11/2024 | 01/29/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509614 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the left leg looked mottled compared to the right, but pulses were difficult to obtain in either leg. the doctor gave additional orders to give fluid bolus, correct bicarbonate, and rewarm the patient, and eventually decided best course of action would be to remove pump and do surgical closure. care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. b.7 added information as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-25222. d.4 revised serial number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25222. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12536 | 1220648-2024-25223 | 21047161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047161&pc=OZD | 12/11/2024 | 01/31/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576987 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support for a high-risk percutaneous coronary intervention patient. after just over one hour of support the pump was explanted. the access site at the groin had hemostasis issues and failed device deployment. the team observed a vessel damage that was treated by a covered stent. the patient survived. | [5 empty] | The investigation for the vascular damage has been completed. no product was returned and logs are not applicable. the dissection was discovered after the manta device failed, and then balloon tamponade and manual pressure were used for closure. therefore, the root cause of the vascular damage was most likely use-related. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12537 | 1220648-2024-25224 | 21047251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047251&pc=PYX | 12/12/2024 | 01/13/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025536955 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an rp flex placed to escalate from an intra-aortic balloon pump and venous-arterial extra corporeal membrane oxygenation for a patient admitted in with peripartum cardiomyopathy. the pump was explanted after just 3 days and a protek duo was placed due to concerns of hemolysis. the patient was additionally started on continuous renal replacement therapy. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and renal failure has been completed. no product was returned for analysis. clinical details do not specify any patient condition issues. the data logs do not show trending placement signal or motor current spikes before reported hemolysis. the root cause of the hemolysis and renal failure was not determined as no product was returned for analysis and insufficient clinical details were provided.
12538 | 1220648-2024-25225 | 21047288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047288&pc=OZD | 12/12/2024 | 01/28/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550506 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory report and experienced hemolysis and would later expire after withdrawal of care. the patient was admitted to the hospital eight days prior to the impella implant for chest pain, atrial fibrillation with rapid ventricular response to rule out cardiac tamponade/sepsis/loculated pericardial effusion. the patient went into pulseless electrical activity and was intubated with a pericardial drain placed. the patient had sustained ventricle tachycardia and was taken to the catheterization lab for the impella cp device implant. thereafter, the patient was transferred back to the intensive care use and dark amber urine was observed. after a couple of hours, the urine switched to red tinged. the impella was repositioned, and it is unknown, however, if the repositioning resolved the hemolysis. the patient would expire after approximately 16 hours start of support. the patient had inoperable and non-intervention cardiac issues. and care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. a5 ethnicity was inadvertently selected incorrectly on manufacturer device report 1220648-2024-25225. g1 reporting contact email was inadvertently entered incorrectly on manufacturer device report 1220648-2024-25225.
12539 | 1220648-2024-25226 | 21047338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047338&pc=OZD | 12/11/2024 | 01/13/2025 |  | 12/31/2024 |  |  | 12/11/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, occlusion of flow was noted with peel-away sheath in place. the peel-away sheath was removed and the repositioning sheath was advanced. contrast down guide-wire re-access port showed minimal flow, not total occlusion, so placing a femorofemoral bypass with antegrade sheath was discussed. the patient¿s pulses were un-dopplerable upon arrival at baseline. it was noted that once the impella was placed and the distal perfusion catheter was placed. it was noted that the foot was pink versus purple. the family decided to withdraw care and the patient expired. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12540 | 1220648-2024-25227 | 21047442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047442&pc=OZD | 12/12/2024 | 02/12/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592554 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after the impella cp was implanted a retroperitoneal hematoma occurred in the icu. two units of blood products were administered and the impella was removed. the patient remained stable. | [5 empty] | The investigation for the vascular damage has been completed. the pump was not returned. clinical information provided is very limited. the hematoma was observed behind the peritoneal cavity. therefore, the root cause of the vascular damage issue was not determined since the product and sufficient information was not returned. b.2 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25227. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-25227 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-25227 incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12541 | 1220648-2024-25228 | 21047514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047514&pc=OZD | 12/12/2024 | 01/31/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579684 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for patient support. following insertion, it was noted that the pump was unable to advance through the subclavian artery. the pump was removed and an impella cp device was subsequently placed for patient support. there was no resulting patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella 5.5 was aborted as the surgeon was not able to pass through the subclavian as per the clinical description provided. h.6 code 4627 and 4115 was previously entered incorrectly on the initial manufacturer device report number. new codes were added. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not discarded but the device was not returned. type of investigation code has been updated to reflect not returned. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12542 | 1220648-2024-25229 | 21047515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047515&pc=OZD | 12/14/2024 | 02/24/2025 |  | 12/31/2024 |  |  | 12/14/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535848 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and the following day after start of support, the patient experienced ventricular fibrillation and received cardioversion multiple times. the position of the pump was corrected, and the patient was reported to be stable while on impella support. update a couple of days later noted, the patient still sedated and intubated with the need for catecholamines for what the physician considers the cause could be septic status. vital signs are stable and the impella was noted working without any problems. following it was noted catecholamines are slowing being reduced and the patient is now on dialysis. now day five of support, the physician noted the first significant recovery signs with less need for catecholamines. according to the physician's report, nephrological status is slowly recovering due to permanent dialysis. no issues were noted with the impella cp device. the following day slow weaning started. however, thereafter, now seven days after start of support it was noted the patient expired due to an "overall bad" health state. although noted the patient expired days later from "overall bad" heath state, the impella device cannot be excluded as a possible contributing factor when considering the impella-related ventricular fibrillation event. | [5 empty] | The investigation for the reported ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vf was not determined. b.2 revised date of death as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25229. h.6 code 4617 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25229 and was added. | The root cause of the investigation for arrhythmias such as atrial fibrillation (af), ventricular fibrillation, ventricular tachycardia, and atrial tachycardia could not be determined without additional clinical information or a device return. therefore, the root cause of the af also could not be determined without additional information. b.5 the initial manufacturer device report number 1220648-2024-25229 stated that the patient experienced ventricular fibrillation (vf). however, the correct adverse event should have been noted as atrial fibrillation (af). d.3 revised us zip code as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-25229, h.6 clinical code changed from 2130 on the initial manufacturer device report number 1220648-2024-25229 to code 1729 due to information mentioned above in b.5 corrections. h.10 the completed investigation in the initial manufacturer device report number 1220648-2024-25229 was for vf. the investigation has been updated to include af instead of vf. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
12543 | 1220648-2024-25230 | 21047528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047528&pc=OZD | 11/21/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/16/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534642 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 45 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient was thrombocytopenic. short runs of intermittent ventricular tachycardia (vt) were also noted. the patient is asymptomatic with some intermittent suction alarms resolved after albumin. the patient recovered with no sequelae. | [5 empty] | The investigation for the low pump flow, thrombocytopenia, and ventricular tachycardia has been completed. the pump was not returned for investigation. the data log shows several suction alarms with intermittent low pump flow while running at a steady p-level. no improper position-related abnormalities were reported. according to the clinical notes, several fluid boluses were administered to resolve the suction events. short runs of intermittent vt were noted on (b)(6) 2024. additionally, a ct scan revealed a right chest hematoma. the patient was diagnosed with thrombocytopenia during pump use. the cause of the thrombocytopenia issue was not determined since product was not returned and sufficient clinical details was not provided. the cause of the low pump flow issue was patient condition related to low blood volume, based on data log review and provided clinical details. the root cause of the vt could not be determined without additional clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12544 | 1220648-2024-25231 | 21047613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047613&pc=OZD | 12/17/2024 | 01/29/2025 |  | 12/31/2024 |  |  | 12/17/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563134 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump support being placed when the team failed to deliver it due to anatomy/position issues. the team aborted the impella 5.5 and instead placed an impella cp. the patient is stable and the procedure was successful with the impella cp. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12545 | 1220648-2024-25232 | 21047620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047620&pc=OZD | 12/12/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571811 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant due to worsening shock and signs of leg ischemia. the pump was prepped and inserted however, the pump was removed due to inability to pass through tortuous vessel. an impella cp was placed instead. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient had tortuous vessel. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h.6 revised health effect - clinical code as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25232. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-25232 in accordance with updated procedures.
12546 | 1220648-2024-25233 | 21047671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047671&pc=OZD | 10/25/2023 | 09/22/2025 |  | 12/31/2024 |  |  | 10/25/2023 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142964 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant sent an automated impella controller device in for service after experiencing a placement signal issue. the issue was found during a clinical case; however, it did not affect patient support, and the device did not need to be replaced. the patient remained on impella support; however patient outcome is unknown. upon sending the unit for service, it was discovered the purge flag was broken. | [5 empty] | Data analysis: there were no purge alarms found during the investigation. device analysis: upon inspection of the automated impella controller, the purge flag was noted to be missing. root cause: the cause of the purge flag issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the optical connector is very dirty when received from the field. running a pump after cleaning, removed the failures seen in the field. additionally the purge flag was missing which was not reported in the complaint description. the root cause of the aic issue was contamination on the optical connector. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
12547 | 1220648-2024-25234 | 21047722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047722&pc=OZD | 12/06/2024 | 05/01/2025 |  | 12/31/2024 |  |  | 12/06/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553742 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Atrial Fibrillation | Cardiac Arrest | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  | MW5168975 | Yes | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support and during support, the pump stopped and the patient expired. the impella device was successfully placed with plans for protective percutaneous coronary intervention (ppci). the patient was transferred to the unit and the support level was increased to p-6 which was tolerated well by the patient. the patient experienced intermittent atrial fibrillation with rapid ventricular response. amiodarone bolus was being considered by the team. however, support level were decreased to p-5. the patient was stabilized and ppci was planned for the following week. the next day after implant, the patient was weaned off all vasoactive drips. additionally, the patient was moved to a chair with a plan to continue prehabilitation for ppci. four days after start of support, the patient underwent high-risk percutaneous coronary intervention to the left main/left anterior descending and circumflex arteries. the patient was noted to have tolerated the ppci well, and a decision was being considered on whether to leave in the impella 5.5 device one more day. review was completed of the importance of restarting systemic anticoagulation if the pump was left in. the pump was left in, and the following day, however, the pump stopped from p-5 to p-0. the nurse practitioner was advised to attempt to restart the pump at support level p-2, which was unsuccessful. the nurse practitioner stated, "the wires were exposed and shredded at the red impella plug." the nurse practitioner was then advised to immediately get someone who was able to pull the device across the valve due to creating wide open aortic insufficiency. company clinical support arrived on site and the pump had been removed. the patient coded. and the patient expired. clinical support noted upon entering the room to examine the pump, the patient was in a chair, and 3-point fixation was intact. however, the 3rd point on the left side of the red plug was secured to the cable, but not to the patient¿s body. the device was taken into possession and returned for analysis/investigation. | [5 empty] | H.3 the device has been returned but has yet to be evaulated by the manufacturer so the selection was changed to no as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25234. | The investigation of pump stop and vf/vt/af/at has been completed. the data logs were reviewed and the pump stop was confirmed. on 12/11, there is a trigger of alarm 115 with a motor current spike greater to 30000 ma. this is an indication of an electrical failure. leading up to the event, there are no abnormalities noted in the logs. returned product was investigated, and as reported, the catheter was completely separated from the red pump plug. at the location of separation, there was clear necking of the electrical wires and tearing of the kink protector, indicating excessive force was used. based on data log review and returned product, the root cause of the pump stop was determined to be an electrical open. as the necessary information was not provided, the root cause of the vt/vf was not determined. e4 revised to reflect yes as medwatch was received by the fda h9 medwatch mw5168975 added attachment of medwatch mw5168975 included. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12548 | 1220648-2024-25235 | 21047782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047782&pc=OZD | 12/17/2024 | 04/02/2025 |  | 12/31/2024 |  |  | 12/17/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560821 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was received via long-term outcome and quality indicator (loqi) impella registry loqi "the adverse event for limb ischemia was listed as "definitely related" and stated the following: patient had impella device removed on (b)(6) 2024 where distal pulses were noted to be diminished, but present. an arterial duplex was ordered with a plan for continued monitoring. patient was reevaluated that afternoon and an ultrasound demonstrated a high-grade stenosis versus occlusion of her left common femoral artery. her foot was cool and slightly mottled. patients flow did not seem to improve on heparin and the decision was made to take her to the operating room for an lle fasciotomy and left groin exploration. a wound vac was applied after a fogarty embolectomy of the left sfa and lower fasciotomies were performed. patient was started on lovenox for dvt ppx. she was taken back to the operating room on (b)(6) 2024 for a washout of the wound and exchange of a new wound vac. she was taken back to the operating room on (b)(6) 2024 for a lle wound washout, debridement, and wound vac exchange. the patient was brought back to the operating room on (b)(6) 2025 for a partial closure of her medial incision of her lle. the patient was brought back once more to the or on (b)(6) 2025 where the lle wound was fully closed and dressed. the patient was deemed stable enough to be discharged on (b)(6) 2025 where she would follow up with outpatient vascular clinic." | The complainant had placed an impella cp pump to support a patient admitted in with acute myocardial infarction/cardiogenic shock, cancer, and other underlying comorbidities. the impella cp was placed and care decisions were updated from do not resuscitate/do not intubate to a full code. the support was ongoing when there was an access site bleed. the bleed was treated by infusion of two units of red blood cells and tightening of the sutures. the patient survived the event. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the bleeding was controlled by tightening perclose. | B.5 additional information was received. g2 added report source study as that is where the new information was received from. h.6 added codes due to new information received. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12549 | 1220648-2024-25236 | 21047783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047783&pc=OZD | 12/12/2024 | 01/28/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580227 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant, an impella cp pump experienced low flows and suction events. after an echocardiogram was performed, the decision was made to place an impella rp flex device in response to the issue. despite the addition of impella rp flex support, the patient passed away the following day. it is unknown if the complication contributed to the patient's passing. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12550 | 1220648-2024-25237 | 21047861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047861&pc=OZD | 12/17/2024 | 01/21/2025 |  | 12/31/2024 |  |  | 12/17/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577679 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had acute changes seen on head computed tomography scan. it was stated that this was not impella related. however the cause was not specified. the patient was on impella support before the patient expired. | [5 empty] | B.2 the exact date of death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12551 | 1220648-2024-25238 | 21047927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047927&pc=OZD | 12/17/2024 | 01/30/2025 |  | 12/31/2024 |  |  | 12/17/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553175 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had a retroperitoneal (rp) bleed from the impella access site. computed tomography showed rp occluded. heparin was paused. the patient is stable and the procedure was successful. | [5 empty] | The investigation for the vascular damage has been completed. patient's clinical details provided are very limited. the damage was noted on 5th day and self occluded so no additional intervention was needed. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12552 | 1220648-2024-25239 | 21047930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21047930&pc=OZD | 12/12/2024 | 03/07/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568885 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a purge leak was noted coming from the purge sidearm of an implanted impella 5.5 pump. troubleshooting was performed to manage the leak and support with the implanted pump continued. there was no resulting patient harm. | [5 empty] | The investigation of purge leak ¿ pump has been completed. visual inspection found no issue on purge side arm. no leak was found under pressure leak test. the root cause of the purge side arm leak was most likely use related as no problem was found under pressure leak test. d6b added. d9 date device returned to manufacturer added. g1 reporting contact email revised on manufacturer device report 1220648-2024-25239 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
12553 | 1220648-2024-25240 | 21048010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048010&pc=OZD | 12/16/2024 | 01/24/2025 |  | 12/31/2024 |  |  | 12/16/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316943 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) purge lever was suspected to be broken. the aic was exchanged. no patient harm has been reported. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the console was examined and a broken purge disc flag was identified. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12554 | 1220648-2024-25241 | 21048055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048055&pc=OZD | 12/13/2024 | 03/24/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller had a damage power cable with exposed wire. there was no patient involvement at the time the damage was noted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the data logs do not provide information relevant to this complaint / failure mode. the console was returned for investigation. upon arrival, the console underwent testing, which revealed an issue during visual inspection. two of the internal cables within the plug were found to be disconnected. further investigation involved removing the transparent cover of the plug. this revealed that both the green and white cables were disconnected. the issue appears to be the insufficient length of the wire covering, which was not cut according to the specified maximum length of 1-1/8 inches (1.125 inches). this ultimately contributed to the inadequate wire covering, which was noticeably shorter than that of a properly functioning unit. the cables' outer sheath/jacket length was inadequate to prevent pinching at the plug, which is why the cord broke loose. the root cause of the issue was a defective ac power cord cable. e2 and e3 selections were inadvertently reported incorrect on the initial manufacturer device report number 1220648-2024-25241.
12555 | 1220648-2024-25242 | 21048083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048083&pc=OZD | 07/08/2023 | 12/31/2024 |  | 12/31/2024 |  |  | 07/08/2023 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122977 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 support the nurse was performing a cassette change and received an alarm stating "automated impella controller (aic) will not detect the purge disc". it was also reported during a phone conversation with the patient's nurse that the "impella stopped working", and the nurse thought there was a controller failure. as a result of the issues, the device was swapped with another aic. there was no reported patient harm for the aic issues. simultaneously, the patient had an axillary hematoma, at the access site of the impella 5.5 pump delivery. patient was also complaining of tingling in their fingers resulting in a wash out of site. the patient remained on impella 5.5 support until successfully weaned for a heart transplant. | [5 empty] | This is one of two medical device reports associated with the event. refer to initial medwatch with mfr 1220648-2023-02603 for the impella 5.5 serial number (b)(6). data analysis: on 16-apr-2023, automated impella controller (console) with serial number (b)(6) had multiple instances of purge disc not detected. the console was then not used again until 6-jul-2023 where the purge disc was initially able to detect the purge disc before the purge disc was unable to be detected again on 8-jul-2023. there was no evidence in the logs of a "controller failure". most likely, the complainant was trying to say the impella was not functional. device analysis: console (b)(6) was confirmed to have a missing purge flag allowing the purge disc not detected failure mode seen in the field to be reproduced. root cause: the cause of the issue was a missing purge flag. this report is being filed as part of a retrospective review of historical records.
12556 | 1220648-2024-25243 | 21048193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048193&pc=OZD | 12/12/2024 | 09/24/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578579 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella 5.5 support and during repositioning, the pump bumped the septal wall and the patient went into sustained ventricular tachycardia. the patient received multiple shocks and was placed on extracorporeal membrane oxygenation and defibrillated. support was continued. additionally, there was intermittent position in aorta alarms. position was verified and the placement signal was monitored. support was continued and the patient outcome was stable. | [5 empty] | The investigation of ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. e.1 fax number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25243. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25243. new health effect - impact codes have been added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the 5s parameters were stable throughout the case but there was periodic low signal-to-noise ratio (snr0 that corresponds with a lack of good signal being read and improved throughout the case. during the instances of the ¿impella position in aorta¿ alarms the placement signal would rapidly spike while the motor current would lose pulsatility. the alarm correctly met the triggering criteria. however, echocardiography (echo) confirmed the position in the correct place and the impella position in aorta alarms were intermittently triggering leading to infer that the alarm was falsely triggered. the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1. product problem updated. h6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-25243. d10. concomitant medical products and therapy dates updated.
12557 | 1220648-2024-25244 | 21048209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048209&pc=OZD | 12/12/2024 | 02/13/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506404 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an impella cp device implant to support an unknown age patient admitted with cardiogenic shock from an acute myocardial infarction, the catheter "broke." the device was replaced, and the implant was successfully completed. no further information was provided including patient outcome (patient-specific information other than the sex is unknown). however, no harm or injury was noted with the report of the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) has been completed. the impella device was not returned for evaluation. according to the clinical details the root cause of cannula kink was most likely due to use related. b3 date of event was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25244. d6a implant date was inadvertently omitted from manufacturer device report 1220648-2024-25244. d6b explant date was inadvertently omitted from manufacturer device report 1220648-2024-25244.
12558 | 1220648-2024-25245 | 21048237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048237&pc=OZD | 12/13/2024 | 01/31/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump experienced complications with their optical sensor during patient support. the pump was replaced to troubleshoot the issue and support was continued. there was no resulting patient harm. | [5 empty] | The investigation of optical signal issue has been completed. the root cause of the optical signal issue cannot be determined as the product was not returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12559 | 1220648-2024-25246 | 21048343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048343&pc=OZD | 03/28/2022 | 12/11/2025 |  | 12/31/2024 |  |  | 03/28/2022 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1467028( |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [17 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 support it was noted that the automated impella controller's (aic) battery temperature was higher than expected when unplugged. as a result, the medical team proactively changed the aic out with another unit. there was no interruption to patient support, and no patient harm reported. it was additionally reported, that the patient developed some oozing at the impella site. a standard bleeding protocol was followed to treat the bleeding. | [5 empty] | D10 concomitant medical products and therapy dates updated. h1: corrected reportability. h6: added code. | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-07145 for impella 5.5 sn (b)(6). data analysis: complaint date is consistent with complaint intake. on 3/28/22, automated impella controller (console) logs reported a "battery temperature high" alarm and cleared within ~5 secs. this occurred just as ac power was disconnected. long term logs were also reviewed at the time the alarm was triggered. it was observed that battery temperatures were well below the threshold to trigger the alarm. it is most likely that when ac power was disconnected, a single sample of communication was misinterpreted from the batteries by the pbm. this is evident by the fact that the alarm was triggered and cleared in under ~20 secs. device analysis: console arrived in danvers. console was ran in burn-in chamber on pump and purge with no issues for an extended period of time. the root cause of the battery temperature high alarm is most likely due to a misinterpretation of data within a single sample of communication from the batteries to the pbm. however the overall root cause is undetermined. this report is being filed as part of a retrospective review of historical records.
12560 | 1220648-2024-25247 | 21048368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048368&pc=OZD | 12/13/2024 | 02/04/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579670 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the pump was shallow due to the patient anatomy. it is unknown if the pump was repositioned to optimal position. additionally, upon explant, a large hematoma was noted at the access site. covered stents were placed and a vascular plug to the profunda. a clot was also removed from the superficial femoral artery. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the positioning issue, access site bleed, and thrombus has been completed. the product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely use related since there was reported poor hemostatic control during removal. no information related to the positioning issue was found in the clinical information. the root cause of the thrombus could not be determined without additional clinical details. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." h.10 added instructions for use for thrombus as they were inadvertently omitted from the initial manufacturer device report number 1220648-2024-25247. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12561 | 1220648-2024-25248 | 21048408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048408&pc=OZD | 08/05/2023 | 10/21/2025 |  | 12/31/2024 |  |  | 08/05/2023 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1708923 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during impella 5.5 support, the pump stopped and the team was unable to restart. the patient was dependent on the pump. the patient decompensated, cpr was initiated and the patient expired shortly after the pump stoppage. | [5 empty] | The investigation into the patient death has been completed. the data logs show an impella #119 alarm. the device is needed to confirm the cause of the #119 alarm. the cause of the pump stop was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. data analysis: multiple pump stops were confirmed in the rtlogs due to a high motor current. (b)(4). device analysis: full investigation of the console involved running pump and purge cassette and revealed no issues. the cause of the pump stops due to high motor current were not determined given inability to reproduce the issue on the returned console. repair: field service instructed to perform; a full functional check (0042-7316), preventative maintenance procedure (0042-7309).
12562 | 1220648-2024-25249 | 21048432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048432&pc=OZD | 12/11/2024 | 02/14/2025 |  | 12/31/2024 |  |  | 12/14/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575252 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an impella cp device implant for mechanical circulatory support to a patient with cardiomyopathy, the catheter shaft kinked. the kink was found after the impella was inserted into the left ventricle. the position adjustment was unable to be properly performed and as the patient was planned for long-time management due to myocarditis, the physician decided to explant and replace the impella cp device. it was reported there were no health hazard as a result of this event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. the pump was not returned for investigation. as per the clinical information, the catheter was kinked due to excessive force. therefore, the root cause of the positioning issue was most likely use related to excessive force.
12563 | 1220648-2024-25250 | 21048440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048440&pc=OZD | 12/27/2022 | 12/31/2024 |  | 12/31/2024 |  |  | 12/27/2022 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1095652 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the field service team that the automated impella controller (aic) purge cassette door would not latch and stay closed. a loaner aic was sent to the customer to initiate return of the aic. there was no patient harm reported. | [5 empty] | The investigation into the purge door wont close issue has been completed. the automated impella controller (aic) was returned for investigation. the real time logs were reviewed and showed pressure fluctuations that were likely related closing the purge door on the purge tubing. the returned console was inspected on an engineering lab bench. the failure mode was not reproduced. the door opened and closed without issues. there was very minimal buildup of glucose around the purge assembly, making it unlikely to have contributed to the purge door latch mechanism malfunction. the root cause of the inability to close the purge door was most likely due the purge tubing pinched between the door and enclosure (not properly placed in the groove, which allows for full closure of the purge door). this report is being filed as part of a retrospective review of historical records.
12564 | 1220648-2024-25251 | 21048441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048441&pc=OZD | 12/16/2024 | 02/06/2025 |  | 12/31/2024 |  |  | 12/16/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569692 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had bleeding around the access site and there were suction alarms on the impella. blood products were given to the patient. bleeding was controlled with manual pressure and femostop. support was continued. the patient is stable and the procedure outcome was successful. | [5 empty] | The investigation for the access site bleed and low pump flow has been completed. data logs were reviewed which showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the low flow was patient related as the pump had suction alarm due to low cvp and was resolved by blood and fluid given. clinical team didn't complaint about low flow. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12565 | 1220648-2024-25252 | 21048471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048471&pc=OZD | 12/13/2024 | 01/23/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466927 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient lost pulses to the right leg. the reperfusion catheter was clotted off. the staff explanted the pump and the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12566 | 1220648-2024-25253 | 21048479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048479&pc=OZD | 12/13/2024 | 01/31/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511230 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being delivered for a surgical patient. the pump failed to deliver despite all attempts and measures taken to deliver. the 5.5 was aborted and extracorporeal membrane oxygenation placed while family makes medical choices. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass the wire and catheter across the av. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12567 | 1220648-2024-25254 | 21048615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048615&pc=OZD | 12/13/2024 | 01/30/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467743 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and it was aborted. the pump would not cross the aortic valve. the patient¿s outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12568 | 1220648-2024-25255 | 21048640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048640&pc=OZD | 12/14/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/14/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590627 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on day two of impella cp support had access site slightly oozing. the doctor placed a stitch at the site and bleeding was resolved. the patient is stable and survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access s site bleeding is determined to be use issue because the bleeding was resolving by applying the stitch at impella site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12569 | 1220648-2024-25256 | 21048830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048830&pc=OZD | 07/26/2021 | 12/31/2024 |  | 12/31/2024 |  |  | 07/26/2021 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438552 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support for high-risk percutaneous coronary intervention. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by replacing the aic. there was no reported patient harm. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aci was returned for investigation. data analysis of the data logs showed that during case prep, console is booted and the first purge cassette is inserted as the purge cassette begins to build pressure in the transducer, logs show: the console prompts the user to check the pump for kinking, and is unable to complete priming. the priming is again attempted and stopped by the console because of a kink. at this point the line is de-aired and the purge cassette begins to reprime. purge fluid information is entered and no kink is detected. at this point the case start is exited, and a new purge cassette is inserted. at the start of the purge priming, there are several log entries for purge fluid not detected. the aic was swapped out after several attempts to prime the pump are failed with 2 purge cassettes. device analysis determined the root cause as it was observed that the purge disc/transducer retainer was broken, cracked, and lifting off of the purge assembly. the retainer is not completely broken off and the purge flag is in place, which is likely why there were no purge disc detection issues. this report is being filed as part of a retrospective review of historical records.
12570 | 1220648-2024-25257 | 21048839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048839&pc=OZD | 12/01/2024 | 02/13/2025 |  | 12/31/2024 |  |  | 12/15/2024 | 12/31/2024 |  |  | 06/30/2025 |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365211 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Perforation | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, there were suction events. the position was checked with echocardiography and volume was administered. the right ventricle failure was diagnosed and extracorporeal membrane oxygenation support was indicated. the impella cp was moved to the ventricle and it was found that the touhy valve was not closed. the patient evolved unfavorably due to the severity of his clinical condition. additionally, there was a suspicion of perforation/rupture which according to interpretation could have occurred during implantation. the patient expired. | [5 empty] | The investigation of ventricle perforation and low pump flow has been completed. the impella device was not returned for evaluation. based on the limited clinical details and product not being returned, the root cause of the ventricle perforation could not be determined. based on data log review and the clinical details provided, the root cause of the low pump flows was determined to most likely be patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12571 | 1220648-2024-25258 | 21048914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048914&pc=OZD | 12/09/2024 | 02/07/2025 |  | 12/31/2024 |  |  | 12/16/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547020 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support ongoing for a patient admitted in for a high-risk percutaneous coronary intervention. the pump support was ongoing when there was an observation made of product damage. the damage allowed for leakage. the pump was explanted and not replaced. the patient¿s outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The pump was not returned for investigation, and a data log review was not required for this case run. according to the clinical team, blood was discovered coming from the catheter plug and shaft. the clinical team may have referred to a leak from the red handle, which is most likely caused by a hole in the catheter, allowing blood to pass through the catheter and leak from the red handle. the cause of the product damage issue was not determined since product was not returned and sufficient clinical information was not provided.
12572 | 1220648-2024-25259 | 21048979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21048979&pc=OZD | 12/16/2024 | 01/29/2025 |  | 12/31/2024 |  |  | 12/16/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692213 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the console does not recognize the pump. the automated impella controller was removed from service. there was no patient involvement. | [5 empty] | The investigation for the pump not detected has been completed. in the case start wizard at step three, the pump connection was most likely not recognized, or the white catheter plug was not completely inserted into the automated impella controller, so after 20 seconds, epm was reset automatically. after the epm reset, the pump connection was still not recognized. this confirmed the reported failure mode. this could be most likely due to the crystal issue on the impellatronic (epm) board. then the user pressed the next soft button, so the console displayed ¿connect cable: detecting impella screen¿ and ¿connect cable: connect grey connectors screen.¿ this console was tested after arriving in field service. the reported failure modewas reproduced 2/5 times. the console could not recognize the optical test pump connection. the i2c data signal was missing during the unsuccessful attempts. upon replacing the original impellatronic board with the known good, they were unable to reproduce the reported failure mode. the root cause for not recognizing the pump connection was impellatronic malfunction. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12573 | 1220648-2024-25260 | 21049086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21049086&pc=OZD | 12/13/2024 | 01/28/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556254 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. impella 14fx25cm sheath attempted to be placed in right femoral artery (arteriogram looked clear but slightly tortuous), sheath could not be advanced past the iliac bifurcation, instead 14fx13cm sheath was successfully placed. impella cp was inserted via sheath into artery but could not be advanced past the bifurcation. shockwave and buddy wire were suggested but the physician elected to abort case and reschedule. there was no harm to the patient. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had tortuous vessel. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12574 | 1220648-2024-25261 | 21049226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21049226&pc=OZD | 07/31/2023 | 02/06/2025 |  | 12/31/2024 |  |  | 07/31/2023 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1551790 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support, the automated impella controller (aic) had a battery failure after starting. the aic was exchanged for another unit. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. console logs from the reported day of event are consistent with complaint and show battery failure alarm. the cause of the issue was use related (batteries were not engaged after transport).
12575 | 1220648-2024-25262 | 21049233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21049233&pc=OZD | 12/13/2024 | 01/27/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547486 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had significant dissection noted from proximal to distal left anterior descending artery (lad). the doctor (md) called for impella placement and impella cp placed without any issues. md continued to work on lad dissection. decided to stop intervention and send the patient to the unit to rest. in the unit, hematoma was noted in the right femoral artery. pressure was held and femostop placed. heparin drip was stopped until access site was stable. two units of packed red blood cells were given due to patient bleeding from the access site and multiple line sites. it was also reported that the family requested for withdrawal of care and patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12576 | 1220648-2024-25263 | 21050164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050164&pc=PYX | 12/12/2024 | 01/31/2025 |  | 12/31/2024 |  |  | 12/12/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025487353 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on bipella support was heparin-induced thrombocytopenia positive. the patient was switched to argatroban and purge fluid in the impella. the patient is stable and survived the event. | [5 empty] | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25263. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12577 | 1220648-2024-25264 | 21050170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050170&pc=OZD | 12/13/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569106 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had a possible retroperitoneal bleed. the patient received two units of packed red blood cells. the staff is planning on explanting the pump. the patient survived the event. | [5 empty] | The investigation for the vascular damage has been completed. the product was not returned. the cause of the vascular damage issue was not determined since product was not returned and due to insufficient information with unknown relation of rp bleed to impella. the instructions for use for the impella 5.5 for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12578 | 1220648-2024-25266 | 21050193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050193&pc=OZD | 12/11/2024 | 02/06/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371809 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) waveform is noisy and choppy and the position waveform is showing unreliable values. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - loss of opm data has been completed. console logs from the reported day of the event were consistent with the complaint. the logs showed multiple instances of the following errors: "algs suspended opm signal invalid," and "invalid_bad_snr_sample_mask" due to signal-to-noise ratio (snr) falling below the minimum threshold. real time log data indicated low contrast and snr values, while the long term log data showed consistently low snr throughout the case. further analysis of console logs revealed similar symptoms with peak snr and multiple placement signal low alarms during the case and long term log data showed snr was low for most of the case. the console was returned. however, the reported issue could not be reproduced. the 5s parameters were all within specification, and voltage measurements from both the power battery manager and across the optical lamp were within required ranges. the bulb output power also met the necessary specifications. visual inspection revealed no apparent issues, and no abnormalities were found with the automated impella controller (aic) operating alongside a functional pump and purge cassette. the cause of the aic issue was most likely a defective optical bench. b.3 revised date of event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25266. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25266 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25266 and a new code was added.
12579 | 1220648-2024-25267 | 21050259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050259&pc=PYX | 12/13/2024 | 03/10/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  | 11/30/2025 |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024433022 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support. while on support, the patient was found to have a right ventricle outflow tract rupture. rupture was repaired. the doctor felt that this rupture happened during the implant. additionally, the patient stopped making urine. the impella rp flex position was unchanged from previous x-ray - appearing to be in adequate position. due to hemolysis, the impella rp flex was explanted. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella rp smartassist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the ventricle perforation and hemolysis has been completed. following a pull-thru insertion technique, the rv outflow tract was found to be ruptured which required additional repair and was fixed in the same setting. this is not the recommended insertion technique as per rp flex ifu document. the cause of the ventricle perforation was most likely use related as it occurred during insertion with an off-label technique. data logs show noisy ps and cvp with brief suction alarm at start of case that recovers. overall, motor current (mc) and flows trending down throughout the case and placement signal (ps) and cvp trending up. no mc spikes outside of p-level changes, which do not suggest biomaterial ingestion. no other relevant abnormalities observed. returned product showed 3 cannula kinks and small biomaterial around the impeller. the pump passed hemolysis testing mih = 3.63. the patient had cabg and vsd/vsr prior to support. the pump was confirmed to be in proper position with no significant changes. poor patient condition and cannula kinks found on the returned product are potential contributing causes to the hemolysis, however a definitive root cause could not be determined.
12580 | 1220648-2024-25268 | 21050269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050269&pc=OZD | 12/13/2024 | 02/04/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and before the impella implant, the patient had a gastrointestinal bleed. while on support, heparin was put in the purge and gastrointestinal bleed continued. blood products and medication were given to the patient. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the vascular damage has been completed. the pump was not returned. the cause of the vascular damage issue was not determined since the product was not returned and unknown relation to impella with gastrointestinal bleed being observed. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25268 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 coded 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25268 and was added. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-25268 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12581 | 1220648-2024-25269 | 21050291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050291&pc=OZD | 12/13/2024 | 01/23/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568649 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being placed to support a patient admitted in and had a coronary artery bypass graft performed in the operating room. the pump was being placed as being weaned/escalated off extra corporeal membrane oxygenation. the 5.5 failed to deliver due to guidewire and pump damage/kinking. the product was replaced and there was no harm or injury noted. | [5 empty] | The investigation for the delivery issues and product damage has been completed. product was not returned for review. based on clinical descriptions, the root cause of delivery issue was most likely use related. the cause of kinked cannula was most likely use related. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12582 | 1220648-2024-25270 | 21050300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050300&pc=OZD | 09/11/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510634 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Arrhythmia | Hemolysis | Ischemia | Tachycardia | Thrombocytopenia | [7 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The complaints team received three long-term outcome and quality indicator (loqi) notifications that found that: a patient that endured thrombocytopenia and anemia with a possible relationship to the impella device used: likely consumptive. low but stable. heparin-induced thrombocytopenia (hit) sent on (b)(6) resulted negative. on (b)(6) plt 55k/ul; heparin and tricagrelor continued. on (b)(6) plt 66 k/ul. on (b)(6) plt 64 k/ul. hit sent on (b)(6) resulted positive. blood products given prn. heparin stopped and argatroban started. resolved on (b)(6)2024. plt 64 on (b)(6) 55 on (b)(6) 64 on (b)(6) 96 on (b)(6) 191 on (b)(6).¿. the patient recovered with no sequelae. a patient that endured non-traumatic compartment syndrome of right lower extremity (rle) with a possible relationship to the impella device used: on (b)(6) 2024: patient's rle signals stable. ck uptrend to 11k. neuvascular checks and serial ck continued. on (b)(6) 2024: patient's ck trended up to 13k on (b)(6) 2024. it was noted that the right compartment was tense compared to the remainder of the calf and left lower extremity. patient also had sensory loss on the dorsum of the right foot between 1st and 2nd metatarsal heads. patient underwent right leg 2 fasciotomy. the patient recovered with no sequelae. a patient that endured non-sustained ventricular tachycardia with a possible relationship to the impella device used: patient had previous occurrence of nonsustained ventricular tachycardia. patient re-admitted on (b)(6) 2024 for non-sustained ventricular tachycardia. heart rate was monitored. patient continued mexiletine 200 mg po tid and metoprolol xl 25 mg po bid for treatment. the patient recovered with no sequelae. the patient survived the events. | [5 empty] | The investigation for the thrombocytopenia, hemolysis, limb ischemia, and ventricular tachycardia (vt) has been completed. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of thrombocytopenia, limb ischemia, and vt was not determined due to insufficient clinical details. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12583 | 1220648-2024-25271 | 21050310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050310&pc=OZD | 12/14/2024 | 02/03/2025 |  | 12/31/2024 |  |  | 12/14/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508001 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Renal Failure | Thrombocytopenia | [9 empty] | Life Threatening | Required Intervention |  |  |  |  | CONTINUOUS RENAL REPLACEMENT THERAPY | EXTRA CORPOREAL MEMBRANE OXYGENATION | VENTILATOR | [16 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support, computed tomography scan of the head confirmed left posterior lower lobe infarct and right lobe abnormalities. also, the patient was heparin-induced thrombocytopenia positive per panel and was switched to argatroban. additionally, hemolysis is a concerned which it was detected via plasma free hemoglobin. the patient remains on venous-venous extra corporeal membrane oxygenation, impella 5.5, continuous renal replacement therapy, and ventilator. the patient survived the events. | [5 empty] | The investigation for the thrombocytopenia, stroke, and hemolysis has been completed. the impella 5.5 pump was returned for evaluation and was visually inspected. a biomaterial around outflow cage observed, but none wrapped around the impeller. data logs show no motor current (mc) spikes outside of p-level changes; therefore, this biomaterial was likely picked up on explant of the pump as observed in the clinical. no sharp edge or broken blade found on returned product. only data logs available to (b)(6), the pump was explanted on the (b)(6), however this is sufficient data logs for investigation as hemolysis was reported on the (b)(6). no mc spike observed outside of p-level changes during the time of the reported hemolysis. many unknown impella position alarms observed near case start which resolve, no comments in clinical regarding positioning issues. many suction alarms occurred with variable flows and ps. ps trends down slightly throughout the case. flows are consistently noisy and variable throughout support, though within acceptable values for the p-levels. the patient's plasma free was trending up while on support but was fluctuating from 12/11 to 12/13 (50, 130, 110) until hemolysis concerns reported. the root cause of the hemolysis issue was most likely patient condition related. the root cause of the stroke and thrombocytopenia could not be determined without additional information. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12584 | 1220648-2024-25265 | 21050351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21050351&pc=OZD | 12/13/2024 | 09/19/2025 |  | 12/31/2024 |  |  | 12/13/2024 | 12/31/2024 |  |  |  |  | 12/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562488 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The us complainant had placed a cp to support a patient admitted in acute myocardial infarction / cardiogenic shock. the pump was placed but, there was an access site bleed and optical signal issues. the placement signal did not prompt explant; however, the bleed did prompt intervention in blood products and eventual explant. the patient survived the event. | [5 empty] | The investigation for the access site bleed has been completed. clinical details noted that groin bleeding was observed and was able to be first resolved with manual pressure. however, bleeding occurred again and was unable to be controlled and patient required a transfusion and a repair of the left femoral artery after pump explant. the root cause of the vascular damage could not be definitively determined as limited clinical details were provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25265. as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. data logs were reviewed and reoccurring suction alarms and "placement signal low" alarms were observed throughout case. placement signal and mean flow are variable during case. no hard failures of the optical sensor were noted, all optical parameters were stable throughout case. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-25265.
12585 | 1220648-2024-25181 | 21034529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034529&pc=OZD | 01/26/2024 | 01/14/2025 |  | 12/30/2024 |  |  | 12/05/2024 | 12/30/2024 |  |  | 10/31/2025 |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411182 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute thrombocytopenia and an impella access site hematoma both events with a "possible" relationship to the impella device used. an impella 5.5 was selected and implanted in the patient for mechanical circulatory support via right axillary cutdown without complication. the patient was being worked up for left ventricular assist device but noted would transfer out if approved for a heart transplant. the patient was approved for the heart transplant and was transferred to an outside hospital. however, the next day after the impella 5.5 device was implanted, acute thrombocytopenia likely secondary to the impella was noted with no sign of hemolysis. heparin-induced thrombocytopenia (hit) was confirmed two days later. the patient was given agatroban for anti-coagulation which was be later switched to heparin and the issue was resolved. the patient recovered with no sequelae. seven days after confirming hit, a hematoma at the impella site was noted. the hematoma resolved with cardiectomy and removal of right axillary graft during surgery. of note, the impella 5.5 device was removed prior to the new heart being implanted. the patient recovered from the hematoma event with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12586 | 1220648-2024-25184 | 21034620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034620&pc=OZD | 12/10/2024 | 01/28/2025 |  | 12/30/2024 |  |  | 12/10/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570490 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced unresolvable low pump flows. during the impella supported percutaneous coronary intervention, the patient experienced loss of flow from impella cp that could not be corrected with repositioning, changes in preload or breaking of suction. unresolvable low flows led the decision to replace the pump for suspected clot entrainment. it was noted there was no harm to the patient as a result of this event. | [5 empty] | The investigation for the low pump flow has been completed. the returned pump was inspected and there were no visual abnormalities. there was no biomaterial in the outflow cage or by the impeller. the data logs confirm low pump flow for matching p-level. the logs do not show any motor current spikes or positioning issues. the logs show placement signal trending downward, and suction alarms. clinical details mention prior episode of low pump flow that was resolved. it is unknown if any biomaterial was observed in the field. the root cause of low pump flow was most likely patient condition related. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
12587 | 1220648-2024-25186 | 21034646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034646&pc=OZD | 12/09/2024 | 02/07/2025 |  | 12/30/2024 |  |  | 12/09/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539796 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp for mechanical circulatory support and upon removal encountered difficulty. pci was completed and the impella cp was rapidly weaned and removed through the peel-away sheath. the physician attempted to perclose with two pro glide. one perclose failed and another was attempted to be placed that also failed. activated clotting time (act) remained elevated. therefore, the decision was made to insert a long 14 french cook sheath and stabilize the site until act decreased. insertion of the cook sheath was traumatic, and the patient developed a medial site hematoma that required pressure to be held at the site. the patient was transferred to cardiovascular intensive care unit for monitoring until sheath was removed and once removed, manual pressure held for 40 minutes at site. following the site was noted stable and soft. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of removal issues has been completed. no product was returned and logs are not applicable. the failure mode "removal issues" was removed since the pump was able to be explanted.
12588 | 1220648-2024-25188 | 21034649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034649&pc=OZD | 12/11/2024 | 02/07/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764867 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was on an impella 5.5 pump support when the automated impella controller (aic) alarmed intermittently. the back up aic was obtained in case the issue reoccurred. information last known indicated the console has replaced. there was no report of harm to the patient following this event. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Analysis: on four separate occasions (12/25, 12/27, 12/28, and 12/30) during the 51 days of running 5.5 pump, the logs noticed singular 1045 errors for crc mismatchesi. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error 12_27 transient loss.png, the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed.
12589 | 1220648-2024-25189 | 21034657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034657&pc=OZD | 11/16/2024 | 01/29/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546821 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a right anterior chest wall hematoma at the impella insertion site. it was noted to be stable. a vascular surgeon was consulted to which it was suspected to be an iliopsoas hematoma ftom a spontaneous bleed from anticoagulation, as an intraaortic balloon pump was placed on the right femoral artery and unable to visualize any procedures done on the left femoral artery. subsequently, hematology was consulted for spontaneous bleeding of unclear etiology. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding issue was most likely patient condition related, based on provided clinical details. e4 should have been left blank on manufacturer device report 1220648-2024-25189. g2 report source; study was inadvertently omitted from manufacturer device report 1220648-2024-25189. h10 incorrect instructions for use for the impella cp were inadvertently included on manufacturer device report 1220648-2024-25189.
12590 | 1220648-2024-25182 | 21034746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034746&pc=OZD | 12/09/2024 | 01/15/2025 |  | 12/30/2024 |  |  | 12/09/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557827 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient had ongoing hemolysis. positioning was checked with echocardiography. volume status was adequate. the decision was made to remove impella. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were positioning alarms triggered during the case but resolved quickly. logs also showed placement signal was trending down and declining since start of case. several impella position unknown alarms were also observed during the case. product was returned in damaged condition (missing pump plug). visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. although biomaterial was found in the pump housing but no evidence indicated that biomaterial caused hemolysis. most likely, it came through the graft during removal as no motor current (mc) spike interaction seen in data logs. the root cause of hemolysis was most likely patient condition related
12591 | 1220648-2024-25190 | 21034775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034775&pc=OZD | 12/11/2024 | 11/04/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571095 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant had placed a impella cp pump to support a patient admitted in with ventricular tachycardia and acute myocardial infarction and cardiogenic shock. the patient was transferred from community to tertiary center for discussions of care escalation and/or transplant. the pump stopped after 15 days and was explanted. the pump was replaced and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25190. g1 reporting contact email revised on manufacturer device report 1220648-2024-25190 in accordance with updated procedures.
12592 | 1220648-2024-25183 | 21034797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034797&pc=OZD | 12/09/2024 | 09/30/2025 |  | 12/30/2024 |  |  | 12/09/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569570 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on day 18 of impella 5.5 support had a gastrointestinal bleed and systemic heparin was stopped. on day 23 of impella support, impella in aorta alarm occurred. the impella was turned down to p2 and the doctor repositioned the 5.5 at bedside. after repositioning the impella, alarms briefly resolved, and flows increased from 3.5 to 4.8 on p7. shortly after, the patient became more tachypneic. the patient complained of a headache and blurry vision. a code stroke was called. the doctor was at bedside and patient was intubated for respiratory distress. the doctor noted asystole on echocardiography and a code was started. the impella decreased to p2, and compressions initiated, return of spontaneous circulation achieved. patient received three units of packed red blood cells and one unit of platelets. vasopressor and inotropic medication started. post code, the impella continued to have alarms with correct positioning. pump failure was suspected. the doctor made the decision to explant impella and cannulate for extra corporeal membrane oxygenation. the next day care was withdrawn and the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vascular damage - peripheral vessel, saturated flow, and stroke/cva has been completed. based on clinical details and data log review, the root cause of the saturated flow was determined to most likely be biomaterial. as this clinical information was not provided, the true root cause of the stroke/cva and vascular damage - peripheral vessel could not be determined b2 date of death was omitted from manufacturer device report 1220648-2024-25183. e4 should have been left blank on manufacturer device report 1220648-2024-25183. h6 health effect - clinical code 1770 was omitted from manufacturer device report 1220648-2024-25183. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. vascular damage peripheral: data logs are not relevant to the complication and product was not returned for investigation. clinical details provide limited information as to what could have caused the bleed and when it resolved. due to limited clinical details and product not being returned, the root cause of the vascular damage could not be determined. saturated flow: data logs were reviewed and the saturated flows were confirmed on 12/14. leading up to this, there was a clear trend up in the placement signal (ps) while the motor current steps down. this is followed by pump flows trending down while the p-level remained constant. when pump flows got to their lowest, suction alarms started to trigger and the user dropped the p-level. then shortly after, there is a sharp motor current spike and severe drop in motor speed without a change in p-level. this is followed by elevated motor current and sustained saturated flows for the majority of the remaining support. these are all indications of biomaterial ingestion. this is supported by the clinical detail that after the gi bleed on 12/12, heparin was put on pause. product was not returned for investigation. based on clinical details and data log review, the root cause of the saturated flow was determined to most likely be biomaterial. optical signal issue: based on the findings of data log review, the optical signal issue failure mode had to be investigated. the impella position in aorta alarms were confirmed. very shortly after a motor current spike and drop in motor speed, the first alarm triggers. at the time of the mc spike, snr, contrast, and gain all have sudden drops that all recover immediately. based on the finds that the root cause of the saturated flows was biomaterial and that clinical details reported position was confirmed with echo, the biomaterial likely caused the trigger of the positioning alarms as well. when the biomaterial got ingested into the cannula, it likely reduced the pulsatility of the pump flow, causing the algorithm to believe the pump was in a single heart chamber. product was not returned for investigation. based on clinical details and data log review, the root cause of the optical signal issue was determined to most likely be biomaterial. stroke: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation; timing of the stroke in relation to impella support (during or post-implant); detailed description of the symptoms experienced by the patient; neurological examination findings and any focal deficits observed; diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic); laboratory test results, including coagulation profiles and cardiac markers; any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications; response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology; as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
12593 | 1220648-2024-25185 | 21034845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034845&pc=OZD | 12/10/2024 | 01/20/2025 |  | 12/30/2024 |  |  | 12/10/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548318 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and encountered positioning issues that eventually resulted in device replacement. the impella cp was successfully placed in high-risk percutaneous coronary intervention (hrpci) setting successfully. two preclosures were used. the device inserted successfully without complications. inlet directed towards apex of heart, positioning optimal under fluoroscopy. repositioning sheath in place with t-lock secured as well as tuohy-borst valve. the decision was made to bring patient to the operating room for extracorporeal membrane oxygenation (ecmo) placement. during the hrpci case, patient¿s oxygen saturation was in the 50s. the patient was brought to operating room and during sterile prep and draping, the impella cp was mistakenly pulled back into aorta. a transesophageal echocardiogram (tee) confirmed the pigtail was not across valve but right at the forefront of the aortic valve. the decision was made to keep impella pump performance level at p-2 during cannulation of ecmo and to fix positioning after. during cannulation, patient-maintained pressures with help of norepinephrine and epinephrine. once cannulated and ecmo running, tee guidance utilized with attempt at repositioning the impella cp pump back across valve. multiple attempts were made with reduction in ecmo flows for filling pressures as well as pushes of inotropic agents to improve valve opening. both were unsuccessful in helping with crossing the valve. decision was mad to attempt a buddy wire technique. the physician was unsuccessful with using j-tip wire to cross valve. the decision was made to keep vessel access with 0.035 stiff wire through reaccess port and proceed with second impella cp device implantation. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since the pump was mistakenly pulled back into aorta and couldn't be inserted again successfully while the 2nd pump was placed successfully. the instructions for use for the impella cp with smartassist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
12594 | 1220648-2024-25187 | 21034846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21034846&pc=OZD | 12/11/2024 | 02/03/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  | 04/30/2025 |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349650 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and immediately encountered positioning issue with mitral regurgitation that would ultimately lead to explanting the device. the patient had a small aortic arch and after the first impella cp implant that was unsuccessful, this second impella cp was attempted. after implant, suction alarms immediately occurred after starting and only resolved at p1. flows would not go above 1 l/min. transesophageal echocardiogram showed the catheter in nearly the exact location of the previous impella cp, in the papillary structures. multiple attempts and maneuvers used to reposition, and, however, were unsuccessful. the impella cp pump continued to track to the same position in the papillaries causing mitral regurgitation. the patient was continued on low dose of primacor and epi. the decision was made to explant the impella but leave introducer in place in case an intra-aortic balloon pump (iabp) may be needed. the chest closed and dressed. however just prior to transferring the patient out of operating room, anesthesia noted the patient starting to require increasing doses of drips and rising lactate. the physicians requested one more attempt be made to place an impella cp device in catheterization lab to see if an optimal position could be achieved. however, this third pump attempt was unsuccessful as well. ultimately, the decision was made place an iabp for at least some mechanical support. the patient was reported to be stable following the event. | [5 empty] | The investigation of heart valve damage and low pump flow has been completed. the pump was returned for analysis. visual inspection found no issues on the pump. data logs confirmed low flows immediately after pump ramp up with flows less than one liter. suction alarms were confirmed. no motor current spikes outside of p level changes. the cause of the low flows issue and heart valve damage issue was most likely improper positioning of the device due to patient anatomy's structure. positioning issues is captured under low flow failure mode. as noted in the impella instructions for use: impella cp with smartassist system section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.9 revised as yes was inadvertently omitted from device returned to manufacturer? question on the initial manufacturer device report number 1220648-2024-25187. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25187 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25187. codes were added to health effect - impact code. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-25187. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. this is one of three medical device reports associated with this event. refer to medical device report for impella cp serial number (b)(6) (pump 1) with date of event 11-dec-2024/identifier ending in (b)(6) and manufacturing report number 1220648-2024-25159 for impella cp serial number (b)(6) (pump 3).
12595 | 1220648-2024-25191 | 21035267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21035267&pc=OZD | 12/10/2024 | 01/31/2025 |  | 12/30/2024 |  |  | 12/10/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579591 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, high purge pressure alarm occurred. tissue plasminogen activator (tpa) protocol was started. purge system blocked alarm occurred and tpa was infusing. the plan was made to exchange the pump. the patient survived the event. | [5 empty] | The investigation for the high purge pressure has been completed. data logs were reviewed and a purge pressure trend was confirmed, though there were no triggers of high purge pressure alarms. in the middle of the night on 12/10, the purge pressure began rising over the course of 3.5 hours through the end of the case. at the same time, the motor current started drifting upward without any shift in p-level. during that rise, there were distinct steps up in the motor current that align with spikes in the ps. that being said, there were no clear motor current spikes that correlated with the start of the purge pressure rise or any other abnormalities leading up to it. returned product was investigated and there was biomaterial in the outflow cage and purge gap. there were no kinks or other abnormalities noted on the pump/catheter. the pump and purge cassette were connected to a console and the purge pressure started to trend up, however, once the pump was started, the purge pressure resolved to normal levels during the 2 hours it was running. at the same time, the motor current was variable and pump flows were low during this entire run. a window was cut in the catheter and there was no blood ingress into the purge line, suggesting a purge kink was not the cause of the purge pressure rise. based on the data log review and returned product, the root cause of the high purge pressure was determined to most likely be biomaterial. added h6 impact codes 4644, 4629. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12596 | 1220648-2024-25192 | 21035425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21035425&pc=OZD | 12/11/2024 | 01/24/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576344 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and it would not cross the aortic valve due to patient anatomy. the implant was aborted. the patient remained stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12597 | 1220648-2024-25193 | 21035551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21035551&pc=OZD | 12/11/2024 | 01/06/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143529 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a battery failure during use in the catheterization lab. there was no documented patient harm resulting from the failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-25193 was submitted in accordance with updated procedures. h.6 code 4114 was reported incorrectly on initial manufacturer device report number 1220648-2024-25193 as the device was returned. h.10 revised additional manufacturer narrative as it was reported incorrectly on initial manufacturer device report number 1220648-2024-25193 as the device was returned.
12598 | 1220648-2024-25194 | 21035696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21035696&pc=OZD | 12/11/2024 | 01/28/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526495 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. transesophageal echocardiogram showed pigtail in papillary muscles. multiple attempts to reposition were unsuccessful. every attempt ended with impella tracking back into same location within papillary structures. the pump was removed and two other pumps were used. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the positioning issue has been completed. the reported situation was that the pump was providing less than optimal flows. the root cause of the low pump flows was most likely due to improper positioning. every repositioning attempt ended with impella tracking back into same location within patient condition's papillary structures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12599 | 1220648-2024-25195 | 21035977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21035977&pc=OZD | 04/26/2021 | 12/30/2024 |  | 12/30/2024 |  |  | 04/26/2021 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239154 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during an impella connect installation, the automated impella controller booted up with a battery failure alarm. despite troubleshooting, nothing resolved the battery failure alarm. there was no patient involvement. | [5 empty] | The automated impella controller (console) was received for investigation. upon inspection and start up, there were no bars on the battery display. event logs show that the console was not powered on (and most likely not plugged into ac) for over 9 months. the cause of the battery failure alarm was due to depleted batteries. this report is being filed as part of a retrospective review of historical records.
12600 | 1220648-2024-25196 | 21036846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21036846&pc=OZD | 12/11/2024 | 01/24/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560084 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with a impella cp device for mechanical circulatory support and experienced limb ischemia. the impella cp was placed using a right femoral access. there was no placement concerns noted and the patient received intervention to culprit left circumflex artery. the patient was staged to return for left anterior descending artery intervention. following the procedure, discussions were had regarding leaving the impella cp in overnight to allow the patient to rest. however after checking distal perfusion, there was reduced flow to the patient's right leg. the decision was made to wean and explant the impella cp device. the groin was closed with two perclose and an angioseal with achieved hemostasis noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12601 | 1220648-2024-25197 | 21036932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21036932&pc=OZD | 12/11/2024 | 01/28/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563826 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. the first pump experienced a pump stop (see case (b)(4). a new pump was difficult to position and caused a lot of ectopy. the doctor elected to pull the pump out, run it in sterile normal saline and attempted to be placed again over the previously used platinum plus. when advancing the pump over the wire, the pump was caught and the wire was "shredding". the entire system was removed and a new pump was prepped and placed. no harm to the patient was noted. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely use issue since non-abiomed product was used which against the recommendations per instructions for use (ifu). | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention . section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12602 | 1220648-2024-25198 | 21037284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21037284&pc=OZD | 12/09/2024 | 01/27/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support due to needing a high-risk percutaneous coronary intervention with the plan to wean and explant post percutaneous coronary intervention. after the procedure, the patient became hypotensive and had limited flow distal to the impella. the impella was weaned and explanted due to concern of the patient developing limb ischemia. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12603 | 1220648-2024-25199 | 21038616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21038616&pc=OZD | 12/12/2024 | 01/30/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530520 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Limb Pain | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient with complex cardiac and medical history. the pump was placed but there was limb pain and ischemia. the decision was made to explant and exchange pump but the family decided to withdraw care instead and the patient expired. | [5 empty] | The investigation for the pain at insertion site and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the pain at insertion site and limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12604 | 1220648-2024-25200 | 21039167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21039167&pc=OZD | 12/13/2024 | 01/28/2025 |  | 12/30/2024 |  |  | 12/13/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507106 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support the patient had ventricular tachycardia. the patient required cardioversion. the patient was given lidocaine. the staff stated that this was not a positioning issue but likely an issue with the patient's heart failing. support was continued and the patient survived the event. | [5 empty] | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12605 | 1220648-2024-25201 | 21039755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21039755&pc=OZD | 12/13/2024 | 01/31/2025 |  | 12/30/2024 |  |  | 12/13/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573210 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following insertion, an introducer became occlusive to the patient's vessel. the procedure was subsequently aborted as a result of the noted condition. the patient¿s outcome is unknown. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12606 | 1220648-2024-25202 | 21041113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21041113&pc=OZD | 12/11/2024 | 02/06/2025 |  | 12/30/2024 |  |  | 12/11/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535964 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and experienced pericardial effusion. the patient received two liters of blood, and the pump was explanted. the patient was reported to be in critical condition following this event. | [5 empty] | The investigation for the ventricle perforation has been completed. the pump was returned for investigation. a data log review was not required for the case. no physical damage was observed on the returned product. according to the clinical team, pericardial effusion (left ventricle perforation) occurred during impella use. time of the occurrence and the patient condition information were not provided. the cause of the ventricle perforation issue was not determined without sufficient clinical information. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ d.9 device returned date/information was added. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25202. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25202. h.6 code 4641 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25202. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-25202 in accordance with updated procedures. | The device has not been returned for investigation. upon completion of the investigation, a supplemental will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12607 | 1220648-2024-25157 | 21044414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044414&pc=PYX | 12/09/2024 | 09/30/2025 |  | 12/30/2024 |  |  | 12/10/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560902 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hypovolemia | [12 empty] | Required Intervention | Death |  |  |  |  | IMPELLA 5.5. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction with an ejection fraction of 10-15% was initially implanted with an impella cp device for mechanical circulatory support (mcs) and approximately two days later, was escalated to an impella 5.5 and added the impella rp flex support for the right ventricle. however, the patient would eventually expire. cardiothoracic surgery turned down the patient for coronary artery bypass graft surgery and recommended high-risk percutaneous coronary intervention (hrpci). ahf also consulted and recommended escalation to impella 5.5 from the impella cp (two days after start of this support) in which hemolysis was experienced. the physician made the decision to take the patient to catheterization lab for high-risk percutaneous coronary intervention (hrpci) and evaluated need for escalation after revascularizing. during the hrpci, the patient decompensated, had ventricular fibrillation arrest, and required vasopressor support. the cardiothoracic surgeon was called, arrived at the catheterization lab, and decided to implant impella 5.5 in the lab. venous bleeding was noted in impella 5.5 site axillary cutdown. surgicel utilized, packing, and blood transfused and oozing resolved. the patient had decreased flows on impella 5.5, persistent suction events, and the right ventricle was dilated on the transesophageal echocardiogram (tee). the patient continued to decompensate and had another ventricular fibrillation arrest while closing impella 5.5 site. decision made to emergently place impella rp flex. the patient was transferred to the unit on bipella (impella 5.5 and rp flex) support...continues in h10. | The complainant also reported that there were placement signal not reliable alarms. audio was disabled and support continued. |  |  |  |  | Two days later in the morning, it was noted the perclose at the initially implanted impella cp site (left subclavian artery) popped during the patient's coughing fit. manual pressure was held for one hour, packed red blood cells and platelets were given. the patient was taken to the cardiovascular operating room for repair. plasma free hemoglobin was down to 60. the patient went back to the operating room that afternoon for an angiogram of the left subclavian artery and bleeding was noted. a covered stent was deployed. further noted, the impella 5.5 was pointing towards the mitral on tee, the impella 5.5 was successfully repositioned and now pointing toward apex. bilaterally chest tubes, right side for pleural effusion and left side for hemothorax. hemoglobin was 7 g/dl with continuous bleeding from chest tubes, and a plan to give multiple blood products. there were suction events on bipella pumps noted likely from volume loss, and albumin and products were infused. next step plan: fix fluid status, monitor chest tube output, and [re]start continuous renal replacement therapy with a possible return to catheterization lab later the same day or the following day to evaluate coronaries. however, the patient would expire the following day. the patient had an ejection fraction of just 10-15% and was very unwell going into the procedure for the impella. the cause of death is unknown. there were impella-related events that occurred during impella mcs and there is not, however, enough information to exclude impella rp flex as an associated factor in the patient¿s outcome. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of three medical device reports associated with the event(s). refer to manufacturing report numbers 11220648-2024-25113 (for impella cp) and 1220648-2024-25097 (impella 5.5). | The investigation for the low pump flow has been completed. the pump was not returned for investigation, and the data log was not provided for this case. according to the clinical notes, the placement signal was lost on the second day, and suction alarms were reported on both devices, likely due to blood loss on the third day. the cause of the low pump flow issue was not determined since the product was not returned and data log was not provided for this case. b.7 hyperlipidemia was reported incorrectly on the initial manufacturer device report number 1220648-2024-25157. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25157 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25157. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the optical issue was not determined since product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal description. h6 updated to include placement signal coding.
12608 | 1220648-2024-25203 | 21044473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044473&pc=PYX | 12/12/2024 | 01/29/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025569703 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. while on support, a pump failure occurred. there was a spike in motor current and drop in flows then "impella stopped" alarm. pump ingested clot. the doctor removed and replaced rp flex with a new one. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop and thrombus has been completed. pump was not returned for investigation. data log shows the pump stop with motor current spike to 1600 after 30 min of case run. pump was explanted and implant new one. the cause of the pump stop issue was most likely biomaterial, based on data log review. the thrombus failure mode was removed as it was addressed under the pump stop failure mode. e4 should have been left blank on manufacturer device report 1220648-2024-25203. g1 reporting contact email revised on manufacturer device report 1220648-2024-25203 in accordance with updated procedures. g4 pma/510(k)number was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25203. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12609 | 1220648-2024-25204 | 21044499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044499&pc=OZD | 12/12/2024 | 01/31/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582223 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and during the implant, the pump would not cross the aortic valve. the catheter was bending and the staff could not advance the pump. additionally, when attempting to remove the 0.018" guidewire, it kinked and the device could not be removed through the sheath. the whole sheath and impella with wire had to be explanted and a new sheath and device was placed. there was no harm to the patient. | [5 empty] | The investigation of delivery issues and product damage (guidewire damage - great vessel) has been completed. no product was returned and logs are not applicable. due to the patient being obese with tortuous vessels and aortic stenosis, the root cause of the delivery issues was most likely patient-condition related. the failure mode "product damage (guidewire damage)" was removed since there were no guidewire fractures/detachments. e4 should have been left blank on manufacturer device report 1220648-2024-25204. g1 reporting contact email revised on manufacturer device report 1220648-2024-25204 in accordance with updated procedures. h6 medical device problem code 2682 was inadvertently omitted from manufacturer device report 1220648-2024-25204. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12610 | 1220648-2024-25205 | 21044592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044592&pc=OZD | 12/12/2024 | 12/30/2024 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support had compromised flow in their impella leg. the pump was being weaned for explant at the time the condition was noted. the pump was subsequently successfully explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
12611 | 1220648-2024-25206 | 21044598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044598&pc=OZD | 12/12/2024 | 01/16/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578887 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when impella cp support was initiated, the impella cannula had a kink just below the outflow. impella flows had reduced. the physician decided to proceed with the percutaneous coronary intervention (pci) procedure with impella support at p-6. impella flows and support monitored closely throughout duration of procedure. pci performed and impella was weaned and explanted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the product damage kink has been completed. the returned product was inspected and the cannula kink was observed. the clinical details do not mention any other failure modes. the root cause of the product damage (cannula kink) was most likely due to a kinked cannula.
12612 | 1220648-2024-25207 | 21044600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044600&pc=OZD | 12/12/2024 | 01/28/2025 |  | 12/30/2024 |  |  | 12/12/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [10 empty] | Disability | Life Threatening | Death | Required Intervention | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (with unknown patient-specific information other than sex) with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and an access site artery perforation occurred and led to the patient's demise. the patient was brought to the catheterization lab post code due to severe mitral disease. the impella was implanted bleeding at the sheath was noted. an angiogram was recommended to check for arterial damage and if no arterial damage move to change the sheet. the team stated the groin seemed stable after exchange to an impella long sheath (it is unknown if an angiogram was completed). the physician proceeded with the percutaneous coronary intervention via percutaneous coronary angioplasty of right coronary artery. at end of pci case, the patient had an uncontrolled hematoma. recommendations made for an angiogram to check for arterial damage and the angiogram revealed perforation of femoral artery. vascular surgery was called; however, the patient would code and expire in the catheterization lab. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of vascular damage - peripheral vessel has been completed. data logs are not relevant to the complication and the product was not returned for investigation. clinical details mentioned that the patient's femoral artery was calcified. however, the vessel was large enough for the sheaths. no other details were provided that would suggest that excessive force was use or any other patient conditions were the cause. based on limited clinical details provided, the root cause of the vascular damage could not be determined. b.2 revised selections as several were reported incorrectly on the initial manufacturer device report number 1220648-2024-25207. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25207 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address section as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-25207. g.7 should have been left blank on the initial manufacturer device report number 1220648-2024-25207. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-25207 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
12613 | 1220648-2024-25208 | 21044632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044632&pc=OZD | 12/11/2024 | 02/07/2025 |  | 12/30/2024 |  |  | 12/13/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513908 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and immediately after implant lactate dehydrogenase was found to be elevated. the support level was lowered from p-6 to p-4 which improved the condition. it is unknown, however, if there was additional treatment including blood transfusion to resolve the hemolysis. the patient was reported to be stable following the event; at end of support, the patient was successfully weaned, and the device was explanted. | [5 empty] | The investigation of hemolysis has been completed. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. e.1 revised address line 2 as information was inadvertently omitted from initial manufacturer device report number 1220648-2024-25208. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25208 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12614 | 1220648-2024-25209 | 21044634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044634&pc=OZD | 12/13/2024 | 02/03/2025 |  | 12/30/2024 |  |  | 12/13/2024 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566593 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the removal of the impella cp the physician cut down to remove impella device and placed 5 stitches to obtain hemostasis of the vessel. the physician stated that there was significant bleeding around the impella arteriotomy site. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely use issue since additional stitches were needed to obtain hemostasis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12615 | 1220648-2024-25210 | 21044654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044654&pc=OZD | 12/14/2024 | 02/24/2025 |  | 12/30/2024 |  |  | 12/14/2024 | 12/30/2024 |  |  | 10/31/2025 |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409450 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported an attempted impella 5.5 device implant was unsuccessful to a patient with postcardiotomy cardiogenic shock. a cardiac catheterization at an outside hospital one day prior to the impella implant attempt revealed multiple vessel disease with the inability to place an intra-aortic balloon pump due to the patient's difficult vasculature. the decision to made to place impella in right axillary artery. innominate artery showed calcification as well as left axillary artery and some calcification noted on right axillary artery. the physician was unable to place impella 5.5 device via direct approach due to small aorta and bypass cannula location. a right axillary cutdown was performed and a 10mm vascutek gel weave graft was sewn in place. the impella pump was primed greater than 15 minutes and torque removed. the physician was able to obtain ventricular access and unsuccessfully attempted to advance the impella 5.5 device over a 0.18 guidewire due to vascular anatomy. troubleshooting included attempts by another physician and buddy wire inserted for extra support. however, ultimately, the impella 5.5 device was unable to be advanced due to vascular anatomy and the case was aborted. patient outcome status and subsequent support/treatment is unknown. however, there was no report of harm or injury with report of the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy.
12616 | 1220648-2024-25211 | 21044743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044743&pc=OZD | 07/06/2022 | 12/30/2024 |  | 12/30/2024 |  |  | 07/06/2022 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136980 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | Black Or African American | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the automated impella controller (aic) had a controller failure alarm. the alarm self-resolved but the soft buttons would not work on the aic. the aic was switched out with another unit. there was no patient harm reported. | [5 empty] | Data analysis: automated impella controller (console) event logs revealed that on the complaint occurrence date, there was only a self-resolved controller error alarm (instead of controller failure). logs from the reported day of event are consistent with complaint and show multiple "high guard level" errors, indicating possible electrostatic discharge events (esd), which are known to be causing keyboard to be temporarily unresponsive. after 4 minutes of the keyboard being unresponsive the console was shut down. device analysis: the issue was not reproduced in (b)(6) and the console was found to be within specification and passing all testing. it is most likely that some undetermined esd event occurred, but aic did not receive any permanent damage. this report is being filed as part of a retrospective review of historical records.
12617 | 1220648-2024-25212 | 21044789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21044789&pc=OZD | 06/21/2023 | 09/18/2025 |  | 12/30/2024 |  |  | 06/21/2023 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464036 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an automated impella controller (aic) presented with an ¿optical sensor not supported alert message ¿. to resolve the issue, the console was exchanged for a backup. this complaint was related to a clinical procedure and there was no patient harm reported. | The complainant reported a patient in st elevation myocardial infarction was admitted to the hospital where it was discovered she had a total occlusion of the left anterior descending coronary artery (lad). the patient's lad was opened, the patient remained stable. prior to leaving the catheterization lab, the patient went into ventricular tachycardia and fibrillation where she was rescued by her automatic implantable cardioverter defibrillator. the patient was transferred to the icu where she continued to decline and decompensate overnight. the family wished to escalate care of the patient, and the physician decided to place the impella to wean the inotrope support as a last effort, and an impella cp device was implanted for mechanical circulatory support. it was reported upon connecting the impella cp to the automated impella controller (aic) with the serial number (b)(6), there was a message for "optical sensor not supported". the pump was disconnected and reconnected several times, however this did not resolve the issue. the aic was swapped with a backup unit, and the pump was successfully connected. it was noted after impella supported was initiated, the patient was in good and stable condition and was brought back to the icu. in the icu, the aic alarmed for placement signal low. stat echocardiography was ordered for the patient, and it was observed the impella cp was too deep within the ventricle. the physician was able to successfully pull the impella back per best practices, and "great" waveforms on the aic were noted. as reported on (b)(6) 2023, the patient continued to decline despite maximal mechanical and medical support with worsening liver and renal function. there was a need for escalation of vasopressor and inotropic support on the morning of (b)(6) 2023 with poor prognosis. it was noted that the impella was placed and functioning well, however, due to multiple metabolic derangements the patient continued to decline. the collective decision was made by the cardiology and pulmonary teams in conjunction with the family to transition the patient to do not resuscitate and comfort care measures only. therefore, the decision was made to withdraw mechanical circulatory support and the impella was removed at the bedside. the patient tolerated the removal well and hemostasis achieved. the patient later expired. although the patient's presenting condition may be more likely have impacted the event, the aic cannot be excluded as a possible contributing factor in the patient's demise after a malposition event of the impella which may have contributed to lower than expected support. the malposition of the impella may have been resultant of the optical sensor issue reported on the first aic that was in use. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information for the initial mfr 1220648-2024-25212 was discovered with the ongoing investigation. patient information was provided in sections a1, a2, a3, and a5. the patient expired during impella support, and therefore b1, b2, b5, and h1 have been updated to include this new information. upon completion of the investigation, a supplemental report will be filed. this is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2025-25520 for impella cp pump with sn (b)(6). | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the placement signal issue has been completed. console logs from the reported day of event are consistent with complaint detailed analysis of console logs revealed that prior to the clinical case. after console was received in danvers, the issue was reproduced. the root cause of the automated impella controller issue was a firmware corruption. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
12618 | 1220648-2024-25215 | 21045036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21045036&pc=OZD | 11/27/2021 | 12/30/2024 |  | 12/30/2024 |  |  | 11/27/2021 | 12/30/2024 |  |  |  |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 support, there was kinking in the drive line. it was noted there was a kink in the impella cable before the red connector. there were no alarms or running issues. abiomed support recommended stabilizing the cable with tape and a mouth spatula. there were no further issues with the cable, the patient remained on support until successfully weaned off for transplant surgery. | [5 empty] | D4 serial, lot, udi information is unknown. h4 manufacturing date is unknown. this is one of two medical device reports associated with the event. refer to initial medwatch with date of report (b)(6)2024 for impella 5.5 serial number (b)(6). this report is being filed as part of a retrospective review of historical records.
12619 | 1220648-2024-25216 | 21045041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21045041&pc=OZD | 11/27/2021 | 11/11/2025 |  | 12/30/2024 |  |  | 11/27/2021 | 12/30/2024 |  |  | 06/30/2023 |  | 12/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117243 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 support, there was kinking in the drive line. it was noted there was a kink in the impella cable before the red connector. there were no alarms or running issues. abiomed support recommended stabilizing the cable with tape and a mouth spatula. there were no further issues with the cable, the patient remained on support until successfully weaned off for transplant surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the event. refer to initial medwatch with mfr 1220648-2024-25215 for the automated impella controller with unknown serial number. this report is being filed as part of a retrospective review of historical records.
12620 | 1220648-2024-25149 | 21025783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025783&pc=OZD | 12/09/2024 | 01/27/2025 |  | 12/27/2024 |  |  | 12/09/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522426 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding. one unit of red blood cells were administered to the patient, manual pressure was applied, and the dressing was redone. the bleeding did not show improvement. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d.6b revised as explant date was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25149. e.1 revised facility name as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-25149. g.1 revised manufacturing site fax number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25149. h.6 code 4641 was entered incorrectly on the initial manufacturer device report number 1220648-2024-25149. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12621 | 1220648-2024-25150 | 21025809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025809&pc=OZD | 12/09/2024 | 01/29/2025 |  | 12/27/2024 |  |  | 12/09/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532096 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. patient anatomy prevented the delivery. the procedure was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella aborted due to vascular anatomy unable to pass through iliac artery. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12622 | 1220648-2024-25151 | 21025819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025819&pc=OZD | 12/10/2024 | 02/11/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569422 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon explant, thrombus was found to have formed above the motor and inlet. it was noted that the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the pump was returned. as all the necessary information was not provided, the root cause of the thrombus was not determined. e.1 revised facility name and city as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-2515. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25151 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact address line 1, reporting contact us city, reporting contact state, reporting contact us zip code and reporting contact country as they were inadvertently omitted from initial manufacturer device report number 1220648-2024-2515.
12623 | 1220648-2024-25152 | 21025822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025822&pc=OZD | 12/10/2024 | 01/30/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570211 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. per the study team, the patient experienced a stroke post explant and thrombus was also identified. actions taken are unknown. no further information was provided. | [5 empty] | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: d9 date device returned to mfr should have been left blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12624 | 1220648-2024-25153 | 21025824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025824&pc=OZD | 05/26/2021 | 02/03/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  | 03/31/2023 |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced possible anemia / thrombocytopenia during impella cp support. no noted intervention was documented regarding the reported condition. it was additionally reported that the patient experienced ischemia following impella cp removal. vascular surgery was performed to resolve the condition. the patient survived the procedure. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombocytopenia and limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the thrombocytopenia cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25153 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25153 and should not have been.
12625 | 1220648-2024-25154 | 21025828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025828&pc=OZD | 12/10/2024 | 02/04/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211686 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) catheter plug port was not recognizing the pump when it was plugged in. it is unknown when this issue was discovered. no further information was provided. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. multiple error messages "invalid_bad_snr_sample_mask," caused by the signal-to-noise ratio (snr) being too low was observed. the first "placement signal not reliable" error was seen. later, the pump was reconnected to the console twice, and the error message "case start: optical sensor system error: reconnect catheter" was observed, along with the peak snr alarm. console was shutdown moments after and likely exchanged. the console was returned for further inspection. while the optical signal issues with a known-good test pump could not be reproduced, an inspection of the analog output from the ccd on the optical bench revealed intermittent signal distortions on the envelope/fringe. the measured "5s" parameters confirmed that the snr was below the threshold. additionally, a visual check found that the blue purge flag retainer had a hairline fracture. the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. the root cause for the broken purge retainer was not determined. e.1 revised name prefix/title as it should not of been left blank on the initial manufacturer device report number 1220648-2024-25154. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25154 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12626 | 1220648-2024-25155 | 21025954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025954&pc=OZD | 12/10/2024 | 01/28/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519218 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump came out of the appropriate position; therefore, the pump was explanted. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was most likely wireless re-access as impella pulled out of the appropriate position and wire-less re-access attempted to cross the valve. the impella device is not indicated for wireless re-access. corrections to the initial mfr report: d9 date device returned to mfr should have been left blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12627 | 1220648-2024-25156 | 21025964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025964&pc=OZD | 12/10/2024 | 01/08/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574094 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was placed, but there were positioning issues noted and ectopy. the team lost left ventricle positioning while attempting to troubleshoot. the pump was explanted without further escalation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. corrections to the initial mfr report: b1-should have selected product problem
12628 | 1220648-2024-25158 | 21025979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025979&pc=OZD | 12/09/2024 | 02/04/2025 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568907 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the doctor states that a patient on ecmella was on support for 3 days with no problems, for unexplained reasons suction alarms were occurring with lower pump flows. the doctor asks if a thrombus in the cannula could be responsible. the pump was explanted. the event had no effect on the patient's outcome. additional information noted care was later withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of saturated flow has been completed. the cut pump was sent back for investigation. biomaterial was observed on the impeller and purge gap. no further investigation was performed due to the cut pump. data logs were not provided for this case run. according to the clinician, the pump was explanted after suction alarms began occurring with low pump flows. the reported failure mode of saturated pump flow was changed to low pump flow as the investigated failure mode, based on the given clinical details. the cause of the low pump flow issue was most likely biomaterial, based on returned product review. d9 date device returned to manufacturer added. d6a implant date was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25158. e4 should have been left blank on manufacturer device report 1220648-2024-25158. g2 report source; foreign was missing from manufacturer device report 1220648-2024-25158. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12629 | 1220648-2024-25159 | 21025980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025980&pc=OZD | 12/11/2024 | 01/30/2025 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531257 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. immediately upon starting pump, suction alarms occurred; unable to flow above p1. a kink was noted at the proximal end of the cannula. all attempts to reposition were unsuccessful. pump was removed, run in a basin and re-inserted but same result. ultimately, the pump was removed and place an intra-aortic balloon pump for at least some mechanical support. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. the returned cp pump was visually inspected. a cannula kink observed. biomaterial observed in purge gap. no broken blade found. no other abnormality seen. data log reviewed: no mc spike or other trend seen. it looks obstruction occurred at p-1 in flow because of the cannula kink, and then positioning alarms occurred while attempted to reposition. suctions occurred. this complaint pump is 3rd pump from the same patient/case. both failure modes (fm) (positioning issue + product damage) changed to fm: low pump flow issue based of data observation with flow issue. the root cause of the low pump flow issue was patient condition related due to patient's small aortic arch/ventricle likely made pump move out of position. e4 should have been left blank on manufacturer device report 1220648-2024-25159. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12630 | 1220648-2024-25160 | 21025988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21025988&pc=OZD | 12/10/2024 | 01/29/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569404 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a 5.5 pump support ongoing when the team observed elevated purge pressure (pp) alarms. the team attempted to troubleshoot by adding tissue plasminogen activator to lyse in the purge fluid but, this measure did not resolve. the pp persisted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the returned product was inspected and biomaterial was found in the purge gap by the impeller. the data logs confirm high purge pressure alarms and show a gradual rise in purge pressure before alarm was triggered. the biomaterial was ejected from the purge gap and was more visible after the lab run. the root cause of the high purge pressure was biomaterial. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-25160.
12631 | 1220648-2024-25161 | 21028281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21028281&pc=OZD | 03/19/2023 | 11/19/2025 |  | 12/27/2024 |  |  | 03/19/2023 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support that the automated impella controller¿s (aic) battery level dropped from 100% to 50%. it was suggested that the physician swap out the console to address the issue, however the physician declined. it was unknown what aic unit had the battery issue. | [5 empty] | D4: udi, serial, model, catalog and lot numbers are unknown. h4: manufacturing date is unknown. the automated impella controller was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12632 | 1220648-2024-25162 | 21028572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21028572&pc=OZD | 10/13/2021 | 12/27/2024 |  | 12/27/2024 |  |  | 10/13/2021 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153987 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an alarm occurred during a training demonstration that the abiomed representative was unfamiliar with. there was no patient involvement. | [5 empty] | The investigation for the reported automated impella controller (console) issue has been completed. complaint date is consistent with logs, on 10/13/2021 errors were noted in the event log. the console was booted and the alarms were reproduced. the console was opened and pbm was inspected. it was observed that capacitors c69 and c70 on the top side of the pbm had leaked and corroded the board. specifically noting that there was corrosion that leaked to a solder through hole and reached the underside of the board. failure mode observed in the field was due to a malfunctioning pbm. the root cause of the malfunctioning pbm is due to capacitors c69 and c70 leaking and corroding the pcb this report is being filed as part of a retrospective review of historical records.
12633 | 1220648-2024-25163 | 21028674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21028674&pc=OZD | 10/27/2021 | 12/27/2024 |  | 12/27/2024 |  |  | 10/27/2021 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263046 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the automated impella controller had an alarm that the medical team described as a ¿weird sounding alarm¿. the aic was not replaced due to the issue and there was no patient harm reported. | [5 empty] | Data analysis: automated impella controller (aic) event logs were unavailable (overwritten) the day after the reported complaint date. aic logs showed that the initial device handshake was successful but later reported that ¿icm not ready for file receive¿. device analysis: the "weird sounding alarm noise" reported in complaint was most likely due to pbm malfunction. this report is being filed as part of a retrospective review of historical records.
12634 | 1220648-2024-25164 | 21028755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21028755&pc=OZD | 05/23/2022 | 12/27/2024 |  | 12/27/2024 |  |  | 05/23/2022 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502212 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during impella 5.5 support, the automated impella controller (aic) alarmed for purge pressure low. the purge cassette was attempted to be replaced; however it was unsuccessful due to a broken purge disc on the aic. as a result of the broken purge disc, the aic was replaced. there was no patient harm reported. | [5 empty] | Data analysis: during the clinical procedure, the impella 5.5 pump sn (b)(6) was disconnected and reconnected to automated impella controller (console) ic8266 to then have eeprom reading issues and the software process became unresponsive while re-reading the eeprom. the console software watchdog detected it, and performed a system restart, as a design intended response to such event. after the restart, as the console booted up, pump sn (b)(6) was still connected, and the console had issues reading the usage section of the eeprom. regardless, the pump was successfully restarted upon reboot. later in the case, the console issued a purge fluid not detected error window due to a broken retainer. device analysis: upon arrival, the console ic8266 was inspected, and the purge pressure sensor retainer was found to be broken. no other issues observed during testing. the cause of the purge fluid not detected error was due to a broken purge retainer. this report is being filed as part of a retrospective review of historical records.
12635 | 1220648-2024-25165 | 21029247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21029247&pc=OZD | 09/24/2021 | 12/27/2024 |  | 12/27/2024 |  |  | 09/24/2021 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090364 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the complainant that the automated impella controller was having connectivity issues. during the investigation of the device, it was discovered that the battery was defective. there was no patient involvement. | [5 empty] | Device analysis: the console was tested thoroughly on an engineering lab bench. the complete failure of the handshake mechanism was not reproduced. while in storage in danvers, the console's battery failed due to cell imbalance. the battery failure was due to a defective battery 1 (serial no: (b)(6) ). the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
12636 | 1220648-2024-25166 | 21029272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21029272&pc=OZD | 04/07/2021 | 12/27/2024 |  | 12/27/2024 |  |  | 04/07/2021 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp with smart assist was connected to the automated impella controller with serial number (b)(6), but the screen to connect the grey connectors appeared as if it wasn¿t plugged in. there was no patient harm reported. | [5 empty] | Data analysis: given the data logs were deleted by the sw upgrade to v8.5, an image of console¿s log download screen showed the console did not generate a log file for the impella and likely did not properly recognize it (see image attached). device analysis: the cause of the failure to read the pump eeprom was most likely due to a software defect (in-house). however, there is a chance that this could be an intermittent glitch of the impellatronic¿s epm circuitry, which could be resolved by simply power cycling the epm. console was upgraded by field team to sw v8.5. this report is being filed as part of a retrospective review of historical records.
12637 | 1220648-2024-25167 | 21029613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21029613&pc=OZD | 12/04/2021 | 12/27/2024 |  | 12/27/2024 |  |  | 12/04/2021 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116347 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the complainant that the automated impella controller was having connectivity issues. during the investigation of the device, it was discovered that the purge disc retainer had snapped off. there was no patient harm reported. | [5 empty] | Data analysis: imc logs are not relevant to this failure mode. device analysis: console arrived in danvers. purge assembly was inspected. it was noted that the purge disc retainer had snapped off. this report is being filed as part of a retrospective review of historical records.
12638 | 1220648-2024-25168 | 21029888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21029888&pc=OZD | 01/28/2022 | 12/27/2024 |  | 12/27/2024 |  |  | 01/28/2022 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117647 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by the complainant that the automated impella controller was having connectivity issues. during the investigation of the device, it was discovered that the console unexpectedly crashed and restarted during the clinical procedure. there was no patient harm reported. | [5 empty] | Automated impella controller (console) logs from the reported day of event are consistent with complaint and show that on (b)(6) 2022 data streaming was disabled, as per software design, after the console unexpectedly crashed and restarted. there was no deficiency of the rlm indicated in aic logs or rlm journal. prior to analyzing console logs, its timestamps had to be recalculated due to the fact that real-time clock on the 4-in-1 was not keeping the time due to coin-cell battery failure. correct time reference was established by comparing characteristic events in aic logs (e.g. handshake) with corresponding entries in the rlm journal. the issue of either unexpected crash or sudden power loss to the epc (while power to the rlm was uninterrupted - as evidenced in rlm journal) was not reproduced during extensive testing, but the most likely component responsible for this type of malfunction is either the epc itself or the carrier board which provides power to the epc and to the rlm via usb ports. this report is being filed as part of a retrospective review of historical records.
12639 | 1220648-2024-25169 | 21030256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030256&pc=OZD | 09/10/2024 | 02/13/2025 |  | 12/27/2024 |  |  | 12/05/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510636 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced post-operative bleeding requiring multiple transfusions. the patient received two units of packed red blood cells, 2 units of platelets, one fresh frozen plasma and 10 cryo overnight plus protamine. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | H.6 health effect - impact code 4641 was inadvertently omitted from manufacturer device report number 1220648-2024-25169 and was added. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the vascular damage - peripheral vessel was not determined since the product was not returned and there is not enough clinical information. d.1 revised brand name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25169. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25169. e.1 revised us phone number as it was inadvertently omiited from initial manufacturer device report number 1220648-2024-25169.
12640 | 1220648-2024-25170 | 21030309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030309&pc=OZD | 08/11/2023 | 01/24/2025 |  | 12/27/2024 |  |  | 12/09/2024 | 12/27/2024 |  |  | 04/30/2025 |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was successfully supported with the impella. the impella was weaned at the end of the procedure and decision was made to place an intraortic balloon pump. after support was concluded, a notification was received indicating that the patient endured ventricular fibrillation with an unknown (undetermined) relationship to the impella device used. additional information noted the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25170.
12641 | 1220648-2024-25171 | 21030310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030310&pc=OZD | 12/10/2024 | 01/23/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578315 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was brought to the cardiac catheterization lab where it was found that the patient¿s right coronary artery was down with ostial to mid lesion. lesion ballooned and the patient began to become unstable . impella called for and cardiopulmonary resuscitation initiated. after impella placed and return of spontaneous circulation achieved, the patient began to have suction events and the patient was not pulsatile. swan was placed, and pulmonary artery pulsatility index was 0.27. echocardiography was ordered and tamponade discovered. pericardiocentesis drain placed. after 1.5l removed, the patient was taken to the operating room (or) for exploration of the chest. it was found in the or that the impella pigtail was protruding through the ventricular wall. the patient expired in the operating room. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of ventricle perforation has been completed. the cause of the ventricle perforation issue was not determined since product was not returned and sufficient clinical information was not provided. e4 should have been left blank on manufacturer device report 1220648-2024-25171. g1 reporting contact fax number missing on manufacturer device report 1220648-2024-25171.
12642 | 1220648-2024-25172 | 21030357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030357&pc=OZD | 12/10/2024 | 01/28/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  | 11/30/2025 |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428186 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 cannula was not able to pass the anastomosis of axillary artery. after several unsuccessful troubleshooting techniques, the 5.5 insertion was aborted. the patient's outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass beyond anastomosis and the insertion was aborted.
12643 | 1220648-2024-25173 | 21030373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030373&pc=OZD | 12/10/2024 | 01/28/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571326 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the pump became malpositioned. the device was repositioned successfully but the patient experienced ventricular tachycardia (vt). the patient was put on amio and systemic heparin. the patient survived the event. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12644 | 1220648-2024-25174 | 21030432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030432&pc=OZD | 12/10/2024 | 01/31/2025 |  | 12/27/2024 |  |  | 12/10/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the pump was repositioned due to occasional ectopy. high purge pressure alarms were also noted. cassette and purge were replaced by the bedside nurse. pressure normalized, but closer to the upper limit and flow was in normal range but very low limit. tissue plasma activator use began at bedside and pressures dropped and flows increased on the purge. there were no further alarms. the patient continued with impella support. | [5 empty] | The investigation of vf/vt/af/at and high purge pressure has been completed. the impella device was not returned for evaluation. according to the data log and clinical details the root cause of the high purge pressure was most likely a use related issue as a prolong bag/cassette change occurred immediately before purge pressure began to increase. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6b explant date added e4 should have been left blank on manufacturer device report 1220648-2024-25174. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12645 | 1220648-2024-25175 | 21030483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030483&pc=OZD | 10/31/2024 | 12/27/2024 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1751012 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [18 empty] | IMPELLA CP | [18 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. as previously reported for the impella cp: while on support, the flows were down on the impella despite additional volume. additionally, the placement signal and flow metrics on the automated impella controller (aic) would not populate during and after insertion of the impella cp. the staff disconnected and reconnected the white power plug but noticed no change. the decision was made to exchange the pump for another impella cp. the following pump worked as expected. the patient coded, the family made the patient do not resuscitate, care was withdrawn and the patient expired. there is not enough information to exclude the impella cp device as an associated factor in the patient's demise. during investigation of the reported complaint against the impella cp pump, additional information noted the automated impella controller presented an optical signal issue. this is one of two related reports for the reported event. this report is for the automated impella controller device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12646 | 1220648-2024-25176 | 21030625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030625&pc=OZD | 12/11/2024 | 01/30/2025 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025549927 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in and having decisions being made. the team was considering left ventricular assist device or transplant. the patient was transferred to the tertiary center and extra corporeal membrane oxygenation was added. the pump stopped after 10 days. the pump was explanted and replaced. the patient survived the event. | [5 empty] | The investigation for the pump stop has been completed. pump was not returned for investigation. data logs were investigated and there was a motor current spike observed at the end of the case along with purge pressure rise and purge flow drop. the pump was tried restarting but motor current spike and similar pattern observed leading to pump stop. therefore, the root cause of the pump stop issue was most likely biomaterial. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12647 | 1220648-2024-25177 | 21030653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030653&pc=OZD | 12/11/2024 | 02/07/2025 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509360 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a 5.5 support ongoing for a patient admitted in cardiogenic shock. the pump had purge pressure alarms, and the purge system blocked after 6 days. the pump remains on for support with tissue plasminogen activator lysis in the purge. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was returned for investigation. data log was not provided for investigation. no physical abnormalities were reported on the returned pump. priming was successful with the test purge cassette, and no issues were encountered with purge pressure. according to the clinician, the purge system block alarm appeared after approximately 144 hours of operation and was successfully resolved after lysis therapy. the cause of the high purge pressure issue was most likely biomaterial, as the clinical team suggested that the lysis therapy successfully resolved the issue, and the issue was with the returned product. d6b explant date added. d9 date device returned to manufacturer added. g2 report source; foreign was missing from manufacturer device report 1220648-2024-25177. e4 should have been left blank on manufacturer device report 1220648-2024-25177.
12648 | 1220648-2024-25178 | 21030730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030730&pc=PYX | 12/11/2024 | 01/15/2025 |  | 12/27/2024 |  |  | 12/11/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517160 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support. while on support, there was a concern for hemolysis. there was increased plasma free hemoglobin and lactate dehydrogenase (ldh). impella was removed due to hemolysis concerns. the patient survived the event. | A notification that was received on 12/22/2024 via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of thrombus in outflow : ¿thrombus in the outflow¿. the patient outcome is unknown. |  |  |  |  | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added: b5-updated narrative with failure mode thrombus g3-new aware date h6-clinical code 4440.
12649 | 1220648-2024-25179 | 21030734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030734&pc=OZD | 12/12/2024 | 01/30/2025 |  | 12/27/2024 |  |  | 12/12/2024 | 12/27/2024 |  |  | 11/30/2025 |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424971 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for impella implant, a sterility issue was encountered with the catheter. a second pump was subsequently utilized for implant. there was no patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of packaging issues- sterility has been completed. the product was returned and evaluated. the product was not returned inside the packaging. no abnormalities were found with the pump. no pictures of the packaging were returned. the root cause of the sterility issue was not determined since the product was returned outside of the packaging and insufficient clinical detail was provided. d.9 date device returned/information was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25179 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 2199 was reported incorrectly on initial manufacturer device report number 1220648-2024-25179 and a new code was added.
12650 | 1220648-2024-25180 | 21030744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21030744&pc=OZD | 12/12/2024 | 01/31/2025 |  | 12/27/2024 |  |  | 12/12/2024 | 12/27/2024 |  |  |  |  | 12/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the physician attempted to place the impella cp via the right axillary, and it was unsuccessful. the impella was discarded. the patient is stable. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through the vasculature as per the clinical description provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25180 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12651 | 1220648-2024-25101 | 21017797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21017797&pc=OZD | 02/17/2023 | 10/13/2025 |  | 12/26/2024 |  |  | 02/17/2023 | 12/26/2024 |  |  | 12/31/2024 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293170 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support after decompensating status post left anterior descending balloon and stent placement. the patient plan of care was to transfer to another hospital for escalation of care and to escalate patient from the impella cp to an impella 5.5. it was reported on (b)(6) 2023 that during patient transfer to another hospital, the automated impella controller (aic) was swapped. it was reported immediately after swapping the aic, there was a purge system failure alarm. to resolve the issue, another aic was exchanged. support was delivered to the patient without interruption and there was no patient harm reported for the reported aic issue. additionally, on (b)(6) 2023, the patient experienced signs of hemolysis including red colored urine and a haptoglobin of 55 (no units reported) and ldh of 2600 (no units reported). the patient's plan of care was to escalate to the impella 5.5, and the impella cp p level was lowered to address the hemolysis. later that day on (b)(6) 2023 the patient was taken to the operating room to explant the impella cp and escalate the patient to the impella 5.5. the impella 5.5 was first placed in the left ventricle with the impella cp still placed. it was noted that the physician was unable to explant the impella cp due to a clotted off rewire access port. the decision was then made to perform a surgical cutdown and vascular closure using a bovine patch. it was reported that the patient was experiencing intermittent bradycardia, so a transvenous pacer was placed. the impella cp was successfully explanted, and impella 5.5 support was initiated. on (b)(6) 2023 it was noted that the patient was confirmed covid positive and the patient was treated for a mixed picture of septic and cardiogenic shock. on (b)(6) 2023 systemic heparin was discontinued as the patient was bleeding at iv lines. the patient was noted to have to arrhythmias on (b)(6) 2023. the patient was noted as not being a candidate for advanced therapies. on (b)(6) 2023 the patient care was withdrawn and placed on comfort measures only. the patient expired while on impella 5.5 support. the impella cp and impella 5.5 relationship to the patient's demise may be unlikely but cannot be excluded. the patient presented in a decompensating state and the patient was noted to be covid positive. the impellas cannot be excluded as a possible contributing factor given the impella related events of hemolysis, bleeding, bradycardia,and sepsis as these may have added to an already complex situation even though the patient was not a candidate for advanced therapies. | [5 empty] | This is one of three medical device reports associated with the event. refer to initial medwatch for pump number one the impella cp with mfr 1220648-2024-25098 and initial medwatch for the automated impella controller with mfr 1220648-2024-24895. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
12652 | 1220648-2024-25107 | 21018110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018110&pc=OZD | 12/05/2024 | 01/23/2025 |  | 12/26/2024 |  |  | 12/05/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025559462 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Failure to Sense | [5 empty] | Abscess | Sepsis | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient developed an abscess at the access site. the patient had a fever with positive cultures. multiple antibiotics were started. support continued. furthermore, there was high purge pressure reported. tissue plasma activator was added to the purge and support continued. additionally, there were placement signal alarms where it stated impella in aorta. transesophageal echocardiogram was performed verified good pump position across the aortic valve. repositioning attempt was made to try and clear the alarm, but the alarm continued throughout the day. concern was brought forward to the team that the pump was alarming incorrectly, and the underlying cause may be flow saturation/biomaterial ingestion. the decision was made to remove the pump due to these concerns. impella 5.5 was removed and an intra-aortic balloon pump was placed in left femoral. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation infection/sepsis, thrombus, placement signal issue, and high purge pressure has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of thrombus and infection was not determined. the cause of the optical signal issue and high pressure purge was most likely biomaterial interaction leading to false position in aorta alarms per log and clinical review
12653 | 1220648-2024-25106 | 21018111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018111&pc=OZD | 12/07/2024 | 01/16/2025 |  | 12/26/2024 |  |  | 12/07/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had a stroke. a computed tomography scan showed ischemic conversion to a hemorrhagic stroke. it was further reported that the patient arrested overnight and the patient was a do not resuscitate therefore the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12654 | 1220648-2024-25108 | 21018161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018161&pc=OZD | 12/08/2024 | 02/07/2025 |  | 12/26/2024 |  |  | 12/08/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575835 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient experienced a loss of pedal pulse. the right foot was turning blue and cool to the touch. reperfusion cannula was inserted and the pulse returned to the patient¿s foot. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12655 | 1220648-2024-25109 | 21018192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018192&pc=OZD | 11/24/2024 | 02/11/2025 |  | 12/26/2024 |  |  | 12/06/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025498062 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had a significant decline in cardiac function and extra corporeal membrane oxygenation was added. the next day, an echocardiogram showed no aortic valve opening, so impella cp was inserted. after cp was added airway bleeding was observed and disseminated intravascular coagulation (dic) occurred, so bleeding control was poor, and output could not be maintained. a blood transfusion was administered but the patient expired of multiple organ failure. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the vascular damage has been completed. clinical information provided mentions that there was bleeding but the relationship to impella is unknown. therefore, the root cause of the vascular damage issue was not determined since product was not returned.
12656 | 1220648-2024-25110 | 21018213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018213&pc=OZD | 12/08/2024 | 01/29/2025 |  | 12/26/2024 |  |  | 12/08/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571954 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a failed delivery attempt to place an impella cp device for mechanical circulatory support. the patient arrived at the hospital, and a computed tomography angiography revealed patient to be having an aortic dissection. the patient was accepted and taken to the operating room and a bentall procedure was completed. the patient required venoarterial extracorporeal membrane oxygenation (va ecmo) cannulation to come off bypass and following was transported back to intensive care unit. after being back in the unit for approximately one hour, the decision was made to consult cardiology for placement of impella cp pump. computed tomography angiography of the left common femoral artery showed no contraindication for the impella insertion. after the sheath was advanced, the physician proceeded with trying to obtain left ventricle access across the aortic valve for about 1.5 hours. multiple steps and techniques taken including giving calcium for increased ejection, going up and going down on ecmo flows, and attempting to pace the patient at a faster rate. aortic valve angiogram revealed no contrast was making it into the left ventricle and appeared that the aortic valve was not opening. multiple physicians were consulted for advice. a transthoracic echocardiogram was completed and noted the new bioprosthetic aortic valve had already developed a small thrombus, and the mitral valve had a large thrombus. decision was made to abort impella insertion with plans to take the patient back to the operating room to place a left ventricular vent. however, after consulting with the patient's family the decision was made to transport the patient back to the intensive care unit and bolus/infuse heparin to see if clot thrombus could be resolved. after further discussion with the family, it was noted the family likely will decide to withdraw care. however, no update to indicate the patient's demise have been received at this time. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issues has been completed. the pump was retuned for investigation and no abnormalities were noted. the root cause of delivery issue was determined to be patient condition. the guide wire was unable to pass the aortic valve because the valve was not opening and the insertion was aborted. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25110. a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-25110 in accordance with updated procedures.
12657 | 1220648-2024-25111 | 21018224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018224&pc=OZD | 12/06/2024 | 01/08/2025 |  | 12/26/2024 |  |  | 12/06/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562293 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and a bridge to transplant. approximately 10 days after start of support, the report noted the patient experienced pain at insertion site and down right bicep, discoloration noted to chest and down right arm. the patient taken to operating room where muscle bleeding (generalized) was found and as well as small arterials that were clipped or cauterized. no graft anastomosis bleeding was found. the graft to hub was intact with no signs of bleeding. cautery and anticoagulation were utilized it control, stop bleeding. the site closed with jp drain, and there was a plan to stop heparin for 24 hours. the patient remained on support and was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12658 | 1220648-2024-25112 | 21018228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018228&pc=OZD | 09/02/2022 | 09/25/2025 |  | 12/26/2024 |  |  | 09/02/2022 | 12/26/2024 |  |  | 11/30/2023 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022155727 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | Tachycardia | [9 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient cannulated with extracorporeal membrane oxygenation was implanted with an impella 5.5 for escalated mechanical circulatory support. during patient support visit, (6 days after start of support) it was noted placement signal (ps) and motor currents (mc) were invisible for 1-2 minutes and there were no alarms. the issue resolved itself; all curves became visible again and pump running. it is unknown the time in the day the event occurred. however, it is noted in an emergency call center log that the automated impella controller (aic) was scheduled to be replaced (b)(6) 2022 morning around 9:00am. no further information/details are available surrounding this event. 6 days after the initiation of impella support, on (b)(6) 2022, the physician reported that the patient had an upper gastrointestinal tract bleeding, with the question if heparin could be removed and if so for how long. the patient's partial thrombin time (ptt) was reported to be elevated at around 70 seconds. abiomed clinical support recommended first to remove systemic heparin and to reduce heparin in the purge solution if necessary to a minimum and to closely control the ptt or activated clotting time (act) and the purge pressure. possible consequences when removing heparin were discussed. the patient outcome as a result of this event is unknown. 13 days later on (b)(6) 2022, there was renewed increase in inflammatory parameters with suspicion of sepsis. this complication was managed with a change in antibiotics. however, the complication was not successfully managed (as noted). the patient experienced renal failure due to the sepsis and was placed on dialysis. it was also reported the patient experienced episodes of ventricular tachycardia. although inflammation markers and liver values were improving, the patient would expire two days later on (b)(6) 2022. given the totality of the information although lacking details regarding the patient's status pre-impella support and while on support, the impella cannot be conclusively excluded as a contributing factor to the outcome given the impella related events experienced. | [5 empty] | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-24998 for the automated impella controller with serial number (b)(6). the impella automated controller device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation of vascular damage - peripheral vessel, infection/sepsis, vf/vt/af/at, and renal failure has been completed. this complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
12659 | 1220648-2024-25113 | 21018276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018276&pc=OZD | 12/08/2024 | 01/30/2025 |  | 12/26/2024 |  |  | 12/08/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500731 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient experienced hemolysis. the patient was escalated to an impella 5.5. after impella cp removal it was noted that the site popped following a coughing episode. manual pressure was held for an hour and the patient was administered packed red blood cells and platelets. the patient was subsequently taken for to the operating repair. a covered stent was deployed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major most likely was patient movement related while patient coughing. the cause of the hemolysis issue cannot be determined since the complaint device not returned for analysis and data logs not available to review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12660 | 21018283 | 21018283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018283&pc=OZD | 12/05/2024 | 12/18/2024 | 12/18/2024 | 12/26/2024 | 12/18/2024 | 12/26/2024 |  | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 2025501405 |  | 0048-0003 |  | No answer provided |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Middle-aged male with a pmhx (past medical history) hypertension and hyperlipidemia who presents as a stemi (st-segment elevation myocardial infarction) now s/p (status post) des x2 (drug-eluting stent) to lad (left anterior descending artery) with a reduced ef (ejection fraction) and impella cp placement with pump failure taken emergently to the or (operating room) now s/p impella cp removal, primary repair of right cfa (common femoral artery) and iliofemoral thrombectomy on inotropic and vasopressor support. per doctor (hvi ccu attending of record at the time), he doesn't believe that anything related to our care was unusual or contributed to the device failure. regarding if it was the pump that malfunctioned versus the console, he believes it was the pump that malfunctioned as far as he can tell and the abiomed rep agrees. the console seemed okay but was showing data consistent with impending pump failure and then the pump failed. he doesn't think there is a way the console could cause that. regarding additional care he cannot identify anything other than the timing of the removal. the patient required a bit of vasopressors during and after the case but that could have happened later too. it could be theoretically argued that the patient's heart might have benefited from longer duration of unloading, but there's no evidence of that in his case. overall, he doesn't think the patient was harmed and it may have helped the patient getting the device out faster.
12661 | 1220648-2024-25114 | 21018360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018360&pc=OZD | 12/08/2024 | 01/28/2025 |  | 12/26/2024 |  |  | 12/08/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573297 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, hematoma was noted at impella site. partial thromboplastin time was greater than 80. the patient also experienced bleeding from the nose. heparin was withheld and the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding ¿ major has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was patient condition related to coagulation disorder.
12662 | 1220648-2024-25115 | 21018429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018429&pc=OZD | 12/09/2024 | 01/27/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536194 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. upon removal, the doctor noticed dissection of the right femoral after impella explant. it is unclear if the dissection was pre-existing during initial intra-aortic balloon pump placement. hemostasis was achieved with perclose 2x, angioseal, balloon inflation, and manual pressure. the patient survived the event. | [5 empty] | The investigation for the reported vascular damage has been completed. no product was returned for evaluation. clinical details noted that patient had previous intra-aortic balloon pump (iabp) placed in same right groin access site as the impella pump. it was not noted if vessel dissection occurred when iabp or impella was inserted. additionally, it was noted that the patient had calcified arteries. the root cause of the vascular damage could not be definitively determined as it was not clear when the vascular damage occurred. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-25115 incorrectly. no ifus are needed to be reported for this report. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
12663 | 1220648-2024-25116 | 21018434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018434&pc=OZD | 12/11/2024 | 01/14/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507079 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a patient awaiting transplant of the heart and kidneys. the patient suffered a cerebrovascular accident on day 17. the team removed all systemic heparin. patient was treated by thrombectomy and remains on support with the 5.5 pump. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issue could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25116. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25116as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 codes 4614 and 4625 were reported incorrectly on the initial manufacturer device report number 1220648-2024-25116. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-25116 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12664 | 1220648-2024-25117 | 21018452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018452&pc=OZD | 12/11/2024 | 01/24/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590689 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that on day one of impella cp support some groin site bleeding was noted. manual pressure and angle of insertion was maintained. bleeding resolved. it was also reported that overnight the patient experienced acute right sided weakness and mental status deterioration. the patient underwent a computed tomography scan and received tenecteplase (tnk). heparin on hold at his time per hospital¿s post-tnk protocol. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: b2- selected death and added date of death. b5- patient outcome and mso review statement added. h1- changed to death. h6-impact code added 1802.
12665 | 1220648-2024-25118 | 21018463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018463&pc=OZD | 12/08/2024 | 01/27/2025 |  | 12/26/2024 |  |  | 12/08/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578299 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. after inserting a 6 french sheath, the right femoral artery was found to be occlusive. a rim catheter was used to go up and over to the left femoral artery; however, it was also vasoconstricted. angiography of subclavian artery was performed and vessel was less than 4mm. impella insertion was aborted due to the patient¿s vascular anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
12666 | 1220648-2024-25119 | 21018538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018538&pc=OZD | 12/12/2024 | 02/12/2025 |  | 12/26/2024 |  |  | 12/12/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580849 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that on day one of impella 5.5 support, a high purge pressure system blocked alarm occurred. tissue plasminogen activator was added to the purge. the next day impella was removed as lysis was unsuccessful in bringing down the high purge pressure and rising motor current. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. since neither product nor data logs were returned, the root cause of the high purge pressure could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25119.
12667 | 1220648-2024-25120 | 21018542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018542&pc=OZD | 12/12/2024 | 01/29/2025 |  | 12/26/2024 |  |  | 12/12/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, unable to doppler distal pulses in right or left foot. the patient was taken to the catheterization lab to shoot runoff of right leg. distal runoff showed no flow past impella sheath. the doctor decided to explant the impella. a thrombectomy to the right leg was performed, and flow restored down right leg. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12668 | 1220648-2024-25121 | 21018550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018550&pc=OZD | 06/15/2023 | 12/26/2024 |  | 12/26/2024 |  |  | 06/15/2023 | 12/26/2024 |  |  | 01/31/2025 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317679 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device and extracorporeal membrane oxygenation for mechanical circulatory support. the patient was then diagnosed with right heart failure and venoarterial extracorporeal membrane oxygenation was also cannulated. it was reported while on impella cp support, a suspected thrombotic event in the impella caused suction alarms, turbulent outflow area, and flows that displayed 0/0/0. the decision was made to explant the impella cp for no impella flow, however the decision was made to withdraw care and the pump was not replaced. the patient subsequently expired. there is not enough evidence to exclude impella as an associated factor in the patient's demise. | [5 empty] | Visual inspection found biomaterial in the pump housing. no other issue were found on the rest of the pump. the cause of the issue was biomaterial. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
12669 | 1220648-2024-25122 | 21018561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018561&pc=OZD | 12/12/2024 | 02/03/2025 |  | 12/26/2024 |  |  | 12/12/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576582 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a bipella support ongoing for a patient admitted in with cardiogenic shock and on extra corporeal membrane oxygenation. the bipella support was running when on day three of the support, the team observed thrombocytopenia. the heparin was removed from the purge in the rp and 5.5 pumps. the support continues to run with transplant conversations. | [5 empty] | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. d.6b explant date was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25122 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12670 | 1220648-2024-25123 | 21018585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018585&pc=OZD | 12/09/2024 | 01/29/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489604 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during insertion of a 14 fr long introducer, the sheath kinked due to a sharp bend in the patient's artery. sheath insertion was ultimately successful and the impella cp pump was implanted without complication for patient support. there was no patient harm from the kinked introducer. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ | The investigation of introducer damage ¿ mechanical has been completed. the pump was not returned for investigation. according to the clinical report, the patient¿s left femoral artery (lfa) was tortuous but suitable for insertion. after inserting the long impella sheath, a slight kink near a sharp bend in the artery was noted on fluoroscopy. the physician advanced the impella catheter past the kink with slight resistance and successfully placed it in the preferred location. the cause of the introducer damage issue was most likely patient condition related to diseased/small vessels condition, based on given clinical details. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-25123 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
12671 | 1220648-2024-25124 | 21018672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018672&pc=PYX | 12/09/2024 | 01/30/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578380 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. shortly after implant, the pump prolapsed in the right ventricle and could not be repositioned. the pump was subsequently replaced to resolve the issue. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on clinical details provided, the root cause of the positioning issue was determined to most likely be a use issue since the pump reportedly prolapsed back into right ventricle upon introducer removal. d3 manufacturer us zip code was reported incorrectly on manufacturer device report 1220648-2024-25124. e4 should have been left blank on manufacturer device report 1220648-2024-25124. g1 manufacturing site us zip code was reported incorrectly on manufacturer device report 1220648-2024-25124.
12672 | 1220648-2024-25126 | 21018690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018690&pc=OZD | 12/09/2024 | 01/29/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support ongoing for a patient admitted in and found to have multivessel disease and coronary artery disease which prompted coronary artery bypass graft consult. the team noted signs of hemolysis and this prompted the team to consult with nephrology. the pump was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12673 | 1220648-2024-25125 | 21018693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018693&pc=OZD | 12/09/2024 | 01/13/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541760 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient had hematoma at impella access site. manual pressure was held and patient received two units of packed red blood cells. physician thought the hematoma might have been from the patient moving and causing bleeding around the sheath. physician decided to wean and remove the impella. the femoral artery was closed with two percloses. the patient outcome was noted as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to patient movement since the hematoma might have been from the patient moving and causing bleeding around the sheath. e4 should have been left blank on manufacturer device report 1220648-2024-25125. h6 code 1888 was reported incorrectly. new code was added to health effect - clinical code on manufacturer device report 1220648-2024-25125.
12674 | 1220648-2024-25127 | 21018712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018712&pc=OZD | 12/09/2024 | 01/08/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567294 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. shortly after the pump was started, the waveforms and flows were inaccurate and non-functioning. the pump was exchanged for a new one. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. product and data logs were not returned for review. the cause of the low pump flow issue was not determined since product, logs were not returned and due to insufficient clinical information.
12675 | 1220648-2024-25128 | 21018764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018764&pc=OZD | 12/09/2024 | 01/16/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574172 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had ectopy. the pump showed deep and the physician repositioned the pump. support continued. additionally, during ambulation, the patient had a syncopal episode and fell forward. this caused the catheter to be severed from the red plug on the device. the impella motor stopped. the pump was explanted and the patient did not need a replacement at this time. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of ectopy and product damage (bond failure) has been completed. the returned pump was inspected and aligning with the clinical details. the catheter shaft was severed from the red plug on the device. there were no other physical abnormalities with the device. the clinical details mentioned patient movement at the time of damage. the root cause of the product damage (bond failure) was determined to be due to be a use-related issue. as the necessary information was not provided, the root cause of the ectopy was not determined.
12676 | 1220648-2024-25129 | 21018897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018897&pc=OZD | 12/07/2024 | 02/24/2025 |  | 12/26/2024 |  |  | 12/10/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211689 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller error alarm which self-resolved. the aic was exchanged. there was no harm to the patient. | [5 empty] | The investigation for the console (aic) controller failure-yellow alarm has been completed. data logs were reviewed which showed that the console displayed 'controller error (loss of comm (pbm1)) ¿ alarm' (event #1023), which was resolved automatically in 30 seconds (imc_logs_ic4008.pdf). this confirms the reported failure mode. support was then continued for an additional 26 minutes. the console was returned for evaluation. the controller error could not be reproduced by wiggling the data cable between the pbm and the 4-in-1. no signs of corrosion were found on either side of the pbm. all cables were appropriately connected found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). the root cause of the controller error was not determined due to the inability to reproduce the reported failure mode. added h6 impact code 4629. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12677 | 1220648-2024-25130 | 21018957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21018957&pc=OZD | 12/10/2024 | 01/24/2025 |  | 12/26/2024 |  |  | 12/10/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573502 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. four units of packed red blood cells were given to the patient and a femstop and suture were applied to the site. support continued. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, access site stabilized with fem stop and purse string suture as per the clinical description provided.
12678 | 1220648-2024-25131 | 21019620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21019620&pc=OZD | 12/04/2022 | 10/21/2025 |  | 12/26/2024 |  |  | 12/04/2022 | 12/26/2024 |  |  | 10/31/2023 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147872 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | Thrombocytopenia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a decrease in platelet count and heparin induced thrombocytopenia (hit) was suspected. on the same day (b)(6) 2022, it was noted the patient was experiencing symptoms of hemolysis was noted bloody urine. the patient was placed on dialysis on (b)(6) 2022 after noted red urine for several days. on (b)(6) 2022 it was noted that the patient had an onset of sepsis and heparin was removed from the purge solution as a result of the suspected hit. no further information was provided on the cause or treatment of the sepsis. the patient expired on support 8 days later. there was no cause of death reported. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records.the investigation of infection/sepsis, hemolysis, and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12679 | 1220648-2024-25132 | 21019668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21019668&pc=OZD | 12/10/2024 | 02/12/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553997 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan had a cp pump support placed for the support of a patient admitted in acute myocardial infarction / cardiogenic shock. the patient was additionally supported by extra corporeal membrane oxygenation. the cp was explanted after one day with signal issues. the pump was explanted and replaced by a 5.5 pump. the patient survived the event. | [5 empty] | The investigation of optical signal issue has been completed. the impella cp was returned for evaluation. upon visual inspection, no biomaterial was present. the pump passed the laser light continuity testing. the pump's optical parameters were checked and no abnormalities or hard failures of the optical sensor were observed. data logs were reviewed and long term logs showed that shortly after thepump was initiated, repeated "impella in aorta" alarms were observed. suction alarms occurred at the same time. optical parameters at time of issue showed no hard failures of the optical sensor. additionally, real time log at the beginning of the case showed motor current spike with a speed deviation. one minute later, a second motor current spike with a spike in purge flows was also seen. right after the second motor current spike, pump flows become saturated. clinical details noted that pump flows and waveforms were abnormal at the beginning of the case and "impella in aorta" alarms were occurring. the pump position was confirmed via echocardiogram. it was noted that alarms resolved when the pump was turned up to p-8 and the pump remained in the patient overnight before escalation to the impella 5.5. the root cause of the optical signal issue was most likely due to biomaterial as multiple motor current spikes were seen in the logs at the time of the "impella in aorta" alarms. e.1 revised last name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25132. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12680 | 1220648-2024-25133 | 21019780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21019780&pc=OZD | 01/04/2022 | 12/26/2024 |  | 12/26/2024 |  |  | 01/04/2022 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395260 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | No Audible Alarm | Obstruction of Flow | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported on (b)(6) 2022 while on impella 5.5 support, the automated impella controller (aic) did not alarm for inappropriately low purge flow readings after the impella 5.5 catheter became kinked. the catheter was unkinked to correct the purge flow issue. the aic was not replaced, and the medical staff was advised to monitor more closely and abiomed support team was also involved and asked to pay attention to impella connect for any further purge flow issues. the patient was noted to have developed sepsis, experienced atrial fibrillation, arrhythmias, and the pump required repositioning during impella 5.5 support. the patient¿s care was withdrawn, and the patient expired. it is unknown how long the purge flow was low, and how this may have affected the patient and therefore there is not enough evidence to exclude the aic as an associated factor in the patient's demise. | [5 empty] | This is one of two medical device reports associated with the event. refer to initial medwatch with date of report 26-dec-2024 for pump impella 5.5 with serial number (b)(6). data and device analysis were unable to be performed because the console was not returned for evaluation and logs were not provided for analysis. the cause of the low purge flow was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
12681 | 1220648-2024-25134 | 21020013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21020013&pc=OZD | 01/04/2022 | 04/16/2025 |  | 12/26/2024 |  |  | 01/04/2022 | 12/26/2024 |  |  | 08/31/2023 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139270 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | Obstruction of Flow | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Arrhythmia | Atrial Fibrillation | Sepsis | Tachycardia | [9 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported on 04-jan-2022 while on impella 5.5 support, the automated impella controller (aic) did not alarm for inappropriately low purge flow readings after the impella 5.5 catheter became kinked. the catheter was unkinked to correct the purge flow issue. the aic was not replaced, and the medical staff was advised to monitor more closely and abiomed support team was also involved and asked to pay attention to impella connect for any further purge flow issues. on 06-jan-2021 the patient developed sepsis during impella 5.5 support, and acutely declined. the patient was then re-intubated during a vacuum-assisted closure change. later that evening the patient had an episode of tachycardial atrial fibrillation. due to cardiac arrhythmia and septic shock, the patient¿s catecholamine was increased. the patient began improving on (b)(6) 2021, and it was noted on 09-jan-2021 that after repositioning the impella the impella was running smoothly again. there was no information provided as to why the impella had not been running smoothly before the need to reposition the pump. on (b)(6) 2021 the patient continued to experience episodes of atrial fibrillation and arrhythmias and required levosimendan and an increased dosage of norepinephrine. the patient was then withdrawn from care, and expired. there is not enough evidence to exclude impella 5.5 as an associated factor in the patient's demise. | [5 empty] | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-25133 for the automated impella controller with serial number (b)(6). the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issues have been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
12682 | 1220648-2024-25135 | 21021580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021580&pc=OZD | 12/05/2024 | 09/30/2025 |  | 12/26/2024 |  |  | 12/05/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501315 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Renal Failure | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured thrombocytopenia with a possible relationship to the impella device used: on (b)(6) 2024, patient has thrombocytopenia at 64. (b)(6) - stable platelet count. the patient recovered with no sequelae. | Additional information provided from abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported worsening acute kidney injury with relationship listed as possibly related. patient was started on continuous renal replacement therapy and was hemodynamically stable off pressors. transitioned to hemodialysis (hd) on (b)(6). line holiday for mrsa (b)(6), blood cultures were clear since (b)(6) . last hd on (b)(6) with renal recovery. patient was started on spironolactone, had a decrease in potassium, and continued on torsemide. additional right lower extremity (rle) compartment syndrome was reported with relationship of possibly related. loqi states: post va ecmo cannulation, patient had right lower extremity compartment syndrome requiring emergent bedside fasciotomies x4 with necrotic muscle noted at time of fasciotomy. (b)(6) underwent bedside cauterization of rle fasciotomy site for persistent serosanguinous drainage. (b)(6) rle debridement and partial complex would closure. the deep posterior compartment had muscle that appeared nonviable and nonreactive to cautery. remained critically ill with non- healing, infected fasciotomy wounds with loss of pulsatility. discussed risks and benefits with family and they elected to proceed with right above the knee amputation. on (b)(6).. | Additional information provided from abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported worsening acute kidney injury with relationship listed as possibly related. patient was started on continuous renal replacement therapy and was hemodynamically stable off pressors. transitioned to hemodialysis (hd) on (b)(6). line holiday for mrsa (b)(6), blood cultures were clear since 8/15. last hd on (b)(6) with renal recovery. patient was started on spironolactone, had a decrease in potassium, and continued on torsemide. additional right lower extremity (rle) compartment syndrome was reported with relationship of possibly related. loqi states: post va ecmo cannulation, patient had right lower extremity compartment syndrome requiring emergent bedside fasciotomies x4 with necrotic muscle noted at time of fasciotomy. (b)(6) underwent bedside cauterization of rle fasciotomy site for persistent serosanguinous drainage. (b)(6) rle debridement and partial complex would closure. the deep posterior compartment had muscle that appeared nonviable and nonreactive to cautery. remained critically ill with non- healing, infected fasciotomy wounds with loss of pulsatility. discussed risks and benefits with family and they elected to proceed with right above the knee amputation. on (b)(6). | The investigation uncovered placement signal issues. |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. the root cause cannot be determined as no product was returned. | The investigation for the thrombocytopenia, limb ischemia, renal failure, and sepsis has been completed. no product was returned for evaluation. clinical details states that the pump remained supporting patient for an additional 6 days before explant. without the pump nor sufficient clinical details, the root cause of the thrombocytopenia, renal failure, limb ischemia, and sepsis could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. b5-added additional narrative h6-clinical code 2067, 1942, 2041, impact code 4644. | No product was returned for evaluation. data logs were reviewed and some variability in placement signal and pump flow as well as snr. rt logs from time of alarm show no hard failures of optical parameters. clinical details noted that ecmo was decannulated after alarms began. pump remained supporting patient for an additional 6 days before explant. the root cause of the optical signal issue cannot be determined as no product was returned to confirm failure of optical sensor. thrombocytopenia: the cause of the issue was not determined due to insufficient clinical details. renal failure: the cause of the issue was not determined due to insufficient clinical details. limb ischemia - irreversible: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. infection/sepsis: an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. | B1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-25135. d10 concomitant medical products and therapy dates updated. e4 should have been left blank in the initial report.
12683 | 1220648-2024-25136 | 21021586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021586&pc=OZD | 06/27/2024 | 01/16/2025 |  | 12/26/2024 |  |  | 12/05/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501313 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Death | Required Intervention | [23 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured left thrombus. the patient also endured ventricular tachycardia (vt) and ventricular fibrillation (vf) arrest. the patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported ventricular fibrillation (vf) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf and thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: d4 updated model number.
12684 | 1220648-2024-25137 | 21021588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021588&pc=OZD | 08/13/2024 | 02/11/2025 |  | 12/26/2024 |  |  | 12/05/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510633 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening thrombocytopenia with a possible relationship to the impella device used: worsening thrombocytopenia at 86k. received two units of packed red blood cells, one unit of platelets, and one unit of fresh frozen plasma. platelet goal of >100. 8/21 - platelets stable. the patient recovered with no sequelae. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25137. g2 report source; study was missing from manufacturer device report 1220648-2024-25137.
12685 | 1220648-2024-25138 | 21021599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021599&pc=PYX | 08/09/2024 | 02/11/2025 |  | 12/26/2024 |  |  | 12/05/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478818 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening thrombocytopenia and acute chronic anemia with a possible relationship to the impella device used: patient developed mild thrombocytopenia post-operation on (b)(6) 2024. thrombocytopenia worsened after rp impella implant (b)(6) 2024. baseline platelet post-procedure: 108 k/ul on (b)(6) 2024 at 12:42. nadir platelet post-impella implant: 82 k/ul on (b)(6) 2024 at 04:00. resolution platelets: 129 k/ul on 17 aug 24. action taken: transfuse rbc prn. the patient recovered with no sequelae. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined.
12686 | 1220648-2024-25139 | 21021600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021600&pc=OZD | 12/09/2024 | 01/30/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561724 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the implant, during the removal of the peel away sheath, the physician cracked the sheath while the distal end was still in the patient's vessel. arterial bleed back led staff to hold pressure at site and sheath was peeled away and removed. the repositioning sheath was inserted in suboptimal conditions and the impella in ventricle alarm appeared. pump visualized with fluoroscopy and was noted to be folded over itself twice in the ventricle. repositioning was attempted however, it was not successful. the physician removed and replaced the pump. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Impella cp was returned for evaluation. no abnormalities were observed on the repositioning sheath. a kink in the cannula was observed just distal of the outflow cage. no peel-away introducer was returned for evaluation. no data logs were returned for evaluation. the root cause of the introducer damage was most likely use related. the root cause of the positioning issues was most likely a use related issue as it occurred when the repositioning sheath was being inserted. d9 date device returned to manufacturer was omitted from manufacturer device report 1220648-2024-25139 e4 should have been left blank on manufacturer device report 1220648-2024-25139. h3 device not returned to manufacturer for evaluation should have been marked no on manufacturer device report 1220648-2024-25139. h6 code 1888 and 4627 were reported incorrectly on manufacturer device report 1220648-2024-25139. new codes were added to health effect - clinical code and health effect - impact code. h6 type of investigation 4114 device not returned was submitted in error on manufacturer device report 1220648-2024-2513 as the device was returned.
12687 | 1220648-2024-25140 | 21021602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021602&pc=OZD | 12/09/2024 | 02/12/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204376 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) system self-check failed. the aic was exchanged. there was no harm to the patient. | [5 empty] | Additional information provided in h6 and h10. the investigation for the console (aic) self-check error has been completed. data logs were reviewed during the initial boot-up, the console restarted automatically due to the communication issue. after restarting, due to the persistent communication issue, the console displayed the ¿system self check failed ¿screen. this confirms the reported failure mode. this console was returned for evaluation and reproduced the reported failure mode ¿ corruption was confirmed after comparing the original firmware with the original production firmware. the root cause for the ¿system self check failed¿ screen was firmware corruption, but the cause of the corruption was not determined. corrected information provided in d2. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12688 | 1220648-2024-25141 | 21021604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021604&pc=OZD | 12/09/2024 | 01/28/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding. the patient was administered two units of packed red blood cells. the patient continued support and was noted to be stable. the patient was bridged to an additional device for transplant work up. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25141.
12689 | 1220648-2024-25142 | 21021605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021605&pc=OZD | 12/09/2024 | 01/28/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  | 06/30/2025 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364363 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that roughly 14 days into impella 5.5 support, the patient began to succumb to their condition and coded. the surgeon pulled the pump back to check the valve after attempting to treat the condition and the pump subsequently experienced a flow failure coinciding with a high motor current. after multiple attempts to discuss potential options, the decision was made to withdraw the pump and the patient expired. it is unknown if the complications with the pump contributed to the patient's passing. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. pump was not returned for investigation. data logs were not returned as well. without the pump and data logs the failure cannot be investigated further. therefore, the root cause of the pump stop issue was not determined since product and data logs were not returned for investigation.
12690 | 1220648-2024-25143 | 21021662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021662&pc=OZD | 12/09/2024 | 01/30/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565560 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the implantation was aborted due to complications encountered with the patient's anatomy. there was no patient harm resulting from the delivery failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. no product was returned. the root cause of the delivery issue was most likely due to patient condition since the implant of pump was aborted due to patient's anatomy. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25143. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25143 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12691 | 1220648-2024-25144 | 21021676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021676&pc=OZD | 12/09/2024 | 01/08/2025 |  | 12/26/2024 |  |  | 12/09/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured pericardial effusion/tamponade with a possible relationship to the impella device used: patient underwent mitral valve repair / tricuspid valve repair requiring venous arterial extra corporeal membrane oxygenation, was later found to have large pericardial effusion w/ concern for tamponade, returned to operating room for exploration and evacuation of hematoma, final echocardiography on (b)(6) revealed moderate volume complex pericardial effusion. it was reported that the patient/event has not recovered/not resolved. the patient care was withdrawn and patient expired. | [5 empty] | The investigation for the vascular damage has been completed. no product was returned for analysis. clinical details do not mention excessive force or patient anatomy issues. the root cause of the vascular damage - great vessel was not determined as limited clinical information was provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12692 | 1220648-2024-25145 | 21021681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021681&pc=OZD | 12/10/2024 | 01/24/2025 |  | 12/26/2024 |  |  | 12/10/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539666 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. following insertion, it was noted that the device was unable advance beyond the subclavian artery. the implant was aborted and an impella cp device was subsequently placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the implant of pump was aborted due to patient's anatomy.
12693 | 1220648-2024-25146 | 21021688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021688&pc=OZD | 11/17/2022 | 01/29/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  | 07/31/2024 |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023256411 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 27 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded 7 publications. one from team in ct/(b)(6) hospital which was entitled: "one pump for two hearts: using an impella 5.5 micro-axial pump in peripartum cardiogenic shock" which details a cp support in an expectant mother who had an access site ooze. until the day before delivery, she was anticoagulated with heparin, along with a heparin-based purge fluid. a day before delivery, she developed epistaxis and oozing from the impella placement site. anticoagulation was temporarily stopped, and the heparin-based purge fluid was substituted with 5% dextrose but oozing continued. ultimately a decision to proceed with delivery was made due to persistent oozing despite holding heparin and severe fetal growth restriction. the procedure was successful with another device and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. article citation: pillai, a. a., badr, m., mai, x., dilip, a., & bell, j. (2024). one pump for two hearts: using an impella 5.5 micro-axial pump in peripartum cardiogenic shock. asaio journal, 70(11), e159¿e161. https://doi.org/10.1097/mat.0000000000002215 e4 should have been left blank on manufacturer device report 1220648-2024-25146. g2 report source; literature was inadvertently not selected on manufacturer device report 1220648-2024-25146. h10 article citation was omitted from manufacturer device report 1220648-2024-25146. article was inadvertently not attached on manufacturer device report 1220648-2024-25146.
12694 | 1220648-2024-25148 | 21021695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021695&pc=OZD | 12/11/2024 | 01/31/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553018 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp support initiated for a patient admitted with acute myocardial infarction and cardiogenic shock and st elevation myocardial infarction. the pump was explanted when repositioning failed to remedy the pump positioning and there was signs of hemolysis. the pump could not be positioned away from the mitral. the pump was removed and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed. no suctions occurred. positioning alarms occurred in the ventricle and aorta. the root cause of the positioning issue was not determined due to lack of clinical details. the root cause of the hemolysis issue most likely was improper positioning. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25148 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12695 | 1220648-2024-25147 | 21021698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21021698&pc=OZD | 12/11/2024 | 01/13/2025 |  | 12/26/2024 |  |  | 12/11/2024 | 12/26/2024 |  |  |  |  | 12/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573496 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. upon admission to cardiovascular intensive care unit, the patient had weak pulses but were detected by doppler. on the next hour check, the nurse was unable to doppler pulses. very little and slow flow noted to right lower extremity (rle). decision was made to remove impella and pulses returned per doppler to rle. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25147.
12696 | 1220648-2024-25083 | 21011054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011054&pc=OZD | 12/07/2024 | 01/08/2025 |  | 12/24/2024 |  |  | 12/07/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573395 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was primed for insertion. however, physicians made multiple attempts over the course of 1.5 hour to achieve ventricular access but had extreme difficulty passing the aortic valve. several different catheters and wires were used. calcium & epi pushes were used to help the valve to open, as well as breath holds. doctors continued to have difficulty advancing past valve, and expressed concerns for not seeing wire on trans-esophageal echo (tee). additional imaging showed dissection of ascending aorta/ innominate with compromised blood flow to the right upper extremity (rue). case was aborted. impella was primed but not placed. vascular surgery consulted to repair flow to rue and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12697 | 1220648-2024-25084 | 21011137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011137&pc=OZD | 12/07/2024 | 02/11/2025 |  | 12/24/2024 |  |  | 12/07/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558652 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Fluid/Blood Leak | Priming Problem |  |  |  |  | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that in the catheterization lab the patient was under extra corporeal membrane oxygenation (ecmo) and impella cp was being placed. echocardiography signaled air bubbles. the physician opened the purge luer-lock (yellow connector) to further de-air the purge line without using the specific menu for it. it looked like the impella was not completely primed before implant. additionally, after the impella was recently placed, the cp was having suction alarms at p2. motor current waveforms pulsatile but negative placement signal pressure. physician said that the position was optimal, but patient needed a lot of filling. the next morning the cp was running at p2 for 1.2l/min. placement signal recovered, no alarms. it was also reported that the patient developed limb ischemia due to ecmo and impella. fem-fem bypass applied. the patient was stable. the impella was explanted due to the ischemia. the patient survived the event. | [5 empty] | The device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of limb ischemia, positioning issues, and purge leak ¿ pump has been completed. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the failure mode (fm) "positioning issues" was updated to "optical signal issue" since the reported issue was regarding the placement signal and there were no positioning issues. there were no abnormalities found with the returned pump regarding the optical system. log review showed suction and impella position unknown alarms while the placement signal was negative in the beginning of the case. the cause of the optical signal issue was most likely patient condition related since the position was optimal, but the patient had low volume and needed a lot of filling. the fm "purge leak - pump" was updated to "priming issue" since there was no leak reported and the issue occurred during priming. the root cause of the priming issue was most likely use-related as the returned pump was able to prime completely and the clinical details suggest a use issue. b1 was inadvertently omitted from manufacturer device report 1220648-2024-25084. h6 health effect - impact code 4625 was inadvertently coded twice on manufacturer device report 1220648-2024-25084. new code was added to health effect - impact code.
12698 | 1220648-2024-25085 | 21011248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011248&pc=OZD | 12/08/2024 | 02/20/2025 |  | 12/24/2024 |  |  | 12/08/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510640 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A complainant reported the patient was on impella 5.5 support. while on support, ¿pp high¿ problem occurred. there were no confirmed kinks. lysis protocol was sent. the patient remains on support. | The decision was made to withdraw care and the patient expired. the patient expired after having 169.77 hours of impella support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and long term logs from the field showed a rise in purge pressure over a period of approximately four hours with a decrease in purge flow. purge pressure remained high for approximately 15 hours before the purge pressure decreased with a slight increase in purge flow. purge pressure values remained slightly elevated from pre-alarm values for remainder of case. clinical details noted that the current activated clotting time value was not available at time of purge alarms. additionally, tissue plasminogen activator (tpa) protocol was provided to the clinical team but it was not noted if tpa resolved the issue. the root cause of the high purge pressure was most likely due to biomaterial. a.1 revised as identifer was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25085. d.6a and d.6b revised as dates were reviously entered incorrectly on the initial manufacturer device report number 1220648-2024-25085. e.1 revised first and last name as they were reviously entered incorrectly on the initial manufacturer device report number 1220648-2024-25085. g.1 revised reporting contact email was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25085. | B.1 revised to reflect adverse event and product malfunction on manufacturer device report number 1220648-2024-25085. b.2 revised to reflect patient death and date of death on manufacturer device report number 1220648-2024-25085. b.5 revised to include the patient outcome of death on manufacturer device report number 1220648-2024-25085. h.1 revised to reflect death on manufacturer device report number 1220648-2024-25085. h.6 health effect - impact code 1802 was added to reflect patient outcome on manufacturer device report number 1220648-2024-25085.
12699 | 1220648-2024-25086 | 21011346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011346&pc=OZD | 12/08/2024 | 01/22/2025 |  | 12/24/2024 |  |  | 12/08/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590570 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an impella cp implant and during the implant, the guidewire kinked. the wire was removed and boston science v-18 control wire was used successfully. the pump was implanted successfully. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the product was not returned for investigation. data log review was not required for this case run. according to the clinical details, the abiomed guidewire was kinked during the removal of the dilator. no image was provided for the guidewire damage. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided.
12700 | 1220648-2024-25087 | 21011356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011356&pc=OZD | 03/20/2023 | 12/24/2024 |  | 12/24/2024 |  |  | 03/20/2023 | 12/24/2024 |  |  | 10/31/2024 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024291586 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient who arrived to the emergency room with an abnormal ekg and chest pains was implanted with an impella cp for mechanical circulatory support and transferred the patient for a coronary artery bypass graft (cabg) consult. it was reported while on impella cp support, the patient experienced a small retroperitoneal bleed. the medical team planned on explanting the impella cp the following morning and heparin was discontinued for the planned event. however, during the weaning of the impella for explant, the patient became hypotensive and the medical team then decided to abort the explant. the patient was then successfully escalated to an impella 5.5 the following morning for increased support for a planned cabg. the patient's outcome at explant and procedure was noted to be survived. | [5 empty] | The root cause of the patient injury, retroperitoneal bleed, was most likely patient condition related. the relationship between the bleed and impella was unable to be determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
12701 | 1220648-2024-25088 | 21011451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011451&pc=OZD | 04/08/2024 | 02/13/2025 |  | 12/24/2024 |  |  | 12/03/2024 | 12/24/2024 |  |  | 11/30/2025 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439378 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolyzing anemia with a possible relationship to the impella device used: suspected hemolysis and anemia due to impella per peripheral smear and hematologist/oncologist doctor. it was reported that the patient/event had not recovered/not resolved. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. no product returned. it is unclear at what time during support the hemolysis started. data logs showed no abnormalities. no motor current spikes were noted. the cause of the hemolysis is not determined due to lack of clinical details and no product returned. b.2 revised selection since the initial manufacturer device report number 1220648-2024-25088 was submitted as no known intervention was noted. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25088 and was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25088 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 revised as study selection was missing from initial manufacturer device report number 1220648-2024-25088. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25088 and a new code was added.
12702 | 1220648-2024-25089 | 21011457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011457&pc=OZD | 12/06/2024 | 01/23/2025 |  | 12/24/2024 |  |  | 12/06/2024 | 12/24/2024 |  |  | 09/30/2025 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402976 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Perforation | [11 empty] | Required Intervention | Death | Life Threatening | Disability | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) but would expire during the procedure from a perforated artery near the arteriotomy. diagnostic catheterization one day prior to the impella implant procedure revealed severe aortic valve stenosis. the patient returned to intensive care unit for planned transfer to an outside hospital for transcatheter aortic valve replacement (tavr). however, the patient decompensated overnight and was brought to catheterization lab for an urgent valvuloplasty prior to transfer. during the valvuloplasty, patient further decompensated, and decision was made to urgently place an impella cp device for mcs. during the impella implant, the physician encountered difficulty while placing the 25cm peel-away. it was noted, however, there were no problems during the procedure and patient was critical but stable. upon removal of the peel-away sheath the patient began to have reduced blood pressure. fluoroscopy of the left groin revealed no issues. fluoroscopy of the right groin/insertion site revealed a perforation of the iliac proximal to the arteriotomy. manual pressure immediately applied and "up and over" balloon tamponade attempted. the patient continued to decompensate, and cardiopulmonary resuscitation was started and lasted for 15 minutes before time of death was called. the physician stated post procedural area - "it's possible there was a perforation during very difficult placement of the peel-away in tortuous artery that was tamponade by the 14fr sheath. upon removal of the peel-away it revealed the perforation." | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation for the vascular damage has been completed. no product was returned and logs are not applicable. the patient's artery was noted to be tortuous and the md suspected that could have been the cause of the perforation. the root cause of the vascular damage was most likely patient condition related.
12703 | 1220648-2024-25090 | 21011458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011458&pc=OZD | 12/06/2024 | 02/10/2025 |  | 12/24/2024 |  |  | 12/06/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721787 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported two automated impella controllers (aic) were shipped from one facility to another after a patient transfer. upon arrival to the receiving facility, the cardiac catheterization lab director noted that both units had been placed in one box and in spite of a certain amount of cushioning, both aics were broken (visibly cracked in several places around outer aic housing). field service was contacted for return of both devices for repair. there was no harm to the patient as a result of this event. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for automated impella controller serial number (b)(6) with date of event 06-dec-2024 and identifier ending in (b)(6). | The investigation has been completed. data analysis proved uninformative for this console. the console was returned for this investigation and there was significant damage to the console in a number of locations. the front housing had cracks on both the left and right side of the console. additionally, the purge was cracked and missing chunks on the outside and inside. additionally, the rear housing was cracked such that the rear housing could not flushly connect to the front housing. the console was tested and performed as expected except the batttest showed depleted batteries that would not charge. replacing the batteries with known-good batteries and the console was able to charge and power the console as expected. there was no evidence in the automated impella controller or long-term power logs why the batteries were originally depleted. the cause of the console damage to the front and rear housings and purge assembly could not be determined through log and console analysis but was most likely due to a user issue. the cause of the failure to power the console and charge was depleted batteries. it is undetermined why the batteries became depleted. h.10 the investigation results in follow up manufacturer device report 1220648-2024-25090 was inadvertently reported for a different serial number that was returned for investigation at the same time. the correct investigation results were added. | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis proved uninformative for this console. the console was returned for this investigation and a small crack was noticed on the rear housing. the console was tested but performed as expected. the cause of the crack in the rear housing could not be determined through log and console analysis but was most likely due to a user issue. d.9 device returned date/information was added. a.1 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25090. e.1 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25090. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25090 as it is unknown if the initial reporter also sent the report to the food and drug administration.
12704 | 1220648-2024-25091 | 21011462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011462&pc=OZD | 12/06/2024 | 02/10/2025 |  | 12/24/2024 |  |  | 12/06/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401285 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Crack | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported two automated impella controllers (aic) were shipped from one facility to another after a patient transfer. upon arrival to the receiving facility, the cardiac catheterization lab director noted that both units had been placed in one box and in spite of a certain amount of cushioning, both aics were broken (visibly cracked in several places around outer aic housing). field service was contacted for return of both devices for repair. there was no harm to the patient as a result of this event. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for automated impella controller serial number (b)(6) with date of event 06-dec-2024 and identifier ending in (b)(6). | The investigation has been completed. data analysis proved uninformative for this console. there was no evidence in the automated impella controller or long-term power logs why the batteries were originally depleted. the cause of the console damage to the front and rear housings and purge assembly could not be determined through log and console analysis but was most likely due to a user issue. the cause of the failure to power the console and charge was depleted batteries. it is undetermined why the batteries became depleted. a.1 revised as it was previously entered incorrected on the initial manufacturer device report number: 1220648-2024-25091. e.1 revised as it was previously entered incorrected on the initial manufacturer device report number: 1220648-2024-25091. e.4 should have been left blank on the initial manufacturer device report number: 1220648-2024-25091 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The investigation of automated impella controller (aic) - damage before therapy has been completed. data analysis proved uninformative for this console. the console was returned for this investigation and a small crack was noticed on the rear housing. the console was tested but performed as expected. the cause of the crack in the rear housing could not be determined through log and console analysis but was most likely due to a user issue. d.9 revised as device was returned and inadvertently omitted from follow up manufacturer device report number 1220648-2024-25091. h.3 revised as the device was returned for investigation and summary is attached were previously reported incorrectly on follow up manufacturer device report number 1220648-2024-25091. h.6 codes 4203 and 4307 were reported incorrectly on follow up manufacturer device report number 1220648-2024-25091 due the incorrect investigation results were reported on that follow up report. h.10 the investigation results in manufacturer device report 1220648-2024-25091 were inadvertently reported for a different serial number that was returned for investigation at the same time.
12705 | 1220648-2024-25093 | 21011466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011466&pc=OZD | 12/07/2024 | 02/11/2025 |  | 12/24/2024 |  |  | 12/07/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025447619 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient experienced heavy bleeding at all access sites. red blood cells and fresh frozen plasma were administered to the patient. furthermore, the pump had noted suction alarms. volume infusion resolved the low pump issue. support continued. additional information noted that the patient expired after 59.13 hours of impella support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleed and the low pump flow has been completed. many suction alarms occurred. motor current (mc) spike observed around 1.5 hours of impella support followed by sudden drop in flow. clinical states that patient had heavy bleeding around the same time as low flows and that volume was given which resolved the suction issue, however there is a clear mc spike with immediate drop in flows indicating biomaterial ingestion. the root cause of low pump flow issue most likely was biomaterial. the root cause of the access site bleeding most likely was patient condition related because the patient had heavy bleeding on all access sites (impella and ecmo). the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12706 | 1220648-2024-25092 | 21011467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011467&pc=OZD | 12/06/2024 | 01/23/2025 |  | 12/24/2024 |  |  | 12/06/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576320 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed hemolysis, positioning issues and thrombus. post implant of the impella cp, same day, the patient had progressively pink to dark urine and dropping off. at bedside the pump was pulled back under point of care ultrasound, locked at 94cm with formal echo pending. it was further reported after multiple attempts to position the impella cp appropriately under fluoroscopy and transthoracic echocardiogram, the decision made to exchange impella cp due to seesawing of catheter in place and severe hemolysis. thrombus was also reported to have occurred. the patient was eventually escalated to an impella 5.5 device and remained on support for at least an additional 16 days. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis, positioning issues, and thrombus has been completed. after review device and data analysis and clinical details, the root cause of the hemolysis and positioning issues was most likely was improper position due to improper technique. the logs indicate that thrombus could have been ingested towards the end of the case, however this is unlikely to have cause the hemolysis as the hemolysis was reported to have occurred before the ingestion event. as the necessary information was not provided, the root cause of the thrombus was not determined. e4 should have been left blank on manufacturer device report 1220648-2024-25092.
12707 | 1220648-2024-25094 | 21011475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011475&pc=OZD | 12/07/2024 | 01/28/2025 |  | 12/24/2024 |  |  | 12/07/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia and access site bleeding. impella was placed reduced to p-4 with good flows and no suction. no chugging noted on extracorporeal membrane oxygenation (ecmo) during unloading. angiogram of leg showed some flow, but reperfusion catheter placed to assist in limb flow from emco circuit. of further note, post implant same day, the patient had been anticoagulated for ecmo already, given extra heparin for impella placement. act 323 at impella insertion and rising. post procedure oozing was noted around impella insertion site. manual pressure applied and femostop placed. no blood products given. next day, recommendations for groin access site management, a stitch and some epi/lido were applied to the groin and had completely stopped the bleeding. the patient was noted to be neurologically intact. no packed red blood cells were given but one unit of fresh frozen plasma was administered overnight. the patient was planned for escalation to an impella 5.5 which occurred four days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report.
12708 | 1220648-2024-25096 | 21011479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011479&pc=OZD | 12/08/2024 | 01/13/2025 |  | 12/24/2024 |  |  | 12/08/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experience bleeding at the access site. impella inserted post percutaneous coronary intervention independently and was noted hemostasis could not be obtained at impella insertion site after peel away was removed and reposition sheath was advanced into the right femoral artery (rfa). ultimately decision was made by the physician to temporarily remove the impella over a wire, place a 16fr gore dryseal sheath, and re-implant impella. after that was completed, hemostasis was obtained at groin site and distal pulses in right lower extremities were confirmed to be present. the patient was transported back to the unit with a plan to consult vascular surgery for groin management and explant of the impella cp. two days later as per the plan, the patient was taken to operating room for surgical explant of impella cp from rfa. impella was pulled back across aortic valve and turned to p-0 then removed from 16f sheath. cutdown was performed and rfa isolated. 16f sheath removed and thrombectomy performed per protocol. the surgical closure was successful to rfa; distal pulses present. the patient was extubated and transferred back to the unit in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25096.
12709 | 1220648-2024-25095 | 21011481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011481&pc=OZD | 12/08/2024 | 02/11/2025 |  | 12/24/2024 |  |  | 12/08/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558640 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient (unknown race and ethnicity) with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced hemolysis and limb ischemia. two days after start of support, hemolysis was diagnosed by lactate dehydrogenase (ldh) increase. pump position was noted as "good." additionally, marbling of the impella leg was noted. consequently, the decision was made to explant the impella cp device with no option for fem-fem bypass surgery. the patient survived explant but was noted to be in critical condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and limb ischemia could not be determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.3 email should of been left blank on the initial manufacturer device report number 1220648-2024-25095. d.4 revised unique identifier (udi) # as it previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25095. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25095 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-25095 and were added.
12710 | 1220648-2024-25097 | 21011483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21011483&pc=OZD | 12/09/2024 | 01/27/2025 |  | 12/24/2024 |  |  | 12/09/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441818 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a patient admitted in with acute myocardial infarction / cardiogenic shock. the team first placed a cp but then escalated to the 5.5. the access site of the 5.5 had a bleed that was treated. the team used surgicell, packing, and transfused blood. the pump remained on for support despite the bleeding. the patient eventually expired after care was withdrawn. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25097.
12711 | 1220648-2024-25098 | 21012211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21012211&pc=OZD | 02/17/2023 | 10/09/2025 |  | 12/24/2024 |  |  | 02/17/2023 | 12/24/2024 |  |  | 08/31/2024 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1673122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Difficult to Remove | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support after decompensating status post left anterior descending balloon and stent placement. the patient plan of care was to transfer to another hospital for escalation of care and to escalate patient from the impella cp to an impella 5.5. it was reported on (b)(6) 2023 that during patient transfer to another hospital, the automated impella controller (aic) was swapped. it was reported immediately after swapping the aic, there was a purge system failure alarm. to resolve the issue, another aic was exchanged. support was delivered to the patient without interruption and there was no patient harm reported for the reported aic issue. additionally, on (b)(6) 2023, the patient experienced signs of hemolysis including red colored urine and a haptoglobin of 55 (no units reported) and ldh of 2600 (no units reported). the patient's plan of care was to escalate to the impella 5.5, and the impella cp p level was lowered to address the hemolysis. later that day on (b)(6) 2023 the patient was taken to the operating room to explant the impella cp and escalate the patient to the impella 5.5. the impella 5.5 was first placed in the left ventricle with the impella cp still placed. it was noted that the physician was unable to explant the impella cp due to a clotted off rewire access port. the decision was then made to perform a surgical cutdown and vascular closure using a bovine patch. it was reported that the patient was experiencing intermittent bradycardia, so a transvenous pacer was placed. the impella cp was successfully explanted, and impella 5.5 support was initiated. on (b)(6) 2023 it was noted that the patient was confirmed covid positive and the patient was treated for a mixed picture of septic and cardiogenic shock. on (b)(6) 2023 systemic heparin was discontinued as the patient was bleeding at iv lines. the patient was noted to have to arrhythmias on (b)(6) 2023. the patient was noted as not being a candidate for advanced therapies. on (b)(6) 2023 the patient care was withdrawn and placed on comfort measures only. the patient expired while on impella 5.5 support. the impella cp and impella 5.5 relationship to the patient's demise may be unlikely but cannot be excluded. the patient presented in a decompensating state and the patient was noted to be covid positive. the impellas cannot be excluded as a possible contributing factor to the overall even given the impella related events hemolysis, bleeding, bradycardia, and sepsis as these may have added to an already complex situation even though the patient was not a candidate for advanced therapies. | Us complainant reported that a patient was on impella cp for hemodynamic support. during support it was reported that they decreased the cp to p6 and increased epi support due to suspected hemolysis. the plan was to start systemic bivalirudin. decision was made by heart team to escalate support to impella 5.5, the patient was taken to or for 5.5 and cp explant, however, they were unable to explant cp due to clotted off rewire access port, decision for surgical cutdown and vascular closure using bovine patch. the patient was still having intermittent bradycardia so transvenous pacer placed, pocket closed and patient transferred to icu for further management. patient expired. |  |  |  |  | This is one of three medical device reports associated with the event. refer to initial medwatch for pump number two the impella 5.5 serial number (b)(6) and initial medwatch for the automated impella controller with mfr 1220648-2024-24895. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. updated date of death.
12712 | 1220648-2024-25099 | 21012233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21012233&pc=OZD | 07/07/2023 | 10/07/2025 |  | 12/24/2024 |  |  | 07/07/2023 | 12/24/2024 |  |  | 01/31/2025 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311972 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was cannulated with extracorporeal membrane oxygenation and placed on continuous renal replacement therapy. the patient was admitted with severe mitral regurgitation and a low ejection fraction of 15%. it was reported that just after impella was placed and during impella cp support, the patient experienced bleeding and a hematoma in right groin after the patient was anxious and began moving. the nurse held pressure, compressed the hematoma, and heparin was withheld. the patient's leg was placed in a brace to prevent further movement. on (b)(6) 2023 it was noted that heparin was withheld as the patient¿s liver function laboratory results were elevated, and the patient¿s platelet count was low with a result of 40,000 platelets per mcl. on (b)(6) 2023 the patient¿s overall health continued to decline and was in multi system organ failure. the patient appeared jaundiced. on (b)(6) 2023, the patient was stable, however the patient¿s partial thrombin time was supra therapeutic with no coagulation medications given. two days later, on (b)(6) 2023, the patient's care was withdrawn per family request after an acute decline in condition overnight and the patient expired. although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. | Us complainant reported a patient was on impella cp for hemodynamic support. during support it was noted of bleeding and a hematoma on right groin. the nurse held pressure and was compressing hematoma, heparin was held. additionally, it was noted that there was no anticoagulation due to elevated lft's and platelets. care was ultimately withdrawn and the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding - major is most likely due to the patients movement that was reported at the time of the bleed. the optical signal issue was investigated and deemed to be an automated impella controller (aic) related issue which is noted in (b)(4). this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12713 | 1220648-2024-25100 | 21012276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21012276&pc=PYX | 11/25/2024 | 01/08/2025 |  | 12/24/2024 |  |  | 11/25/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529163 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support and during support, the patient had atrial fibrillation (af). the pump position was confirmed and the staff was planning on explanting the following day. additionally, the patient had a hematoma at the access site. intervention for the hematoma is unknown. the patient expired after 61.22 hours of support after care was withdrawn. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25100.
12714 | 1220648-2024-25102 | 21014433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21014433&pc=OZD | 12/08/2024 | 02/12/2025 |  | 12/24/2024 |  |  | 12/08/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502457 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced hemolysis. placement of the pump was confirmed, and one unit of red blood cells was administered to the patient. furthermore, the patient experienced ventricular tachycardia (vt). the patient was shocked to convert out of vt. additional information noted that the patient expired after 42.78 hours of impella support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the reported issues has been completed. device history record review confirmed that the pump passed all post sterile inspection checks. data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved quickly. the device was not received for evaluation. the root cause of the reported issues was unable to be determined as limited clinical details were provided and no device was received for evaluation.
12715 | 1220648-2024-25103 | 21014434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21014434&pc=OZD | 12/12/2024 | 02/03/2025 |  | 12/24/2024 |  |  | 12/12/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571161 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a patient deteriorate after the deployment of the transcatheter aortic valve replacement. the team placed the cp but found the need to explant it due to bleeding at the hub. the team observed bleeding at the blue hub and first attempted to remedy with pressure and surgical glue. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The pump was returned for inspection and no abnormalities were noted upon investigation. the root cause of access site bleeding is determined to be use issue because of the bleeding from the blue butterfly. e4 should have been left blank on manufacturer device report 1220648-2024-25103.
12716 | 1220648-2024-25104 | 21014436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21014436&pc=OZD | 12/12/2024 | 02/05/2025 |  | 12/24/2024 |  |  | 12/12/2024 | 12/24/2024 |  |  |  |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651060 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a controller error outside of patient support. there was no patient involvement. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller error-yellow alarm has been completed. data logs were reviewed which confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. console was sent back for investigation but the controller error failure mode was not reproduced. rlm was able to connect to ic_test and send video to impellaconnect.com but once a pump was connected, the rlm rebooted itself over multiple tests. eventually, the rlm no longer rebooted itself shortly after plugging in a pump and the rlm was run overnight. there were multiple rlc reboots overnight but the module re-connected each time to ic_test as designed. 15 hours into the test, the usb power cycled which caused the module to reboot, the firewall to reset and load back not connected to ic_test also by design. the current to the rlm is around 0.5 amps well below the 1.25a threshold for the usb. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
12717 | 1220648-2024-25105 | 21014438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21014438&pc=OZD | 12/12/2024 | 01/21/2025 |  | 12/24/2024 |  |  | 12/12/2024 | 12/24/2024 |  |  | 12/31/2024 |  | 12/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291837 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | BOSTON SCIENTIFIC V-18 CONTROL WIRE 0.018 IN | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump to support a surgical patient. the patient had the pump placed for planned mitral valve repair and coronary artery bypass graft. the pump was placed via axillary after care escalated from an intra-aortic balloon pump. the 5.5 pump had an access site bleed from the graft and was treated by heavy tie and graft lock repositioning. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since non-abiomed product like bos. sci. v-18 control wire 0.018 inch was used. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added d10-concomitant medical products.
12718 | 1220648-2024-25044 | 20999326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999326&pc=OZD | 12/09/2024 | 01/16/2025 |  | 12/23/2024 |  |  | 12/09/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant that was aborted due to patient anatomy. the patient¿s outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since the implant of the pump was aborted due to patient's anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12719 | 1220648-2024-25045 | 20999718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999718&pc=OZD | 12/05/2024 | 02/07/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1259338 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during preparation, the automated impella controller (aic) had a system self-check failure at start up. consequently, the aic was exchanged. there was no report or indication of delay in treatment or harm to the patient. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. according to the device and data analysis the root cause of the system self check failure was an impellatronic and power battery manager (pbm) board malfunction.
12720 | 1220648-2024-25046 | 20999784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999784&pc=OZD | 12/05/2024 | 01/21/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538077 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Respiratory Arrest | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support, experienced a hematoma, and later expired. the patient was brought to the catheterization lab where the left main artery was found to be 99% occluded. the physician elected to place an impella pump for hemodynamic support. the impella cp was placed to the right femoral artery without an issue. percutaneous coronary intervention was completed with a drug-eluting stent in the left anterior descending to the left main. post impella and stent placement, the patient suffered a respiratory code. cardiopulmonary resuscitation was performed. the patient was intubated and return of spontaneous circulation was achieved. the left femoral artery was accessed to determine if the stent was patent, and it was patent. impella placement was rechecked with satisfaction. the patient was transferred on support to another hospital. upon arrival, echocardiogram was completed and the physician was happy with the impella placement. however, oozing from the groin was noted and was resolved with manual pressure. the following day, an echocardiogram was completed and showed an ejection fraction of 40% and good impella position. distal pulses were present. however, a decrease in urinary output was noted, and the patient had been spiking fevers. computer tomography scan showed a retrocolic hematoma. treatment to address is unknown. over the next days, the impella position was noted to be optimal with appropriate flows and distal pulses present with doppler. however, the patient would expire on the fourth day of support. at one point, the patient was deemed not a candidate for escalation, and the family decided to withdraw care. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is determined to be patient movement. the patient was transferred by flight to ormc and bleeding was resolved by manual pressure.
12721 | 1220648-2024-25049 | 20999881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999881&pc=OZD | 12/05/2024 | 01/21/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525357 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp device. despite attempting to pre-treat the vessel, the staff was unable to insert the 14 fr introducer and pump for implant. the implant was aborted as a result of the resistance. there was no patient harm. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues and the difficult insertion has been completed. the product was not returned for investigation. according to the clinical details, the patient¿s vessel was treated with a balloon and several other methods, but the introducer and pump could not be inserted. resistance was mostly felt at the innominate artery. the cause of the delivery issue was most likely patient condition related to diseased/small vessels. the cause of difficulty in inserting/removing the introducer was most likely related to the patient's condition, based on the given clinical details indicating that the patient had diseased/small vessel condition.
12722 | 1220648-2024-25047 | 20999891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999891&pc=OZD |  | 01/10/2025 |  | 12/23/2024 |  |  | 10/22/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542953 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automatic impella controller triggered a controller error while supporting the patient. the patient was escalated to the impella 5.5 from impella cp device without complications and transferred to the unit for further management and recovery. after start of support (day unknown), there was a loss of placement signal and a controller error for less than 30 seconds. it was noted, however, the impella 5.5 pump continued to flow with no change in the flow and without change in the patient's hemodynamics. the patient was reported to be in stable condition following the event. information indicates, the aic remained with the impella 5.5 pump to support the patient until successfully weaned and explanted approximately six days after start of support. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. after reviewing the console and the date logs the cause of the aic issue was a software issue. e4 should have been left blank on manufacturer device report 1220648-2024-25047. | The automated impella controller has been received from the customer; however, the evaluation and analysis of the device is not complete. upon investigation closure, a supplemental mdr will be filed. note: the event occurred on or between impella 5.5 device implant and explant dates of (b)(6) 2022 - (b)(6) 2024. the actual date of the event is unclear and has been represented as blank.
12723 | 1220648-2024-25048 | 20999900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999900&pc=OZD | 12/05/2024 | 01/23/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025439876 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 for mechanical circulatory support and approximately 12 days after start of support, the patient was noted in critical condition with various bleedings in the nasopharyngeal area. the physician paused anticoagulation. additionally, there was sudden high purge pressure alarm triggered. the purge flow was 1.5 ml/h and purge pressure was 1070mmhg. motor current was stable under performance level p-6 with 340/300ma. two days later no improvement in the purge flow and purge pressure was noted, and palliative care was initiated due to the patient's poor prognosis. following, the care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of high purge pressure and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: high purge pressure: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ e4 should have been left blank on manufacturer device report 1220648-2024-25048. h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2024-25048.
12724 | 1220648-2024-25050 | 20999916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999916&pc=OZD | 12/06/2024 | 02/07/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568288 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing for an electrophysiology (ep) ablation patient. the ablation catheter has become entangled with the pigtail of the impella and has torn off. the ablation was nevertheless successfully performed and the pigtail was successfully retrieved. the patient survived the event and no harm noted. | [5 empty] | The investigation of product damage (detached inflow/outflow ¿ great vessel has been completed. the pump was returned for investigation. a data log review was not required for this case. the pigtail of the pump was found sliced on the returned product. no further investigation was required. according to the clinical report, the ablation catheter became entangled with the impella pigtail, tore off, but the procedure and retrieval were successful. the cause of the cannula/pigtail detachment issue was an interaction with another device(ablation catheter). d9 date device returned to manufacturer added. instructions for use : impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controller section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: warnings: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ e4 should have been left blank on manufacturer device report 1220648-2024-25050. h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2024-25050. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12725 | 1220648-2024-25051 | 20999920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20999920&pc=OZD | 12/06/2024 | 01/13/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient lost flow down the impella leg. additionally, the patient had elevated creatine levels with no urine output however, the presence of hemolysis was not confirmed. the impella was removed. the patient was stable after the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and the limb ischemia has been completed. data logs showed generally pump ran without issue during support. there were motor current (mc) spikes, positioning and suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12726 | 1220648-2024-25052 | 21000120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000120&pc=OZD | 12/09/2024 | 01/13/2025 |  | 12/23/2024 |  |  | 12/09/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577780 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient lost pulses to the left lower extremity. vascular was consulted. no further information was provided. the patient expired after 104.08 hours of patient support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12727 | 1220648-2024-25053 | 21000125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000125&pc=OZD | 12/08/2024 | 01/31/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Loss of Data | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Complainant in japan reported the automated impella controller (aic) had optical signal issues. the position waveform suddenly disappeared while on support. the aic had been used beyond its service life and the staff replaced the aic. after replacement, the position waveform appeared and support continued. there was no patient harm. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue has been completed. according to data and device analysis the root cause was a defective optical bench. during the investigation a controller error was discovered. the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)¿ was determined to be a firmware issue. g2 report source; foreign was missing from manufacturer device report 1220648-2024-25053.
12728 | 1220648-2024-25054 | 21000138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000138&pc=OZD | 12/10/2024 | 02/03/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) had a controller shutdown. this issue was discovered during preparation. no further information was provided. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. the root cause of the intermittent, unexpected controller shutdown could not be determined, as there was no clear indication of the cause, and the issue could not be reproduced. e4 should have been left blank on manufacturer device report 1220648-2024-25054.
12729 | 1220648-2024-25055 | 21000234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000234&pc=OZD | 12/05/2024 | 09/30/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548454 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that when finishing impella 5.5 insertion the patient started having atrial ectopy, perfusing every few beats. amio was given and the impella was pulled back and the new measurement was 5cm across atrioventricular and 44cm at skin. the procedural outcome was the patient survived surgery. | The complainant also reported that there was a placement signal not reliable alarm. the audio was disabled, and the team monitored motor current and flows. the patient was stable. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for analysis. the data logs confirm placement signal alarms. no identified optical signal issue was identified upon evaluation of the 5s parameters. contrast and the signal-to-noise ratio (snr) decreased while placement signal was out of range and gain, lamp, and light were stable at the time of the alarm. the raw optical signal was increasing, suggesting a possible increase in pressure in the heart. the device was not removed due to this. the root cause of the optical signal issue was not determined as no product was returned for analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-25055. b7 other relevant history, including preexisting medical conditions updated. d5 updated. d10 concomitant medical products and therapy dates updated.
12730 | 1220648-2024-25056 | 21000258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000258&pc=OZD | 06/21/2023 | 12/23/2024 |  | 12/23/2024 |  |  | 06/21/2023 | 12/23/2024 |  |  | 01/31/2024 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023171575 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Death | Yes | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. post impella implant, the patient condition was classified as critical and the patient who had been on extracorporeal membrane oxygenation (ecmo) since before impella support, was switched to venovenous ecmo. five days into impella support, it was reported there was a small leakage at the purge and low purge alarms were triggered. the physician was advised on a purge bypass procedure, or to replace the pump. the medical team determined pump replacement was not an option, and the plan of care was to attempt to wean the impella within the next two days. three days later, per the icu physician, the patient was in septic shock and the patient was started on dialysis. the purge leak issue was unchanged, with reported acceptable pump values. the following day, it was noted that the pump was running with no issues, but the patient continued to decline. the patient continued to experience septic shock and multiple organ failure. the purge arm leak remained unchanged, and the pump was running a p-7 with no issues reported. it was noted that the patient later expired this day. additional information was received regarding the sepsis and it was noted the patient developed sepsis due to the surgical procedure (bypasses) and the subsequent centrally placed ecmo with initially only a temporarily closed thorax. given the totality of the information, cause of death is more likely sepsis and multiple organ failure per the treating physician. per the treating physician the sepsis appears to have been caused by bypass surgery and subsequent centrally placed ecmo and there was no allegation or concern that impella caused or contributed to the patient's sepsis. however, given the purge sidearm leak, the consequential low purge pressure and the critical status of the patient, the continual purge sidearm leak/low purge pressure cannot be excluded from the contributing to the outcome. | [5 empty] | The cause of the purge arm leak was a damaged sidearm due to an unknown reason since the complaint device not returned for investigation. insufficient data provided to determine a root cause. this report is being filed as part of a retrospective review of historical records.
12731 | 1220648-2024-25057 | 21000406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000406&pc=OZD | 01/17/2023 | 10/20/2025 |  | 12/23/2024 |  |  | 01/29/2023 | 12/23/2024 |  |  | 08/31/2024 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023256977 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Death | Life Threatening | Other | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to a decision for patient care (recovery vs left ventricular assist device vs palliative care). it was reported after the delivery and just after initiation of support of the impella 5.5, the patient experienced ectopy which was caused by the position of the impella. the impella was repositioned and left deeper than recommended in the left ventricle to avoid further ectopy. the following day on (b)(6) 2023 the patient experienced ectopy when moving, however it was noted the patient was still well supported on the impella. the patient was scheduled to verify impella position, however it is unknown if the impella required repositioning. two days after, on (b)(6) 2023, the patient experienced more episodes of ectopy, however it was noted that the patient was hemodynamically stable. the impella was found to be close to the septum, and the medical team considered repositioning the impella, however it is unknown if the pump was repositioned. it was noted that the patient's overall health was very poor, and the patient had been placed on total parenteral nutrition (tpn) several days prior. on (b)(6) 2023 the patient was noted to still have ectopy episodes, with no noted resolution however the patient was noted to be hemodynamically supported. on (b)(6) 2023, the performance level of the pump was lowered to p-3. and this was noted to improve the patient's ectopy, however when laying flat the patient experienced premature ventricular contractions/ventricular tachycardia. on (b)(6) 2023 it was reported a clot was visualized on the impella 5.5 inlet cage with transesophageal echocardiography that was performed on (b)(6) 2023. it was additionally reported that the patient's heparin had been intermittently withheld due to a low platelet count. the patient's family decided to withdraw care and pursue comfort measures only. the patient expired shortly after care was withdrawn. there is not enough evidence to exclude impella as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: updated manf. email address.
12732 | 1220648-2024-25058 | 21000453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000453&pc=OZD | 12/05/2024 | 01/08/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569601 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a damp area on the filter where the purge tubing connects. the tubing was changed however, after a night it was damp again. it was recommended to change the purge composition and to replace the pump however support continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge-pump leak has been completed. no product or images were returned for analysis. the root cause of the purge leak - pump was not determined as no product or images were returned and insufficient clinical information was provided.
12733 | 1220648-2024-25059 | 21000611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000611&pc=OZD | 01/01/2024 | 01/16/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  | 10/31/2025 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024411412 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed's long-term outcome and quality indicator impella registry indicating that a patient endured a progression of ventricular tachycardia during impella 5.5 support. when the tachycardic condition persisted, amiodarone was started and a cardioversion was completed. the condition was deemed to be possibly related to the use of the impella device. the patient was stable following the intervention and survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of ventricular tachycardia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g.2 study. selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25059 and was added.
12734 | 1220648-2024-25060 | 21000709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000709&pc=PYX | 12/06/2024 | 01/13/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025437965 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp flex for the support of a patient admitted in for a high-risk percutaneous coronary intervention. impella rp flex was placed but during the replacement, there was significant blood loss. the bleeding caused a hemoglobin drop so the team gave two units of blood. after 25 hours of support, the team explanted and escalated to venoarterial extracorporeal membrane oxygenation. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual. section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12735 | 1220648-2024-25061 | 21000819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000819&pc=OZD | 12/05/2024 | 02/07/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506301 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during the implant, the patient had access site bleeding. the access site was from an intra-aortic balloon pump (iabp). a cut down was needed in order to replace the iabp with the impella. blood products were given. it was further reported that the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details provided. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25061. e.1 revised facility name as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25061. g.1 revised reporting contact email as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25061 h.6 code 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25061 and was added.
12736 | 1220648-2024-25062 | 21000898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000898&pc=OZD | 12/10/2024 | 01/22/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555930 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was attempted to be placed but the case was aborted due to the patient¿s vascular anatomy. impella cp was placed instead to the right axillary artery and the patient was cannulated for peripheral venous-arterial extra corporeal membrane oxygenation as well. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, due to patient¿s vascular anatomy being torturous, impella implant was unsuccessful.
12737 | 1220648-2024-25063 | 21000943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21000943&pc=OZD | 11/30/2024 | 02/06/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553380 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Necrosis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed for a patient admitted in with coronary artery disease. the pump was placed but after 10 days the patient expired from intestinal necrosis. the direct cause of intestinal necrosis is unknown, so the relationship with impella is unclear. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - great vessel has been completed. the root cause of the vascular damage - peripheral issue was not determined since insufficient clinical information provided and the relation to impella is unknown. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25063.
12738 | 1220648-2024-25064 | 21001001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001001&pc=OZD | 12/06/2024 | 01/23/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572042 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient with acute myocardial infarction and cardiogenic shock. the pump was inserted and found to have pump flow issues. the pump was removed and replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation low pump flow has been completed. data was not returned for review. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was most likely biomaterial ingestion.
12739 | 1220648-2024-25065 | 21001328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001328&pc=OZD | 12/06/2024 | 02/03/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548194 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 14f peel away sheath could not be forwarded in the pelvic anatomy. the calcium load was greater than expected, so a kink occurred, and the sheath had to be discarded as it was defective. the procedure was completed without impella and there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. e.1 name prefix/title, first name and last name are unknown. | The investigation of introducer damage ¿ mechanical has been completed. as per the given clinical, the peel away sheath couldn't be advanced due to excessive calcium load, causing a kink. the cause of the introducer damage issue was most likely patient condition related to diseased vessel. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25065.
12740 | 1220648-2024-25066 | 21001644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001644&pc=OZD | 12/07/2024 | 01/30/2025 |  | 12/23/2024 |  |  | 12/07/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569836 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump being placed to support a patient admitted in acute myocardial infarction/cardiogenic shock. while prepping the impella, the scrub technician forgot to connect the yellow to yellow leur causing the pump not to prime. the pump was then inserted through the sheath into the body. at that point, the scrub technician realized the mistake, removed the pump, connected the yellow to yellow and attempted to prime. per the report, the pump would not prime at that point. that pump was aborted and a new pump was used. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the priming issue has been completed. data logs were not available for review. returned product was investigated, though only the pump was returned, not the purge cassette. there were no visual abnormalities noted, so the pump was connected to a purge cassette in lab and successfully primed. it was then run for 3 hours without any issues. clinical details report that after not following the priming procedure correctly initially, they attempted to restart and the pump could not be primed. without data logs to review the exact timeline of events in the case and the purge cassette being returned, the root cause of the priming issue could not be determined.
12741 | 1220648-2024-25067 | 21001697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001697&pc=OZD | 12/06/2024 | 01/27/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551876 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support an acute myocardial infarction / cardiogenic shock patient. the pump was placed and was supporting for approximately two weeks when the patient was observed to have a cerebrovascular accident with herniation. the patient had possibly underlying brain lesions. the pump additionally stopped after the 12 days and was unable to restart the pump. the family decided to withdraw care, and the patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into stroke/cva and pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12742 | 1220648-2024-25068 | 21001756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001756&pc=OZD | 12/10/2024 | 02/03/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) experienced a controller error alarm during patient support. the aic was swapped and no further issues were encountered. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) controller error-yellow alarm has been completed. console logs from the reported day of event are consistent with complaint and confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. reported controller error alarm and temporary loss of placement data was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
12743 | 1220648-2024-25069 | 21001822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001822&pc=OZD | 12/06/2024 | 01/16/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1706584 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that the automated impella controller (aic) had been locked in a closet and not plugged in for an extended period of time. when turned on, a battery failure alarm was displayed, and the aic was unable to recharge when plugged in. no patient harm has been reported. | [5 empty] | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. after reviewing the imc logs, the reported battery failure issue was confirmed. the console was tested after arriving in field service. battery failure alarm issue was able to be reproduced. the root cause of this issue was depleted batteries which resulted in battery lockout. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12744 | 1220648-2024-25070 | 21001846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001846&pc=OZD | 12/10/2024 | 01/15/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, there was no flow down the right side (impella side) while the peel away was left in. the provider opted to place a six french antegrade sheath placed for distal perfusion cannula, connected via male connector to the left femoral artery in case no flow following peel away. limb ischemia was resolved. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12745 | 1220648-2024-25071 | 21001965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21001965&pc=OZD | 12/06/2024 | 01/27/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  | 09/30/2025 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024400653 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that three days post impella 5.5 explant, patient suffered catastrophic embolic stroke while ambulating around unit prior to discharge. per the doctor, a computed tomography (ct) scan showed residual thrombus in affected axillary artery extending into the aortic arch. the doctor feels that the stroke is impella related. the patient was appropriately anticoagulated once surgical bleeding with controlled pod 3. anti-xa was used with hospital protocol for heparin titration. patient remained on heparin for duration of impella therapy once started, and it was stopped at removal. patient remained in critical condition in the intensive care unit at this time. after impella was explanted, trans-esophageal echocardiography was performed at the time of explant and did not show any clot of surgical aortic valve replacement or in the left ventricle. heparin was continued into the operating room and stopped at time of explant. anti-xa was therapeutic at that time at 0.51. heparin was not restarted post-procedure. biomaterial was removed from the residual graft at the time of explant; there was no biomaterial on the impella itself on removal. after the patient suffered the cerebrovascular accident, imaging showed a basilar artery large vessel occlusion. a ct scan also showed a large thrombus in the right axillary artery proximal to the site of the graft anastomosis. thrombectomy was performed, however the patient was unable to have any meaningful neurological recovery and was transitioned to hospice care on (b)(6) and ultimately expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the stroke and thrombus has been completed. no product was returned. the data logs were reviewed and there are no motor current spikes, a trend in the placement signal, positioning alarms or changes in flows without a p-level change. there was no ingestion seen in the logs. the root cause of the thrombus and the stroke could not be determined without additional clinical details and the product evaluation.
12746 | 1220648-2024-25072 | 21002273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21002273&pc=OZD | 12/06/2024 | 01/22/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and they were unable to advance past innominate. the decision was made to abort the implant and there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass past innominate artery due to calcification in the vessels. h6 medical device problem code 2682 was reported omitted on manufacturer device report 1220648-2024-25072.
12747 | 1220648-2024-25073 | 21002956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21002956&pc=OZD | 12/06/2024 | 02/11/2025 |  | 12/23/2024 |  |  | 12/06/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing when there was an observation made of thrombus in the left ventricular (lv) on the day after insertion. the team did not treat the lv clot and support continued until explant at day four of the support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25073.
12748 | 1220648-2024-25074 | 21004520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21004520&pc=OZD | 12/05/2024 | 09/24/2025 |  | 12/23/2024 |  |  | 12/07/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the automated impella controller (aic) had a controller error that resolved on its own in a few seconds. there was no harm to the patient. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) controller failure-yellow alarm has been completed. data logs were reviewed finding that the console displayed ¿placement signal (ps) not reliable (out of range - optical) ¿ alarm,¿ event #1048. real time log confirmed the ps was spiked to 3000. this aligns with the reported failure mode. then console displayed ¿placement signal not reliable (opm invalid signal, optical pump connected]) ¿ alarm,¿ event #1036. which was resolved and reoccurred several times throughout the case. after 15 hours, during support, the console displayed ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045, which was resolved immediately. rt log confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. then, the console and pump were used for another 16 hours and 27 minutes. the console was returned for evaluation. upon running the optical test pump for 1 day 14 hours, unable to reproduce either the psnr or controller error event #1045. the analog output from the optical bench was good, this confirms no issue with the optical bench. found the front panel optical connector contaminated, unable to remove the debris upon cleaning. it is most likely unrelated to the reported failure mode. the root cause for the psnr was not determined due to the inability to reproduce. the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has already been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to contain optical signal coding. d6a / d6b updated. d10 concomitant medical products and therapy dates updated.
12749 | 1220648-2024-25075 | 21006368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21006368&pc=OZD | 12/07/2024 | 01/14/2025 |  | 12/23/2024 |  |  | 12/07/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567249 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support ongoing for a patient admitted in with multiple organ failure and cardiomyopathy. the pump was placed but after just over 1 month the team observed a cerebral hematoma with pupil changes. the team made choice to change heparin in purge to sodium bicarb and monitor for cerebral vascular accident with computed tomography imaging. pump remains on for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
12750 | 1220648-2024-25076 | 21008946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21008946&pc=OZD | 01/02/2024 | 09/19/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  | 10/31/2025 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406225 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used. the patient was a transfer in for higher level of care with mixed presentation of cardiogenic shock and sepsis with worsening lung function. ejection fraction at transfer was 30% and the impella 5.5 was planned for mechanical circulatory support. the impella 5.5 was implanted without incident with future plans for high-risk percutaneous coronary intervention (pci). next day, however, after start of support, likely hemolysis in the setting of impella use was noted. action taken was a total of four units of packed red blood was administered over the course of ten days. the protected pci to the left main artery was completed in between the course of blood transfusions. the patient still required pci of the right coronary artery however, which was noted to be planned for the end of the said week. impella mechanical circulatory support was completed. per the registry, the patient recovered from the event with no sequelae. the patient was successfully weaned, and device was explanted without complications. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. the audio was disable and the team monitored motor current and flow. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. no product was returned.log review did not show any motor current spikes indicating biomaterial ingestion. there were also no positioning issues based on log review, and the placement signal did not exhibit any trends indicating deteriorating patient condition. blood was given to the patient throughout the case, starting on the second day of support. however, it is unknown when the hemolysis resolved in relation to when blood volume was given. the root cause of the hemolysis was not determined since the product was not returned and insufficient clinical details were provided. e.1 revised facility name as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25076. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25076 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2024-25076. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data log review showed a "placement signal not reliable" alarm and placement signal going out of range. the root cause of the optical signal issue was most likely a damaged sensor as shockwave was being used when the optical signal issue occurred. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-25076.
12751 | 1220648-2024-25077 | 21008948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21008948&pc=OZD | 03/09/2024 | 01/14/2025 |  | 12/23/2024 |  |  | 12/05/2024 | 12/23/2024 |  |  | 11/30/2025 |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428652 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp support ongoing for a patient admitted in for a high risk percutaneous coronary intervention. the pump was placed and supported for just under 3 hours and was explanted. the patient was enrolled in abiomed¿s long-term outcome and quality indicator (loqi) impella registry database, and 11 months after the explant, the team noted ventricular fibrillation with a possible relationship to the use of the impella. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of ventricular fibrillation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issue could not be determined. b.2 revised selection as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25077. b.7 added information as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-25077. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25077 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.2 study selectionwas inadvertently omitted from manufacturer device report 1220648-2024-25077 and was added. h.6 codes 1721 ans 4641 were reported incorrectly on the initial manufacturer device report number 1220648-2024-25077 and a new code was added.
12752 | 1220648-2024-25078 | 21009052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21009052&pc=OZD | 12/07/2024 | 01/24/2025 |  | 12/23/2024 |  |  | 12/07/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569572 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with aortic valve surgery. the pump has been supporting the patient for 3 weeks when there was elevated purge pressure alarms observed. team attempted to troubleshoot by checking for kinks and cassette replacement. the alarm was disabled and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and lt logs show a rise in purge pressure over approximately 3 hours before high purge pressure alarms are triggered. motor current and pump flow are slightly variable after purge pressure rise and purge pressure remains high. clinical details noted that cassette change was performed and did not resolve issue. no kinks were observed in purge line in the field. d5w with sodium bicarbonate was being used in purge at time of purge pressure rise and alarm. the root cause of the high purge pressure cannot be definitively determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen."
12753 | 1220648-2024-25079 | 21009057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21009057&pc=OZD | 12/07/2024 | 01/13/2025 |  | 12/23/2024 |  |  | 12/07/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573233 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient had a brain bleed while on support. heparin was discontinued. furthermore, while turning the patient, the impella came out of the ventricle. the impella was completely in the aorta and the doctor did not want to advance the pump. the decision was made to remove the pump. hemostasis was achieved after explant, however the patient expired shortly after the pump was removed. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleed and positioning issues has been completed. product was not returned for review. based on clinical description, the cause of the positioning issue was most likely use related with patient movement leading to pull pump out of aorta. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12754 | 1220648-2024-25080 | 21009069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21009069&pc=OZD | 12/08/2024 | 01/23/2025 |  | 12/23/2024 |  |  | 12/08/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454040 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) alarmed, ¿purge disc not detected¿. the cassette was removed and reinserted, but the alarm persisted. advised to use back up console and there was no harm to the patient. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12755 | 1220648-2024-25081 | 21009075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21009075&pc=OZD | 12/10/2024 | 01/29/2025 |  | 12/23/2024 |  |  | 12/10/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557909 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. roughly 24 days into support, an echocardiogram showed that the pump positioning had shifted with the tip turned toward the mitral valve. the pump appeared to be engaging the papillary muscle in this new position. attempts to reposition the device failed and a surgeon was subsequently consulted to resolve the issue. the following day, the patient suffered from a stroke. support continued for another two days before planned explant. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-25081.
12756 | 1220648-2024-25082 | 21009078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21009078&pc=OZD | 12/11/2024 | 01/31/2025 |  | 12/23/2024 |  |  | 12/11/2024 | 12/23/2024 |  |  |  |  | 12/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580386 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, a hematoma was noted at the impella site. a sandbag was placed on the groin and manual pressure was applied. anti xa was 1.9. additionally, the patient developed hemolysis, urine red. p level was dropped to p5. the doctor was informed of both the hematoma and hemolysis and decided to drop the p level to 4 and take the patient back to the catheterization lab to explant the impella. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and access site bleed has been completed. no product was returned for evaluation. data logs were reviewed and real time log at the beginning of the case shows a small step up in motor current, simultaneous small speed deviation, pump flow becoming slightly more saturated, and a slight step down in placement signal. additionally, a slight drop in purge flow with a small increase in purge pressure was observed at the same time. these data log signature align with a small ingestion of biomaterial into the pump. clinical details noted that red urine was observed after patient check in into unit, pump was turned down from p-7 to p-5. pump was running with d5w at time of hemolysis and sodium bicarbonate purge solution was ordered. it was not noted if hemolysis resolved after pump was explanted. the root cause of the hemolysis was most likely due to biomaterial. clinical details noted that patient with triple vessel disease and known cad had puffiness around the impella access site during pci procedure. at check into unit, a hematoma was present at the access site. the root cause of the access site bleeding cannot be definitively determined as limited clinical details were provided.
12757 | 1220648-2024-25009 | 20986883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20986883&pc=OZD | 12/05/2021 | 12/20/2024 |  | 12/20/2024 |  |  | 11/05/2021 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351867 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella 5.5 for mechanical circulatory support during a percutaneous coronary intervention. after the impella 5.5 was implanted, upon connecting the impella to the automated impella controller (aic), there was an error on the aic for "optical sensor not supported". the white cable was disconnected and reconnected, but the aic would not recognize the pump. the alarm never cleared. there is no further information regarding a back up console and how long before a backup console was obtained and impella was connected. the pump could not be initiated prior to therapy initiation and it is unknown how long of a delay was in place given this event. it cannot be ruled that there was no impactful delay and untoward effects to the patient. information around the outcome, resolution for this event is limited. the decision was later made to withdraw care noting that the patient was unlikely to survive needed revascularization. the patient expired while on impella 5.5 support. the patient's presenting condition may be more likely have impacted the event, however there is not enough evidence to exclude the aic as an associated factor in the patient's demise. | [5 empty] | Data analysis: on the reported complaint occurrence date, automated impella controller (console) logs revealed that the console presented the window/ error ¿optical sensor not supported¿ when an impella 5.5 was connected. device analysis: the issue of console ic6194 not supporting optical pump was due to an undetermined hardware/software glitch that occurred during console's preventative maintenance procedure 3 months prior to the event.: pmd fw version bit was reset between power-cycles, which disabled the optical bench. after re-setting the pmd fw bit to its correct value the optical bench was operational, and the issue was not reproduced after subsequent 100 power-cycles. this case is related to impella 5.5 device with serial number (b)(6). an additional report will be submitted for this device. this report is being filed as part of a retrospective review of historical records.
12758 | 1220648-2024-25010 | 20986966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20986966&pc=OZD | 12/05/2024 | 02/03/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was a spike in purge pressure and there was saturation. the impella cp stopped. the pump was powered down and pulled back into the descending aorta. the impella cp was removed and the patient was stable. the customer would also like the automated impella controller (aic) evaluated for any possible relationship to the pump stop. the patient survived the event. | [5 empty] | The investigation for the console (aic) pump stop has been completed. the console log confirmed the reported failure, showing multiple instances of the "purge pressure high - alarm" and "purge system blocked - alarm" during the clinical procedure. additionally, the "impella stopped (motor current high) - alarm" was triggered only after the pump was unplugged from the console. rt logs indicated inconsistencies in the purge flow ticks and the purge flow. the console was returned for evaluation. visual inspection of the internal circuitry of the console did not reveal any anomaly. the console was run with a known good pump and purge system with no observable anomaly. the root cause for the purge pressure high - alarm"and "purge system blocked - alarm" ) was not related to any problems with the aic. no issues were found within the aic. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. the pump associated with this automated impella controller is catured under manufacturer device report number 1220648-2024-24960.
12759 | 1220648-2024-25011 | 20987055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987055&pc=OZD | 12/05/2024 | 01/21/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546208 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock and care escalation from an impella cp. the pump had high purge pressures on day 10. the team tried to troubleshoot with replacing the purge cassette but that failed and no kinks were seen. the team reviewed the tissue plasminogen activator addition to the purge along with the already running sodium bicarbonate. the patient remains on impella support. | [5 empty] | The investigation for the high purge pressure has been completed. pump and data log were not returned for investigation. purge pressure was reported high on 1oth day of use. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12760 | 1220648-2024-24981 | 20987162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987162&pc=OZD | 02/14/2022 | 12/19/2024 |  | 12/20/2024 |  |  | 10/23/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116341 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, there was controller alarm. the impella cp pump remained implanted for support despite this without harm or injury. | [5 empty] | The investigation for the reported loss of opm data issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. device analysis: console was sent back for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. per the results of the clinical review and investigation, the root cause was not determined since it could not be reproduced. this report is being filed as part of a retrospective review of historical records.
12761 | 1220648-2024-24982 | 20987230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987230&pc=OZD | 07/18/2022 | 12/19/2024 |  | 12/20/2024 |  |  | 10/15/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported automated impella controller (aic) had placement signal (ps) loss during support. troubleshooting did not resolve the issue; therefore, aic was replaced by the backup console. no harm or injury to the patient was reported. while investigation the issue, opm failure was noted. | [5 empty] | The investigation for the reported placement signal issue has been completed. the product was returned, and the placement signal issue was not confirmed. imc logs confirmed the reported failure mode given there were controller error alarms (opm comm crc error ¿ event set#: 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. console (ic3165) was sent back fs for further investigation. the failure mode could not be reproduced. this is a known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm. this report is being filed as part of a retrospective review of historical records.
12762 | 1220648-2024-24983 | 20987386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987386&pc=OZD | 06/05/2023 | 12/19/2024 |  | 12/20/2024 |  |  | 06/05/2023 | 12/20/2024 |  |  | 02/28/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340963 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, leaking was observed from the purge reservoir side arm. attempts were made to change the purge cassette and fluid but issue was not resolved. purge bypass was performed. once completed the purge pressure was significantly increased and the purge flow decreased. impella was successfully explanted and manual pressure held, two perclose deployed, and femostop put in place for additional support. no patient harm was reported. | [5 empty] | The investigation for the reported "purge leak - pump" has been completed. the product was returned and the "purge leak - pump" was not confirmed. leak test was performed on pss & leak was found from the luer neck of filter. the cause of the pss leak was due to a crack at the filter joint. this report is being filed as part of a retrospective review of historical records.
12763 | 1220648-2024-24984 | 20987448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987448&pc=OZD | 06/30/2023 | 12/19/2024 |  | 12/20/2024 |  |  | 06/30/2023 | 12/20/2024 |  |  | 12/31/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325712 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the nursing staff turned the patient to clean him up. it was a lengthy clean up and began getting suction alarms, waveforms changed. it was noted that the impella got pulled back into aorta. the patient went to the cardiac care lab (ccl) for repositioning. unsuccessful reposition. the pump was removed over .035 wire and a new 14 peel away and a new pump was inserted. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as per clinical observations, the pump was mispositioned during a lengthy cleanup process. the doctor attempted to reposition it, but the pump was eventually moved back into the aorta. the cause of the positioning issue was patient movement based on given clinical information. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
12764 | 1220648-2024-24985 | 20987471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987471&pc=OZD | 04/11/2023 | 12/19/2024 |  | 12/20/2024 |  |  | 04/11/2023 | 12/20/2024 |  |  | 10/31/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023275569 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that there was slight leaking of the purge side arm between the pressure reservoir and filter. there were no visual issues observed. the purge flow was 19 millimeters and pressure 380 millimeters of mercury. it was recommended bone wax and monitoring of the purge flow and pressure. the physician was sent purge reservoir and filter bypass documentation and it was emphasized the necessity of a pump change in this situation. the pump was explanted 5 days later, since the system was running stable and without further issues the physician didn`t want to change the pump. there was no patient harm reported. | [5 empty] | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. the cause of the issue was not determined since the product was not returned and insufficient clinical information was not obtained. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
12765 | 1220648-2024-24986 | 20987488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987488&pc=OZD | 06/04/2023 | 12/19/2024 |  | 12/20/2024 |  |  | 06/04/2023 | 12/20/2024 |  |  | 03/31/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341808 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient pulled back on the purge cable resulting in the sterile sleeve separating from the catheter exposing over 10 centimeters of the sheath. a hematoma formed at the site and due to the compromised non-sterile nature of the catheter, the decision was made to explant the pump. patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding and sterile sleeve damage has been completed. based on available information provided by the complainant, the root cause of the access site bleeding issue was determined to be patient movement. the root cause of the sterile sleeve damage issue was use related to patient movement. this report is being filed as part of a retrospective review of historical records.
12766 | 1220648-2024-25012 | 20987622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20987622&pc=OZD | 12/05/2024 | 10/02/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838437 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while on patient support, the automated impella controller (aic) plug, and the power cord came apart. the wires were hanging freely and the plug was still in the wall. a new aic was used. the patient remained stable and the procedural outcome was the patient survived the event. | [5 empty] | The investigation of automated impella controller (aic) - cabling issue has been completed. there was no evidence in the logs of a broken power cord but console was not returned for investigation. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2024-25012. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. data analysis: the console was able to plug and unplug ac power on (b)(6) and after but complaint did not allege loss of ac power. device analysis: the console was returned and the ac cord was found to have extra cable exposed. the cord was disassembled and there was found to be 1.5" of cable confirming the complaint. root cause: the cause of the issue was a defective cable. d.9. device returned to manufacture date added.
12767 | 1220648-2024-25013 | 20988033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988033&pc=OZD | 12/05/2024 | 01/22/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531185 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump for support of a post-operative coronary artery bypass graft/aortic valve repair patient. the pump was placed as care escalation from an intra-aortic balloon pump and the team considered impella cp-rp support. the impella cp was explanted and replaced after just over one hour of support due to lost placement signal. the patient survived the explant. | [5 empty] | The investigation for the optical signal issue has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. the pump passed the laser continuity test for the optical line. all the 5s snr parameters were within the specified range. the pump was tested in a bucket of water and ran as intended. the data log shows that the placement signal reached 3000 with psnr, then after the pump was inserted into the patient's body, the snr trended from 4000 to 0. the snr and contrast remained at very low limits and trendy throughout the case, without obtaining a placement signal. the cause of the optical signal issue was most likely patient condition related, based on returned product investigation and data log review. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12768 | 1220648-2024-25014 | 20988133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988133&pc=OZD | 10/01/2022 | 12/20/2024 |  | 12/20/2024 |  |  | 10/01/2022 | 12/20/2024 |  |  | 07/31/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238332 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding (oozing). the patient was switched to systemic anticoagulation. the patient was noted to have been successfully weaned off device without issues. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12769 | 1220648-2024-25015 | 20988138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988138&pc=OZD | 03/24/2022 | 12/20/2024 |  | 12/20/2024 |  |  | 03/24/2022 | 12/20/2024 |  |  | 11/30/2023 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169303 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was presented for escalation of care from the impella cp to the impella 5.5 post cardiotomy cardiogenic shock (pccs). impella 5.5 was placed without issue. during support, device had multiple days of suction issues, and the patient was given blood and albumin. it was also noted that the patient was converted from atrial fibrillation back to normal sinus rhythm. additional information noted that the device was explanted, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12770 | 1220648-2024-25016 | 20988153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988153&pc=OZD | 12/03/2024 | 01/13/2025 |  | 12/20/2024 |  |  | 12/03/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574130 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support placed for support of a patient admitted in with congestive heart failure. the patient was transferred to a tertiary center as they awaited decision on left ventricular assist device or transplant. the pump was supporting for over a week when there were signs of hemolysis in elevated plasma free hemoglobin and there was ventricular tachycardia that was treated by defibrillations. the pump remains on for support. | [5 empty] | The investigation of vf/vt/af/at and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12771 | 1220648-2024-25017 | 20988551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988551&pc=OZD | 12/04/2024 | 04/24/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025581159 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Reportable for the death per mso review and given the impella-related event, heparin in the purge exacerbated the situation for the patient, there is not enough evidence to exclude the impella as a contributing factor to the outcome. | The complainant had a impella 5.5 and extracorporeal membrane oxygenation (ecmo) support ongoing for a patient admitted in after transcatheter aortic valve replacement and suffering from acute myocardial infarction. the pump purge pressure (pp) rose and to troubleshoot the team adjusted the purge fluid and heparin concentration multiple times. the pump remained on for support despite the higher pp/purge system blocked. later, the patient had a tendency to bleed and was experiencing nosebleeds. managed with ecmo and intra-aortic balloon pump (iabp). heparin was added to the purging fluid, and activated clotting time was extended to over 250 seconds and then escalated to impella 5.5. nosebleeds continued and it was difficult to stop bleeding at the iabp removal site. hemoglobin dropped from 9 to 6. therefore, blood transfusion was performed. it was later reported that the patient expired and the death was unrelated to the impella 5.5 pump. |  |  |  |  | B2 revised from the manufacturer device report 1220648-2024-25017 to include death. b5 mso statement added. h1 type of report revised from the manufacturer device report 1220648-2024-25017 to include death. h6 health effect - impact code revised from the manufacturer device report 1220648-2024-25017 to include death. | The investigation for the high purge pressure and the vascular damage has been completed. pump was not returned for investigation. data logs were not returned for investigation as well. without the data logs and the pump, the failure hpp cannot be investigated. therefore, the root cause of the high purge pressure issue was not determined since product and data logs were not returned. clinical information provided mentions that the patient had a tendency to bleed and was having nose bleeds and bleeding from the iabp site. heparin had been added to the purge. the root cause of the vascular damage issue was not determined since product was not returned and lack of clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12772 | 1220648-2024-25018 | 20988561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988561&pc=OZD | 01/17/2022 | 12/20/2024 |  | 12/20/2024 |  |  | 01/17/2022 | 12/20/2024 |  |  | 09/30/2023 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141888 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, patient had bleeding and thrombocytopenia. the physician asked for systemic heparin drip to be held for three hours for bleeding at site. the patient was also heparin induced thrombocytopenia (hit) positive, so sodium bicarb was added to the purge solution. additional information noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleeding and thrombocytopenia has been completed. the root cause of the access site bleeding could not be determined as insufficient clinical details were provided. the root cause of the thrombocytopenia could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
12773 | 1220648-2024-25019 | 20988587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988587&pc=OZD | 04/06/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 04/06/2023 | 12/20/2024 |  |  | 09/30/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278101 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported the dry seal introducer did not allow for proper positioning and locking to the impella sheath. bleeding was noted around the sheath. additional information noted that the family decided to withdraw care and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported access site bleeding issue has been completed. no device was received for evaluation. the root cause of the access site bleeding issue was determined to be because of a non-recommended introducer kit, as the doctor accessed the insertion site using an unapproved introducer kit. only impella introducer kit is recommended per the instructions for use. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12774 | 1220648-2024-25020 | 20988660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988660&pc=OZD | 12/03/2024 | 01/22/2025 |  | 12/20/2024 |  |  | 12/03/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563963 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue. insertion of the impella 5.5 was attempted via axillary approach, however, the pump would not make innominate turn despite manipulation. patient is of small stature and size and decision was made by the doctor to explant the 5.5 pump and proceed with cp placement instead. the patient¿s outcome is unknown. | [5 empty] | The investigation for the delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related, impella 5.5 would not make innominate turn despite manipulation. patient is of small stature and size as per the clinical description provided. added d9 device return date. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12775 | 1220648-2024-25021 | 20988676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988676&pc=OZD | 12/11/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 12/04/2024 | 12/20/2024 |  |  | 09/30/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409914 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the prime screen and start screen were skipped when the pump was connected during setup. there was no harm to the patient and support was able to continue as planned with the noted pump. upon evaluation of the device, no noted issues were confirmed. however, the potential severity of the event was escalated. | [5 empty] | The investigation was completed. the pump and purge cassette were returned for investigation. no physical damage was observed on the returned product. the cassette was primed without any issues. from the returned log, it was confirmed that the pump case start was initially initiated on ic5543. case start activities were confirmed from the log, which showed that the pump was primed successfully and that case steps were completed in a specified manner. the pump was was not inserted into patient body. several unsuccessful attempts reported to repeat the case start. according to the clinical report, the automated impella controller setup wizard skipped the purge prime step. the root cause of the priming issue was determined to be no problem found based on data logs and returned product.
12776 | 1220648-2024-25022 | 20988738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988738&pc=PYX | 12/08/2024 | 01/28/2025 |  | 12/20/2024 |  |  | 12/08/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551239 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella rp placement one additional hemostatic suture was placed to stop oozing. the next day of impella placement the patient was confused and pulled the impella. suture at impella site had been broken and repositioning sheath had been pulled out slightly. decision was made to explant impella at bedside due to the patient complex history and no options for advanced therapies. the patient was made do-not-resuscitate and expired. | [5 empty] | The investigation for the positioning issue and access site bleed has been completed. no product was returned. clinical details state that when the pump was pulled by the patient, the suture at the impella site was broken and the repo sheath was slightly pulled out. after explant, the vascular resident placed a figure 8 suture to achieve hemostasis while also holding manual pressure. the root cause of the access site bleeding was most likely patient movement. therefore, the root cause of the positioning issue was patient movement. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
12777 | 1220648-2024-25023 | 20988753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988753&pc=OZD | 04/15/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 04/15/2023 | 12/20/2024 |  |  | 01/31/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315230 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after percutaneous coronary intervention (pci) was completed, physician noticed dissection to right femoral artery (rfa) from initial impella insertion, as well as major distal peripheral arterial disease (pad). the physician utilized re-access side port before pulling impella catheter back to sheath, then wiring from left femoral artery (lfa) access ¿up-and-over¿ to rfa to balloon lesion once impella was explanted. impella successfully explanted after weaning pump during procedure. once impella was removed, physician ballooned rfa with manual pressure held to rfa access site. the patient remained stable during procedure, only low dose levo continued. protamine administered while pressure was held to rfa. also, patient pedal pulses were not able to be found, even via doppler. upon nurse evaluation, the pedal pulses was still unable to be dopplered. | [5 empty] | The investigation into the reported issues has been completed. no device was received for evaluation. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the limb ischemia issue was patient condition related to pad vessel conditions. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
12778 | 1220648-2024-25024 | 20988762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988762&pc=OZD | 04/12/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 04/12/2023 | 12/20/2024 |  |  | 11/30/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295347 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient required cardiopulmonary resuscitation (cpr) while preparing to move from the cath lab to cardiovascular intensive care unit (cvicu). impella needed to be positioned several times as it tended to want to come out with any reduction in p-level. the physician elected to leave the impella in deep 5.5 cm on echo at bedside. however, there was inability for impella cp to maintain the position. patient escalated to impella 5.5 as an additional impella device was needed. | [5 empty] | No product was returned. as per clinical information provided, cpr was done for patient support before transferring from cath lab to cardiovascular intensive care unit. after cpr, positioning issue occurred and had to escalate to new product. the cause of the positioning issue was use related due to cpr being performed before the improper position occurred.
12779 | 1220648-2024-25025 | 20988819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988819&pc=OZD | 12/03/2024 | 02/06/2025 |  | 12/20/2024 |  |  | 12/06/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559188 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had a stroke on the ward after impella cp implantation was completed. the patient was transferred to a larger clinic and no other information provided. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12780 | 1220648-2024-25026 | 20988938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20988938&pc=OZD | 04/15/2021 | 12/20/2024 |  | 12/20/2024 |  |  | 04/15/2021 | 12/20/2024 |  |  | 03/31/2023 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022054563 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Life Threatening | Death | Other | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician did not feel like the patient had enough support. sirs suspected. they were unable to stabilize the patient and discussed escalating care. the patient expired after 2.40 hours of impella support. the relationship between the cause of death and the impella is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12781 | 1220648-2024-25027 | 20989174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989174&pc=OZD | 12/06/2024 | 02/11/2025 |  | 12/20/2024 |  |  | 12/06/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527197 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Injury | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and during the implant, the staff could not advance the sheaths. two different sheaths were used and the pump could not advance without kinking. the implant was aborted. the patient¿s outcome is stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella couldn't pass through the vasculature as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12782 | 1220648-2024-25028 | 20989279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989279&pc=OZD | 11/29/2024 | 01/31/2025 |  | 12/20/2024 |  |  | 12/09/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524959 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | Required Intervention | Death |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, ecpella support was continued, but hemorrhagic infarction occurred and life-saving was difficult. the patient later died. the cause of death was heart failure and percutaneous transluminal septal myocardial ablation (ptsma) was performed but the situation was difficult due to outflow tract stenosis. there was no problem with the operation of the impella. the cause of hemorrhagic infarction is unknown, and there is not one single cause such as anticoagulation and dual antiplatelet therapy (dapt). the causal relationship with the impella is unknown. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12783 | 1220648-2024-25029 | 20989346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989346&pc=OZD | 12/06/2024 | 02/12/2025 |  | 12/20/2024 |  |  | 12/06/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560413 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported an impella cp was being prepped for insertion when there was a leak to the purge side arm noted. a different impella was used successfully and there was no harm to the patient. | [5 empty] | The investigation for the purge leak-pump has been completed. no product or logs were returned for evaluation. clinical details noted that a leak from the purge sidearm filter was observed when prepping pump. the root cause of the purge leak was most likely due to a damaged sidearm. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | E.1 revised name prefix/title as it was previously entered incorrectly on the initial manufacturer device report number. g.1 revised reporting contact email as it was previously entered incorrectly on the initial manufacturer device report number.
12784 | 1220648-2024-25030 | 20989605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989605&pc=OZD | 12/09/2024 | 02/19/2025 |  | 12/20/2024 |  |  | 12/09/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568887 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had impella 5.5 support ongoing for a patient admitted in for a high-risk surgery/coronary artery bypass graft. after 12 days of support, the purge sidearm needed taping to stabilize and then no purge fluid/leak was identified. the pump remained on support and after 13 days, care was withdrawn and the patient expired. | [5 empty] | The investigation for the purge leak-pump has been completed. no product was returned for evaluation. clinical details noted that there was no leaking of purge fluid observed in the field. perfusionist noticed sidearm was "wiggling" and was taped for stabilization. additionally, no changes in purge values observed in the field. as the customer reported that no purge leak was observed in the field and only wiggling of the purge sidearm was noted. the root cause of the cosmetic defect cannot be definitively determined as no product was returned. | This complaint, purge fluid sidearm lack of solidity, was inadvertently reported under the medical device reporting regulations and initial report (submitted 20-dec-2024) and supplemental follow-up 1 report (submitted 10-feb-2025) should be disregarded. medical safety review of the event noted the patient was very sick, decompensating and dying. there is no product allegation beyond sidearm lack of solidity with use of tape for stabilization. information at hand does not state or reasonably suggests there was interruption, disruption, change or pause in therapy, and the reporter confirmed there was no purge fluid leak and no change in purge value. the complaint is without information that reasonably suggest the stabilization of the sidearm may have cause or contributed to the patient's demise, and further this type of complaint is not one that is likely to cause or contributed to a serious injury or illness if it were to recur. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
12785 | 1220648-2024-25031 | 20989748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989748&pc=OZD | 12/09/2024 | 01/14/2025 |  | 12/20/2024 |  |  | 12/09/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. the sheath was placed, and an attempt was made to cross iliac artery multiple times with no success. the decision was made to shoot the left femoral to see if it could go contralateral but the left side was even worse. the case was aborted due to vascular anatomy. the patient¿s outcome is stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿.
12786 | 1220648-2024-25032 | 20989844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989844&pc=OZD | 06/25/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 06/25/2023 | 12/20/2024 |  |  | 03/31/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there were noted positioning issues. the patient was switched to venoarterial membrane oxygenation. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. data confirmed about 1.5 hours since the pump started pump was in aorta area. pump level then was down to p-0 to the rest of the case. product and data were not returned for review. based on clinical description, the root cause of positioning issue was due to use issue. this report is being filed as part of a retrospective review of historical records.
12787 | 1220648-2024-25033 | 20989951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20989951&pc=OZD | 11/29/2024 | 02/27/2025 |  | 12/20/2024 |  |  | 12/03/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in had a impella cp support for 7 days for a patient admitted in with cardiogenic shock and cardiomyopathy. the impella cp was weaned and explanted. the team then powered the automated impella controller down but the powering down failed. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the automated impella controller (aic) shutdown issue has been completed. the console logs from the reported event date aligned with the complaint of inability to shutdown. the pump was disconnected (twice) and the user then requested a shutdown of the console. however, the shutdown sequence failed due to unavailable resources. the logs indicated a real time logger error message. subsequently, the pump was reconnected to the aic, then disconnected twice, but the aic remained unresponsive. the aic attempted to shut down, displaying a black screen, but was unable to complete the shutdown process. further investigation revealed that the rlm was frequently rebooting, which likely caused the aic to repeatedly attempt to restore the rlm data streaming connection, resulting in multiple copies of the usb driver process being created. this created a aic software resource shortage. the rlm logs indicated potential issues with the microsd card. the console was tested, but the issue could not be reproduced. additionally, no power supply interruptions were detected between the power battery manager and the 4-in-1 assembly. the issue with the aic not shutting down correctly was identified as a software anomaly by software subject matter experts. the root cause of the aic's inability to shut down was due to an aic software anomaly. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25033. d.9 revised date device returned to manufacturer as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25033. g.1 revised reporting contact email as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25033. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25033.
12788 | 1220648-2024-25034 | 20990231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20990231&pc=OZD | 12/04/2024 | 04/15/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569406 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received from the clinical site confirmed that the patient did not expire while on support. it was reported that the patient was successfully weaned and survived the event. | The complainant reported a patient on impella 5.5 support and that there was a pump stop. when the p level was increased, it caused a spike on the motor current and the automated impella controller (aic) started alarming again. the recommend was given to first to replace aic, however, after the aic was replaced the issue remained. the recommendation given afterwards was to replace the pump. however, care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. |  |  |  |  | The investigation for the pump stop has been completed. impella 5.5 was returned for evaluation. upon visual inspection, biomaterial was present on the impeller blade. pump was electrical tested and all three motor cable lines were measured as open lines. the red handle was opened and ground line of motor cable lines was found to be broken on the catheter side of the red handle. additionally, green motor cable line was found to be necked and bent at monkey fist glue joint. no data logs were returned for evaluation. clinical details noted that a spike in motor current was observed in the field with an "impella defective" alarm occurring. the root cause of the pump stop was most likely due to an electrical open. added d9 device return date. | B.1 revised since previous manufacturer device reports 1220648-2024-25034 were submitted due to updated information received regarding the patients outcome. b.2 revised since previous manufacturer device reports 1220648-2024-25034 were submitted due to updated information received regarding the patients outcome. b.5 revised from the initial manufacturer device report 1220648-2024-25034 to reflect updated information received regarding the patient's outcome. g.1 revised reporting contact email as it was entered incorrectly on the initial manufacturer device report 1220648-2024-25034. g.2 added foreign as it was inadvertently omitted from initial manufacturer device report 1220648-2024-25034. h.1 revised since previous manufacturer device reports 1220648-2024-25034 were submitted due to updated information received regarding the patient's outcome. h.6 health effect - clinical code 4580 and health effect - impact code 1802 were revised since the initial manufacturer device report 1220648-2024-25034 was submitted to reflect updated information received regarding the patient's outcome. codes 4121, 3243, 4251, 50 and 4307 were reported incorrectly on follow up 1 manufacturer device report number 1220648-2024-25034. new codes were added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12789 | 1220648-2024-25035 | 20990483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20990483&pc=OZD | 12/04/2024 | 01/14/2025 |  | 12/20/2024 |  |  | 12/04/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849082 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during that the purge cassette would not plug into the automated impella controller (aic). a piece of plastic/debris was found in the purge cassette holder and once it was removed, the purge cassette went in like normal. it is unknown if the aic was replaced due to this issue. further details are unknown. | [5 empty] | The investigation for the console (aic) purge disc/flag issues has been completed. console logs were reviewed which did not contain any data relevant to this complaint. the console was not returned for evaluation due to the issue being resolved in the field. a piece of plastic was found in the purge cassette holder and once it was removed, the purge cassette worked as normal corrections to the initial mfr report: d2 common device name updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12790 | 1220648-2024-25036 | 20990543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20990543&pc=OZD | 12/04/2024 | 01/29/2025 |  | 12/20/2024 |  |  | 12/04/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565612 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient planned for high-risk percutaneous coronary intervention (pci) was implanted with an impella cp for mechanical circulatory support pre-pci. the patient who had left heart catheterization the prior week was turn-down for coronary artery bypass graft surgery. treatment proceeded with impella supported pci of the left main, ostial left anterior descending, and left circumflex arteries. the right femoral artery (rfa) was accessed, and angiogram obtained with right common iliac and external iliac disease noted with some concern from physician. however, it was noted the vessel appeared angiographically large enough for impella insertion. impella cp device was advanced; however, it was unable to pass noted lesion(s) and therefore removed. the physician assisted with percutaneous angioplasty of mentioned iliac segments, and the impella was reinserted without incident. the pci was completed noted with successful results. the impella cp was weaned and pulled back from left ventricle. the physician, however, was concern with peel away sheath¿s position in the rfa. the physician inquired how to remove peel away sheath and leave impella catheter in vessel while advancing repositioning sheath into arterial tract to advance the guidewire into vessel for closure and maintain access. the remaining perclose failed post explant, and the vascular team was contacted for support in closing the rfa. a manta device placed with noted bleeding around device and covered stent placed. hemostasis was obtained and good closure mentioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be patient condition because of the calcified vessels of the patient. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25036 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4624 was reported incorrectly on the initial manufacturer device report number 1220648-2024-25036 and a new code was added.
12791 | 1220648-2024-25037 | 20990630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20990630&pc=OZD | 12/04/2024 | 01/29/2025 |  | 12/20/2024 |  |  | 12/04/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548651 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed a hematoma, which was noted upon arrival to the unit, and would eventually expire. the patient underwent left heart catheterization (lhc), and percutaneous coronary intervention (pci) was performed to the circumflex artery (date/timeframe unknown). while in recovery, the patient presented with unstable angina prompting a code heart. the patient experienced ventricular tachycardia prompting defibrillation and intubation. it was discovered from the repeat lhc the patient had 100% re-occlusion of the circumflex artery. pci angioplasty and stenting were performed via right femoral artery. the patient was hemodynamically unstable, and impella cp mechanical circulatory support (mcs) post pci was initiated via right femoral artery. after the impella implant, upon arrival to the intensive care unit, a hematoma was noted at the access site. reassessment of angle of entry initiated with femostop support. critical lab value results noted low hemoglobin at 6 g/dl. the patient received blood products (specifics unknown). however, it was noted the family declined escalation of care, and the patient would subsequently expire approximately 17 hours after start of impella mcs. | [5 empty] | The investigation of access site bleeding and vt has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-25037. b2; life-threatening should have not been selected. e4 should have been left blank. h6 health effect - clinical code 2095 was inadvertently omitted. h6 health effect - impact codes 4641 and 4647 was inadvertently omitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
12792 | 1220648-2024-25038 | 20990781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20990781&pc=OZD | 12/05/2024 | 01/27/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571048 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the little bit of oozing from groin side. as a result, it was decided to implement heparin at a later time when activated clotting time (act) is in target range. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25038 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12793 | 1220648-2024-25039 | 20991216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20991216&pc=OZD | 12/05/2024 | 03/10/2025 |  | 12/20/2024 |  |  | 12/05/2024 | 12/20/2024 |  |  |  |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526547 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp procedure when removing the guidewire, the guidewire had a full fracture. the patient¿s vessel was tortuous. there was no harm to the patient. it was also reported the 14f introducer was kinking after the removal of the dilator. the physician switched over to the left groin and a second introducer was used. the second introducer was also kinking as well but impella was able to be removed after the percutaneous coronary intervention procedure with no harm to the patient. | [5 empty] | The investigation has been completed. the introducer and guidewire were returned for investigation. both the returned introducer and guidewire have kinks on them, with no other damage noted. additionally, the returned guidewire was stretched from the tip coil and broke into two pieces. according to the provided clinical information, the initial introducer was kinked after the removal of the dilator. the patient was reported to have a tortuous vessel condition. the operator switched sides, and the same issue occurred with the new introducer. however, the impella explant was successful without patient harm. furthermore, a full fracture of the guidewire occurred during the withdrawal, and the floppy end was torn off between the valves of the introducer, while the noted repository unit was slightly inserted into the introducer. the failure mode for the delivery issue was changed to product damage (guidewire damage - peripheral vessels) since a full fracture was noted on the returned guidewire. the introducer damage issue that occurred during the delivery was addressed under a separate failure mode. the cause of the guidewire issue was most likely use related since the floppy end was torn off between the valves of the introducer, while the noted repository unit was slightly inserted into it. the cause of the first introducer issue was patient condition related tortuous vessel. the cause of the second introducer issue was patient condition related tortuous vessel. d.9 device return date/information was added. g.1 revised reporting contact email as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-25039. revised manufacturing site country as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25039. g.6 30-day selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-25039. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12794 | 1220648-2024-25040 | 20991829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20991829&pc=OZD | 03/15/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 03/15/2023 | 12/20/2024 |  |  | 07/31/2024 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253034 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. echo was called to confirm placement and positioning issues were noted. multiple attempts were made to rotate the pump away from the mitral valve into free space without success. the device was explanted due to the placement issues. no patient harm was reported. | [5 empty] | The investigation into the reported issue was completed. no product was returned for investigation. per the provided clinical information, the root cause of the positioning issue was most likely use related.
12795 | 1220648-2024-25041 | 20992403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20992403&pc=OZD | 06/19/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 06/19/2023 | 12/20/2024 |  |  | 02/28/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2024331669 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a delivery issue involving an impella cp pump. during the initial attempt, the pump failed to cross an existing lesion in the external iliac. the pump was withdrawn and re-inserted for a second attempt, however the sheath kinked at the iliac lesion and the pump once again failed to advance. after a subsequent third failed attempt to deliver the pump, the implant was aborted and the decision was made to reschedule the patient for axillary pump insertion. there was no patient harm. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. per the given clinical details, the root cause of the delivery issue was patient condition related to small/ diseased vessels. per the given clinical details, the root cause of the introducer kink issue was patient condition related to small/ diseased vessels. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
12796 | 1220648-2024-25042 | 20992734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20992734&pc=OZD | 06/22/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 06/22/2023 | 12/20/2024 |  |  | 01/31/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316609 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge tubing kinked during patient transport prompting replacement. it was also noted that the patient developed an access site bleed, however intervention was not required and support with the implanted pump continued. | [5 empty] | Note: the explant date is unknown. product was not returned for investigation. data logs were returned and reviewed. the root cause of the high purge pressure issue was most likely a kinked purge line, as the purge pressure rose rapidly and flows declined. the high purge pressure resolved after the cassette was changed. per the clinical information provided, the root cause of the access site bleeding issue was most likely patient movement during patient transportation. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12797 | 1220648-2024-25043 | 20993113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20993113&pc=OZD | 07/12/2023 | 12/20/2024 |  | 12/20/2024 |  |  | 07/12/2023 | 12/20/2024 |  |  | 04/30/2025 |  | 12/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348506 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella leg had lower than expected perfusion following pump explant. it was noted that the site had been closed using 2x angioseal following pump removal. the staff continued to monitor the low perfusion and further intervention was not deemed necessary. | [5 empty] | The investigation into the reported issue was completed. no device was received for evaluation; however, clinical information was sufficient to determine the root cause. the cause of the reported issue was most likely patient condition related. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
12798 | 1220648-2024-24943 | 20975809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20975809&pc=OZD | 12/03/2024 | 01/30/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564648 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient underwent a cerebral infarction. after implant of the pump, anisocoria was confirmed in the intensive care unit and cerebral infarction was developed. it is unknown whether the cause of the cerebral infarction was due to patient factors or impella. additionally, after the cerebral infarction, the patient passed away. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12799 | 1220648-2024-24944 | 20975961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20975961&pc=OZD | 12/04/2024 | 01/22/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556135 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a cannula kink. after successfully placing the impella cp during active cardiopulmonary resuscitation, the impella cannula could be seen with a kink along with low flows. several unsuccessful attempts were made to reposition the pump. ultimately, the doctor decided to remove the pump and replace it with a new impella cp. the procedure outcome was successful with the other device. | [5 empty] | The investigation for the product damage has been completed. the pump was returned for evaluation. upon visual inspection, a cannula kink was observed near outlet. no biomaterial observed. no broken blade found. data logs were reviewed and found suctions occurred. drop in flow observed after impella insertion which confirms cannula kink that obstructed the flow. the root cause of the low pump flow issue was cannula kink due to improper technique. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12800 | 1220648-2024-24945 | 20975970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20975970&pc=OZD | 06/15/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/15/2023 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1407963 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella console was defective and presented optical signal issues. no patient was involved. | [5 empty] | The investigation into the reported issue has been completed. log showed a couple of issues with the console controller error. the following 5s parameters are abnormal; specifically, the "light¿ parameter triggered the controller error. the console was tested. the failure mode was not reproduced upon testing. the root cause of the optical signal issue is not determined due to the inability to reproduce.
12801 | 1220648-2024-24946 | 20976127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976127&pc=PYX | 12/04/2024 | 01/03/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517708 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding. during implant, after the peel away was removed, the repositioning sheath was advanced and the mattress suture was tightened down. there was still bleeding. despite placing an additional mattress suture, there was still bleeding. therefore, manual pressure was held and the doctor ultimately decided to also place a low pressure femostop and the site was stabilized. three days later, the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | H.1 was reported incorrectly on manufacturer device report 1220648-2024-24946. h.6 was reported incorrectly on manufacturer device report 1220648-2024-24946. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
12802 | 1220648-2024-24947 | 20976136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976136&pc=OZD | 06/19/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/19/2023 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1326578 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that in preparation for a software update the aic was booted up and displayed a battery error. console was not charged when booted first. the error persisted after two full days of charging. no patient was involved. | [5 empty] | The investigation into the reported issue has been completed. logs confirmed the reported failure. there was a battery failure alarm, battery level low alarm and battery failure (low voltage) [v < 12v] - alarm due to cell imbalance. though there is no ltpowerdata from the complaint date, batttest confirms the cell imbalance with cell 1 of battery 1, which caused the battery failure. the console was tested. the failure mode was reproduced while running the aic without the ac power. also confirmed the battery failure by performing batttest in the terminal window and by visual inspection of the battery lcd screen. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to cell imbalance, a known pattern failure. before sending this console back to the customer, field service was instructed to replace the battery kit, validate charging, and perform a full functional check on the console and optical system.
12803 | 1220648-2024-24948 | 20976209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976209&pc=OZD | 11/27/2024 | 01/29/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549391 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after 22 days of impella 5.5 support the purge flow drop until less than 1ml/hr. doctors checked the alpha tocopherol transfer protein, and it was 60. the echocardiogram was good, impella was in good position. the impella was removed. the md decided to put a centrimag because the patient¿s anatomy didn¿t allow for another impella to be placed. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the high purge pressure has been completed. the pump was not returned for investigation. data logs were not returned for review. as per the clinical notes, purge flow dropped to one milliter per hour on the 22nd day of use. the cause of the high purge pressure issue was not determined since the product was not returned and sufficient clinical information was not provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24948 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact emial since the initial manufacturer device report number 1220648-2024-24948 was submitted in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12804 | 1220648-2024-24949 | 20976220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976220&pc=OZD | 12/03/2024 | 01/21/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528939 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to having cardiogenic shock. during support, there were low pump flow alarms and echocardiography showed a left ventricle thrombus. the pump was explanted and an intra-aortic balloon pump (iabp) was placed. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system: section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the low pump flow and thrombus has been completed. the pump was not returned for analysis. data logs were reviewed which showed that suction alarms occurred. motor current (mc) spikes were observed at the same time of first suction followed drop in p-level. flows were within specific limits with p-levels change, but at p-1 increase in flow observed and couple mc spikes observed. a few position alarms in aorta occurred that resolved with quick trouble shooting. a purge spike observed while pump run on p-1. low pump flow changed to saturated flow based on data logs observation. the root cause of the saturated flow issue most likely was biomaterial ingestion (thrombus).
12805 | 1220648-2024-24950 | 20976228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976228&pc=OZD | 12/04/2024 | 01/13/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527395 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed oozing from the impella access site. purse string suture was placed and femstop. the patient is stable. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since hemostasis achieved by adding purse string suture. e4 should have been left blank on manufacturer device report 1220648-2024-24950. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12806 | 1220648-2024-24951 | 20976275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976275&pc=OZD | 12/05/2024 | 01/13/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the groin site oozing and a small hematoma. the oozing groin site oozing was treated by placing another forward tension suture and redressing the site. the small hematoma was treated by applying manual pressure and adding stitches. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since adding additional stitch resolved the bleed. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12807 | 1220648-2024-24952 | 20976279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976279&pc=OZD | 06/21/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/21/2023 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135089 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was an aic battery failure, and that the console was not connecting to impella connect. there was no report of patient harm. | [5 empty] | The investigation for the reported battery cell failure and impella connect - connectivity/wifi issue has been completed. the product was returned, and the issue was confirmed. data analysis: rlm- the console was returned to the field on may 5th. "bswhbiomed" was configured and the console was able to connect and establish data connection and was able to maintain connectivity. aic- there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage (imc ic2438.pdf). additionally, the lt logs showed that battery 2 had a cell imbalance on cell 4. device analysis: rlm - ictest was able to be connected to and have wifi connected and maintained. the engineer who aided in wifi configuration at the hospital confirmed that implementation was successful. aic - the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 0 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). battest confirmed battery 2 cell 4 was outputting 0v. per the results of the clinical review and investigation, the root cause of the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
12808 | 1220648-2024-24953 | 20976295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976295&pc=OZD | 12/04/2024 | 01/09/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576452 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered an delivery issue. the patient's iliac lesion would not let the pump pass with a short 14 french sheath so the pump was pulled out and it was found that there was a thrombus in the inlet. therefore, a new pump with a long 14 french sheath was used instead for the procedure. the patient survived the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to the patient's anatomy. the root cause of the thrombus could not be determined without additional clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12809 | 1220648-2024-24954 | 20976497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976497&pc=OZD | 09/28/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 09/28/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1317758 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller unit had loose screws within the device. there was no patient involvement at the time of the noted issue. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. inspection of the console revealed loose screws inside the aic housing, which were determined to have been the two bottom mounting screws for the impellatronic board. in addition, the aic batteries were from the lot known to be susceptible to failures (single cell failure). no other deficiencies were observed during testing in engineering. it is most likely that the screws were not tightened after the board was reinstalled. the root cause of the aic issue was a loose fastener.
12810 | 1220648-2024-24955 | 20976541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976541&pc=PYX | 09/13/2024 | 01/15/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025543816 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella rp flex device used. acute kidney injury required continuous renal replacement therapy was noted. no history of kidney problems prior to admission were reported. the patient/event has not recovered/not resolved and the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the renal failure has been completed. no product was returned. log review did not show any motor current spikes indicating biomaterial ingestion or suction. there were also no placement signal trends indicating deteriorating patient condition. the root cause of the renal failure was not determined since product was not returned and insufficient clinical detail was provided.
12811 | 1220648-2024-24956 | 20976625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976625&pc=OZD | 09/25/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 09/27/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154174 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the console show an alarm "unexpected console shutdown". the controller was turned off, then controller turned on and runs without problems. no further alarms. since then, no impella connect connection. no harm was done to the patient and the product was returned for investigation. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. the data logs confirmed the reported failure mode. the console unexpected reboot when running an impella cp on ac power. there were no abnormal log entries before or after the unexpected reboot. the console was able to resume patient support (previous p-level) within 2 minutes 30 seconds after one hard reboot. device analysis: the console was tested in collaboration with field service in (b)(6), but the failure mode was not reproduced. the power battery manager circuit board¿s efficient power conversion supply was within normal range for battery and ac mains operation. console was run for 24 hours with known-good pump and purge, but no unexpected reboot was reproduced. the root cause of the unexpected reboot was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
12812 | 1220648-2024-24957 | 20976933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20976933&pc=OZD | 09/14/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 09/14/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an placed on impella 5.5 and an automated impella controller (aic) for mechanical circulatory support. while on support, the inability to download usb data from aic was noted. the aic would not recognize the usb device sent with the ra kit. the aic was replaced without any harm or injury. | [5 empty] | The investigation for the reported controller error (yellow alarm) and shutdown or restart issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs could not confirm the reported issue on the complaint occurrence date ((b)(6) 2022). six days after, there were multiple attempts to download the logs. more than once, the download failed for an unknown reason: logs indicate that these initial attempts to download logs could have been made at the i/o usb port instead of at the rlm¿s usb port given the inability for the aic to detect the rlm beforehand via device handhsake. a successful usb data download was eventually performed (see (b)(6) logs usb data download successful). unreported but potentially related to the usb data download difficulties observed by the user, the console had underwent multiple partial and one full reset due to a failure of the imcgui process. there was also a controller error (loss of kbd communication ¿ event set #24) on one of the first boot ups. these events were not related to a clinical procedure. device analysis: console (b)(6) was investigated in collaboration with (b)(6). the issues downloading logs was reproduced but linked to a different failure mechanism. two attempts were made to download logs via the rlm¿s usb port, and each stopped at 21%, presenting an error while downloading. there was a long duration case which was almost 130 hours on the console¿s drive and caused the usb file transfer to take longer than 90 seconds. with this duration of aic-rlm data streaming-related usb inactivity, the rlm will usb hub will reset by design, causing the usb data download to fail. the kbd-related controller error was due most likely due to an intermittent communication issue during startup (likely related to timing). however, it was not reproduced and not determined. the repeated console reboots (partial and full) were due to a failure of the imcgui process. the reason for the imcgui process failure was not determined. the sw mitigation was successful in restarting the process after the multiple partial and one full reset. per the results of the clinical review and investigation, the root cause was determined to be for the usb data downloading issues in the field were due to multiple console reboots, which were caused by a failure of the imcgui process. the cause of the controller error was not reproduced and not determined. this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records.
12813 | 1220648-2024-24958 | 20977316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977316&pc=OZD | 08/04/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 08/04/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301100 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the aic was unable to get accurate readings/waveforms. there was no patient harm reported. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported day of event are inconsistent with complaint as there were no entries after (b)(6) 2022. analysis of the logs prior to that date suggested that the reported issue was presumably a duplicate of previously reported issue. testing of the console, that was conducted and did not reveal any deficiencies that would be related to placement signal or motor current. nevertheless, console's inspection revealed some irregularities: cracked blue retainer (purge pressure disk retainer), and purge pressure sensor out of spec (fails 1650mmhg test point). the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
12814 | 1220648-2024-24959 | 20977345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977345&pc=OZD | 12/03/2024 | 01/14/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1429248 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) would not advance beyond the priming screen. the aic was replaced to resolve the issue. there was no patient harm. | [5 empty] | The investigation for the priming issue has been completed. console logs from the reported day of event were consistent with complaint and indicated that the console did not detect the pump being connected to it, after the case start wizard was initiated manually. the console kept priming for about five minutes, and did not proceed to step 4 of case start wizard. there were no attempts to re-start case start wizard. the console (aic) was returned for evaluation and was found to operate within specification, and detected a pump each time. what was described as "priming screen did not progress" was in fact a failure to detect pump connected to the console. it was most likely caused by ¿epm crystal issue¿ where a sub-circuitry of impellatronic board occasionally malfunctions. the impellatronic pca does not have any permanent damage, and this intermittent pump detection failure occurred infrequently and already had a software mitigation in place, although not resolving the issue 100% of the time - which might have been the case here. a.5 race was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24959 and was added. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12815 | 1220648-2024-24960 | 20977668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977668&pc=OZD | 12/05/2024 | 01/23/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501405 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was a spike in purge pressure and there was saturation. the impella cp stopped. the pump was powered down and pulled back into the descending aorta. the impella cp was removed and there was no harm. | [5 empty] | The investigation high purge pressure and pump stop has been completed. no product was returned. the data logs show that there is a motor current spike 12 hours into support followed by an immediate purge pressure rise and saturated flows. the purge pressure rises lasts for over 8 hours while purge flows slowly decrease until purge alarms are triggered. a pump stop is not seen in the logs. the root cause of the high purge pressure is most likely due to ingested biomaterial. the root cause of the pump stop cannot be determined as there was no pump stop seen in the data log analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12816 | 1220648-2024-24961 | 20977785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977785&pc=OZD | 10/23/2024 | 12/19/2024 |  | 12/19/2024 |  |  | 10/23/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1367614 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal failure during support. the issue resolved within 10 minutes. there was no harm done to the patient. while investigation the automated impella controller (aic) opm failure was noted. | [5 empty] | Date device returned to manufacturer (field d9) is unknown. the investigation for the reported optical signal issue has been completed. the product was returned, and the optical signal issue was not confirmed. log analysis: this is a known pattern failure in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. this lasts between 2 to 10 minutes, after which all signals go back to normal. device analysis: complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause of the transient loss of optical data is not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
12817 | 1220648-2024-24963 | 20977953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977953&pc=PYX | 12/06/2024 | 01/23/2025 |  | 12/19/2024 |  |  | 12/06/2024 | 12/19/2024 |  |  | 08/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395761 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella rp was inserted but in the first attempt the .027 guidewire was kinked, and no replacement was available. two .018 guidewires were used to deliver the impella rp with ease. a mattress suture was used for good femoral hemostasis. it was also reported that on day two of impella rp support, the impella rp was malpositioned and the impella rp was explanted. the patient is stable and survived the event. | [5 empty] | The investigation for the delivery issues and the positioning issues has been completed. no product was returned. as per the clinical information provided, the guidewire was kinked upon insertion and later 0.18 guidewire were used successfully. therefore, the root cause of the delivery issue was kinked guidewire due to excessive force. clinical also mentions that the pump was malpositioned while transferring the patient from another hospital. therefore, the root cause of the positioning issue was patient movement. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
12818 | 1220648-2024-24962 | 20977963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20977963&pc=OZD | 06/12/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/12/2023 | 12/19/2024 |  |  | 01/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314685 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, high friction was observed while positioning the 0.018" guidewire, and the physician could not insert the guidewire. after removing the easy guide lumen, the physician was concerned about potential portion of the lumen in the catheter. another impella cp was successfully used and the patient was stable. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the red lumen was not returned for review. visual inspection found no issues on the pump. a red lumen was installed to the pump, then gw was inserted and went through the red lumen without issue. the root cause of delivery issue was unable to be determined as the red lumen was not returned for review. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
12819 | 1220648-2024-24964 | 20978055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978055&pc=OZD | 06/14/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/14/2023 | 12/19/2024 |  |  | 11/24/2024 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274585 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped. after the explant, no clot was found. no further information was provided. no patient injury or harm was reported. | [5 empty] | Date/time explant: unknown. the investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no data logs were provided. no physical abnormalities were observed, and the pump ran as expected during the test. the root cause of the pump stop issue was not determined since data logs and sufficient clinical information were not provided. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
12820 | 1220648-2024-24965 | 20978186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978186&pc=OZD | 12/06/2024 | 01/09/2025 |  | 12/19/2024 |  |  | 12/06/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576527 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had hemolysis. the patient's urine was dark red and labs were hemolyzed. the patient was escalated to impella 5.5 support. the patient survived the event. | [5 empty] | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were positioning and suction alarms triggered during the case but resolved. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12821 | 1220648-2024-24966 | 20978300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978300&pc=OZD | 12/06/2024 | 01/29/2025 |  | 12/19/2024 |  |  | 12/06/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551899 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was inserted and percutaneous coronary intervention was completed. post-procedure, the peel-away sheath was removed and the repositioning sheath was advanced. suction alarms were present with low flows. the waveforms were decoupled. the impella position was optimal under fluoroscopy. the physician decided to remove the impella. there was no harm to the patient. | [5 empty] | The investigation for the low pump flow has been completed. the impella cp was returned for evalulation. upon visual inspection, no damage to the pump was observed. no biomaterial was present on the pump. pump was run in a bucket of water for approximately 8 minutes and no low flow was observed. rt log shows slightly saturated pump flow and motor current from start of case. suction alarm at time of drop in pump flows noted. rt log at time of first suction alarms shows drop in flows with slight saturation after motor current decreases. spikes in placement signal also observed right before flows decrease. clinical details noted that pump was on p-auto at the time of suction alarms and device position was checked and confirmed to be correct. the root cause of the low pump flow cannot be definitively determined as the low flow was not reproduced in the lab and no abnormalities with the pump were observed. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12822 | 1220648-2024-24967 | 20978421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978421&pc=OZD | 06/20/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/20/2023 | 12/19/2024 |  |  | 03/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342443 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient on impella support encountered an alarm indicating that the pump was incorrectly positioned in the aorta. the pump was replaced with the new one. there was no harm to the patient as a result of this incident. | [5 empty] | Product was not returned for investigation. according to the clinical data, the initial echo indicated proper positioning and waveform consistency. subsequent alarms, suggesting a potential shift to the aorta, prompted adjustments in impella settings. despite efforts to clarify the situation and plan for further evaluation, the doctor decided to shut down the device, necessitating its replacement. data logs were reviewed and no related alarms were found per reported complaint outcome. data logs did not match the complaint description. there was no issue found during the case, based on data log review.
12823 | 1220648-2024-24968 | 20978491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978491&pc=OZD | 12/05/2024 | 01/13/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568643 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, a computed tomography (ct) scan showed the patient had an ischemic stroke. the patient was taken from the ct to the operating room for a craniotomy. suction alarms occurred on the impella 5.5. p level dropped and they were unable to give volume due to a risk of brain swelling. the patient is stable. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12824 | 1220648-2024-24969 | 20978709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978709&pc=OZD | 07/27/2021 | 12/19/2024 |  | 12/19/2024 |  |  | 07/27/2021 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1315815 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with loss of placement signal. the team replaced the aic and support proceeded. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported day of event are inconsistent with complaint description, and the reported event could not be identified in logs. prior to the day of reported event, console was used on 2021-06-11, with pump 249405, for about 1 hour, with no loss of placement signal. next use was on 2021-07-26, but no pump was recorded. further analysis of logs indicated that pump was not recognized. the manufacturer followed up with the representative for clarification of the nature of the reported event, but received no response - therefore we assumed that "loss of placement signal" was entered in error, and focused on investigating pump not being recognized during case start. logs show that the console was able to detect pump connection but failed to read pump eeprom. this is a new failure mode only found in aic_v8.4 and aic_v8.4.1 and has been documented. in the logs, it was noted that during console boot-up the optical bench state machine (ossiopsens) gets stuck in post_start_state and does not reach post_ok_state. this prevents the pump from being read. the root cause of the reported issue was found to be bugs or defects in software (excludes firmware). a mitigation for the issue has been added to aic_v8.5. this report is being filed as part of a retrospective review of historical records.
12825 | 1220648-2024-24970 | 20978747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978747&pc=OZD | 11/07/2024 | 12/19/2024 |  | 12/19/2024 |  |  | 11/07/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) crashed indicating risk of power/case start issues during training. there was no report of patient involvement. | [5 empty] | (d9) device returned to manufacture date is unknown. the investigation for the reported "aic - shutdown or restart issue" has been completed. the product was returned and the "aic - shutdown or restart issue" was not confirmed. the failure mode was reproduced on an engineering lab bench. despite having the correct supply voltage and power on signal, the 4-in-1 would not power on. momentarily swapping this component with a known-good resolved the boot up problem, and console ran over a long period of time without any issues. the root cause of the boot-up issues was due to a defective 4-in-1. this report is being filed as part of a retrospective review of historical records.
12826 | 1220648-2024-24971 | 20978920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978920&pc=OZD | 12/03/2024 | 01/16/2025 |  | 12/19/2024 |  |  | 12/03/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546813 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support for around three weeks when purge pressure was high and purge system blocked alarms occurred. tissue plasminogen activator protocol was completed twice without success. the impella 5.5 was exchanged for another impella 5.5. there was no harm to the patient. | [5 empty] | The investigation of high purge pressure has been completed. the returned product was inspected and blood was found in the outlet cage by the impeller. there were no physical abnormalities. the pump was run in the lab and high purge pressure was not reproduced. the data logs confirm high purge pressure alarms. it is unknown if tissue plasminogen activator (tpa) was added to the purge. the root cause of the high purge pressure was determined to be biomaterial instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2024-24971. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12827 | 1220648-2024-24972 | 20978961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20978961&pc=OZD | 09/03/2021 | 12/19/2024 |  | 12/19/2024 |  |  | 09/03/2021 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409295 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Staff reported that during support, the automated impella controller (aic) had placement signal values that were not showing. additionally, it was alarming for a controller error. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. the device was received for evaluation. logs are consistent with complaint date. optical 5s parameters were checked with a optical test cavity and observed to be within tolerance. blue connector optical ferrule was inspected and debris and slight damage were found. console was ran on pump and purge with no issue. console interior was inspected where it was observed that the blue fiber directly from optical splitter was routed poorly. the root cause of the absence of placement signal values during support is most likely due to an optical bench malfunction however the malfunction cause is undetermined due to being unable to reproduce. this report is being filed as part of a retrospective review of historical records.
12828 | 1220648-2024-24973 | 20979160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20979160&pc=OZD | 11/05/2024 | 12/19/2024 |  | 12/19/2024 |  |  | 11/06/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1222703 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, automated impella controller (aic) had optical and controller failure alarms. no additional information was provided. | [5 empty] | Date device return to manufacturer is unknown. the investigation for the reported optical signal issue and aic - controller error (yellow alarm) issue has been completed. aic logs confirmed the reported issue. there were multiple instances during the case start procedure where the console presented a screen related to an optical sensor system error. this was triggered due to an invalid/ unreliable placement signal. rt logs revealed definite placement signal issues that were eventually resolved by reconnecting the pump as advised from the user interface. during routine testing, the console alarmed controller error (opm comm loss). this loss of communication during the routine testing was likely unrelated to the placement signal issues from the complaint date. during the electrical safety test, the pump cable (associated with the fixture) was plugged in, which usually results in an impella defective alarm, but no alarm appeared per usual. this was likely an intermittent malfunction of the epm circuitry and was not related to the placement signal issues from the complaint date. the root cause of the placement signal issue in the field was most likely an air gap (or intermittent optical connection within optical system). this report is being filed as part of a retrospective review of historical records.
12829 | 1220648-2024-24974 | 20979420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20979420&pc=OZD | 05/01/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 05/01/2023 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210717 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Operating System Version or Upgrade Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in on an automated impella controller (aic) for mechanical circulatory support. while on support, after changing the purge cassette, the impella controller alarm displayed purge cassette or disk not detected. controller and purge cassette exchanged to a purge cassette with a different lot number and the problem resolved. there was no report of patient harm or injury. | [5 empty] | Manufacture date: unk. the impella automated controller device was discarded by the customer; therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12830 | 1220648-2024-24975 | 20979533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20979533&pc=OZD | 06/23/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/23/2023 | 12/19/2024 |  |  | 12/31/2024 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303397 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported there was blood coming from the red impella plug and the purge pressure was low. a leak was observed, and the purge cassette was replaced. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12831 | 1220648-2024-24976 | 20979720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20979720&pc=OZD | 05/25/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 05/25/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171605 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) console failed to display optical signals. there was no patient harm reported. while investigating the console, a yellow controller alarm was also observed. | [5 empty] | The investigation for the reported "aic - controller error (yellow alarm)" has been completed. the product was returned, and the reported observation was confirmed the failure mode of ps not reliable was reproduced with test pump and pc. receptacle and optical fibers interconnections were inspected but no abnormality was found, 5s parameters are within specification. optical bench was temporarily replaced with a known-good one for troubleshooting, which resolved the ps alarm. console ic3418 was confirmed for placement signal not reliable alarm due to defective optical bench. the exact root cause of optical bench malfunction was not determined. this report is being filed as part of a retrospective review of historical records.
12832 | 1220648-2024-24977 | 20979833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20979833&pc=OZD | 01/12/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 01/12/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1361559 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | Altered State of Consciousness | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the icm does not start when the console is switched on. there was no mention of patient involvement. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported display issue has been completed. the product was returned, and the issue was confirmed. device analysis: failure mode was reproduced during investigation. aic booted up, and it was observed that the icm did not start. usb-connector on the io-panel was unplugged and plugged several times without success. while the micro-usb plug on the icm was pressed, the power was momentarily restored and light mode and blue led would eventually flash, indicating the cable movement/ reducing strain resolved the intermittent backstrap cable issue. a known-good icm module was replaced temporarily without solving the failure. per the results of the clinical review and investigation, the root cause was determined to be related to the "backstrap" cable on the back of aic this report is being filed as part of a retrospective review of historical records.
12833 | 1220648-2024-24978 | 20980682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20980682&pc=OZD | 06/10/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 10/15/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239161 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) optical bench error. no report of any patient harmed, or treatment impacted. in the aic investigation the team observed an opm failure . | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was not confirmed. the cause of the issue was not determined since the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
12834 | 1220648-2024-24987 | 20981085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981085&pc=OZD | 08/02/2024 | 12/19/2024 |  | 12/19/2024 |  |  | 08/02/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143659 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported absence of audible alarms on automated impella controller (aic) during an in-service. there was no patient involved. during investigation, team observed a lack of buzzer/ alarm sound. | [5 empty] | The investigation for the reported "aic - buzzer/ alarm inaudible" has been completed. the product was returned and the "aic - buzzer/ alarm inaudible" was not confirmed. imc logs were reviewed and deemed irrelevant for this failure mode. console was plugged into ac power and battery switch was engaged; console was then powered on. while powered on, the battery switch was then disengaged intentionally to create an alarm. when the battery failure alarm triggered, the red alarm banner was present on screen however no audible tone could be heard. console interior connections were checked and verified. it was observed that the speaker cable was disconnected. after connecting the speaker cable, the same test was run. the alarm tone was audible after reconnecting the speaker wire. the cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
12835 | 1220648-2024-24980 | 20981359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981359&pc=OZD | 10/02/2024 | 12/19/2024 |  | 12/19/2024 |  |  | 10/02/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 120612 |  | 0042-0000-UK |  | Yes |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During investigation of an automated impella controller (aic), a yellow controller failure alarm was noted. there was no patient involvement. | [5 empty] | (d9) device returned to manufacturer date is unknown. the investigation for the reported "aic - controller error (yellow alarm)" has been completed. the product was returned and the "aic - controller error (yellow alarm)" was not confirmed. the cause of the aic issue was a defective component (ribbon cable from kbd). this report is being filed as part of a retrospective review of historical records.
12836 | 1220648-2024-24979 | 20981366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981366&pc=OZD | 06/18/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/20/2023 | 12/19/2024 |  |  | 03/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326090 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. there was noted purge leakage at the connection part between catheter plug and shaft. when the surgical mode was attempted to be used in the operation, the connection part between the catheter plug and the shaft was sticky. the purge fluid was thought to be leaking. however, replacement was necessary for the impella defect. there was no patient harm reported. | [5 empty] | The investigation into the reported issue has been completed. the pump was not returned for investigation. the cause of the purge leak issue was not determined since product was not returned and sufficient clinical information was not provided. this report is being filed as part of a retrospective review of historical records.
12837 | 1220648-2024-24988 | 20981399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981399&pc=OZD | 06/02/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/02/2023 | 12/19/2024 |  |  | 02/28/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324571 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped abruptly. pump was successfully restarted but subsequently replaced with impella 5.5 pump. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned, and the pump stop was confirmed. the root cause of the pump stop was bearing wear due to misuse of the product. this report is being filed as part of a retrospective review of historical records.
12838 | 1220648-2024-24991 | 20981642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981642&pc=OZD | 02/10/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 02/10/2022 | 12/19/2024 |  |  | 02/28/2023 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049804 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the measured purge flow dropped by 50% overnight. the purge cassette and bag were subsequently replaced, and tissue plasminogen activator was added to the purge line to manage the issue. support continued with the implanted pump. additional information noted that the care was withdrawn, and patient expired. use of the device cannot conclusively be associated with the outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported high purge pressure issues has been completed. the impella device was not received from the customer. however, the data logs were sufficient to determine the root cause. the root cause of the high purge pressure issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
12839 | 1220648-2024-24992 | 20981672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981672&pc=OZD | 11/28/2024 | 04/01/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1173720 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the registered nurse connected the pump to a new automated impella controller (aic), but the new aic never asked to restore previous settings and wanted to start a new case. the nurse was instructed to plug back into the original aic; however, the original aic also did not allow to restore previous settings. the nurse had to go through the start case process again to get the new aic running. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - a/c power issues has been completed. the cause of the user not being asked to resume the pump was because the pump was read by the aic as uninitiated. it was not determined why the pump was read as uninitiated since consoles were not returned and the pump was able to be read correctly during testing. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-24992 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12840 | 1220648-2024-24990 | 20981681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981681&pc=OZD | 12/05/2024 | 02/07/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1271403 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | A complainant reported the patient was on impella cp support. the automated impella controller was returned for investigation. during review of the logs, a controller error was observed. | [5 empty] | H.2 device evaluation should have been selected on supplemental manufacturer device report 1220648-2024-24990 instead of additional information. | The investigation for the automated impella controller (aic) controller failure-red alarm has been completed. data logs were reviewed which found to have experienced a 1045 error with optical signals. this is caused by optical data misaligning and the console does not need to be returned for analysis due to the nature of this failure. the console was returned for evaluation. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24990. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24990 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised multiples fields in this section as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24990. h.6 code 4114 was reported previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24990 as the device was returned. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12841 | 1220648-2024-24989 | 20981719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981719&pc=OZD | 12/04/2024 | 01/29/2025 |  | 12/19/2024 |  |  | 12/04/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542855 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, low pump flows were observed. the pump had suction alarms and the left ventricle pressure also became negative. the position was confirmed. the pump was replaced and there was no harm to the patient. | [5 empty] | The investigation for the low pump flow has been completed. the impella device was returned for evaluation and was visually inspected. no biomaterial observed. no cannula kink observed. no broken bladed found. pump connected to aic with purging to reproduce reported low pump flow issue. no low flow reproduced. data logs reviewed finding no saturated flows seen, no motor current spike observed. several suction alarms and placement signal is noisy and trending down as well. no positioning issue observed. the root cause of the low pump flow issue most likely was patient condition related. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12842 | 1220648-2024-24993 | 20981728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981728&pc=OZD | 12/06/2024 | 01/24/2025 |  | 12/19/2024 |  |  | 12/06/2024 | 12/19/2024 |  |  | 10/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423603 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in acute myocardial infarction / cardiogenic shock. the pump was placed but the team observed a hematoma and pseudoanyreusm that the team infused red blood cells back to the patient for and explanted the pump. the explanted pump noted a thrombus at the pump. the patient survived the event. | [5 empty] | The investigation of vascular damage - peripheral vessel and thrombus has been completed. the root cause of the vascular damage and thrombus was not determined since clinical information was inconclusive. e4 should have been left blank on manufacturer device report 1220648-2024-24993. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
12843 | 1220648-2024-24994 | 20981808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981808&pc=OZD | 12/06/2024 | 01/22/2025 |  | 12/19/2024 |  |  | 12/06/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491690 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced some bleeding from their axillary access site following impella 5.5 implant. blood products were given to treat the condition and support with the implanted pump was maintained. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12844 | 1220648-2024-24995 | 20981809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981809&pc=OZD | 08/13/2022 | 05/23/2025 |  | 12/19/2024 |  |  | 08/13/2022 | 12/19/2024 |  |  | 05/31/2024 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219408 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, stroke alert was called for alternate mental status and right sided weakness. additionally, the patient was positive for left frontal lobe infarct. it was also noted that the patient went into ventricular fibrillation (vf) arrest and the family decided to withdraw care. the patient was on impella support for 82.05 hours before the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | To make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation; timing of the stroke in relation to impella support (during or post-implant); detailed description of the symptoms experienced by the patient; neurological examination findings and any focal deficits observed; diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic); laboratory test results, including coagulation profiles and cardiac markers; any relevant procedural or device related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications; response to any previous anticoagulation or antiplatelet therapies; evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12845 | 1220648-2024-24996 | 20981832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981832&pc=OZD | 12/07/2024 | 01/30/2025 |  | 12/19/2024 |  |  | 12/07/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support an acute myocardial infarction / cardiogenic shock patient. the team observed a pericardial effusion of unknown etiology. the team is unsure if the effusion was prior to or post pump placement. the patient survived the event and a drain was placed. | [5 empty] | The investigation for the ventricle perforation has been completed. product was returned and no cannula kinks or other abnormalities were found. the timing of the effusion is unknown, but the patient did have hypotension and narrow pulse pressure prior to impella placement. the pump also required repositioning in the cath lab. however, since the timing of the effusion remains unknown, the root cause of the pericardial effusion was not determined since insufficient clinical details were provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12846 | 1220648-2024-24997 | 20981848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981848&pc=PYX | 12/08/2024 | 09/27/2025 |  | 12/19/2024 |  |  | 12/08/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551163 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella rp flex support and during support, there were placement signal not reliable alarms. position was confirmed and support continued. additionally, the patient had access site bleeding. two units of packed red blood cells were given to the patient, and heparin was withheld. furthermore, the patient had ventricular fibrillation (vf) while on support. the patient was defibrillated, and cardiopulmonary resuscitation was performed. the patient expired after 11.07 hours of impella support after the family decided to withdraw care. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. based on the clinical details provided, the root cause of the access site bleed was determined to most likely be anticoagulation management. as the necessary information was not provided, the root cause of the vt/vf was not determined e4 should have been left blank on manufacturer device report 1220648-2024-24997. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed. returned product was investigated and the pump passed electrical testing and had no other visual abnormalities. based on returned product and data log review, the root cause of the placement signal issue was determined to be sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24997. d8/d9 device returned to manufacturer updated. d10 concomitant medical products and therapy dates updated.
12847 | 1220648-2024-24998 | 20981856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981856&pc=OZD | 09/02/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 09/02/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276515 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the placement signal was lost for a short time, and the aic was successfully changed, however the patient ultimately expired on support. | [5 empty] | The investigation for the reported loss of opm data issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs confirmed the reported placement signal failure mode given there were controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. clinical staff reported no alarms were triggered at the time they observed this behavior, which was per design for the controller error alarm hysteresis. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. however, the motor current was stayed within normal ranges for the set p-level. the reported motor current invisible was not confirmed. device analysis: the console was sent back for further investigation. the failure mode could not be reproduced. this is a known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm. abiomed is aware of the issue and is working on appropriate corrective action. this is a low probability event and is very difficult to reproduce. if needed, the clinical team can monitor patient hemodynamics and impella position with imaging. the failure mode will be monitored for trends. per the results of the clinical review and investigation, the root cause was not determined because the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
12848 | 1220648-2024-24999 | 20981867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981867&pc=OZD | 12/08/2024 | 01/16/2025 |  | 12/19/2024 |  |  | 12/08/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570543 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had access site bleeding with a hematoma. kengreal was discontinued and blood products were given. support continued. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, bleeding from multiple sites noted as per the clinical description provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24999 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 reporting contact fax number was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24999 and was added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12849 | 1220648-2024-25000 | 20981885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981885&pc=OZD | 06/22/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/22/2023 | 12/19/2024 |  |  | 03/31/2025 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335869 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | There was noted suspicion of thrombus in the impella cp cannula due to low flows and frequent inexplicable suction events. decision was made to exchange the pump. the patient was stable. | [5 empty] | Investigation into the reported issue has been completed. visual inspection found no issues on the pump. pump ran without issues under bench test. data logs were not returned for review. based on clinical description, the root cause of low flow was unable to be determined as the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
12850 | 1220648-2024-25001 | 20981888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981888&pc=OZD | 12/09/2024 | 01/14/2025 |  | 12/19/2024 |  |  | 12/09/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590702 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support placed to support an acute myocardial infarction / cardiogenic shock patient. the team was considering coronary artery bypass graft or high risk percutaneous coronary intervention due to underlying condition and comorbidities. the patient developed a cerebellar cerebrovascular accident on day four of the support, and the abdomen was distended due to possible bleed. the patient had four units of red blood cells given. the team and family were discussing the plan of care. the patient had systemic heparin stopped, the pump had no heparin, rather sodium bicarbonate. the family did the make the choice to withdraw care and patient expired. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. b.2 added date of death and required intervention selection as they were inadvertently omitted from initial manufacturer device report number 1220648-2024-25001. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25001 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12851 | 1220648-2024-25002 | 20981905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981905&pc=OZD | 12/09/2024 | 01/28/2025 |  | 12/19/2024 |  |  | 12/09/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596350 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, there were ongoing suction alarms. fellow had attempted reposition earlier. he had pulled back on the impella and readvanced it with no relief of the alarm. the pump's pigtail was interacting with some structure and against the wall. fellow repositioned into aorta. a decision was made to remove impella. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue most likely was patient condition related due to patient anatomy, md repositioned it and alarms did not resolve. pigtail stated to be interacting with some heart structure and against the wall. added h6 impact code (b)(4). | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12852 | 1220648-2024-25003 | 20981981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20981981&pc=OZD | 12/09/2024 | 01/30/2025 |  | 12/19/2024 |  |  | 12/09/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576509 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. during preparation for removal of cp, ischemia was noted in the leg. the patient survived the event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.2 revised selection as other serious was inadvertently omitted from initial manufacturere device report number 1220648-2024-25003. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-25003 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was previously entered incorrecttly on initial manufacturere device report number 1220648-2024-25003 and a new code was added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12853 | 1220648-2024-25004 | 20982014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20982014&pc=OZD | 06/20/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 06/20/2023 | 12/19/2024 |  |  | 12/31/2024 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320176 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump would not cross the aortic valve during attempted delivery due to a kinked catheter. a second pump was utilized for attempted delivery. there was no patient harm. | [5 empty] | Investigation into the reported issue has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
12854 | 1220648-2024-25005 | 20982045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20982045&pc=OZD | 05/25/2023 | 12/19/2024 |  | 12/19/2024 |  |  | 05/25/2023 | 12/19/2024 |  |  | 06/30/2023 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112323 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that impella flow dropped drastically after running for 5 mins. the pump was replaced. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. data was not return for review. the device was received for evaluation. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of low flow was due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
12855 | 1220648-2024-25006 | 20982067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20982067&pc=OZD | 12/02/2022 | 12/19/2024 |  | 12/19/2024 |  |  | 05/09/2024 | 12/19/2024 |  |  | 05/31/2024 |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228648 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient exhibited oozing around the impella incision site. the physician decided to perform percutaneous intervention on the right coronary artery due to the uncontrolled oozing. at the end of the procedure, the impella was successfully weaned and removed, and the access site was closed with perclose x2 and manual compression to achieve hemostasis. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine the root cause of the access site bleeding was determined to be use issue because of the use of scalpel to open the skin. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
12856 | 1220648-2024-25008 | 20982822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20982822&pc=OZD | 12/05/2024 | 01/29/2025 |  | 12/19/2024 |  |  | 12/05/2024 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572753 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella cp support. it was initially reported that a purge cassette failed to prime properly leading to an error message being displayed on the aic. the purge cassette was replaced and the new cassette primed without issue. roughly eleven days into support, arterial bleeding was discovered at the impella cp access site. two units of blood were given while safeguard and a sandbag were placed on the site to treat the condition. the condition was reported to be successfully managed with the documented intervention. despite successful management, care was withdrawn later that same day and the patient reportedly expired. it is unknown if the bleeding complication contributed to the patient's passing. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to patient movement as patient was restless in bed.
12857 | 1220648-2024-25007 | 20982947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20982947&pc=OZD | 05/13/2022 | 03/27/2025 |  | 12/19/2024 |  |  | 05/13/2022 | 12/19/2024 |  |  |  |  | 12/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1072117 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a pump stop noted with the automated impella controller. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. console logs from the reported day of event are inconsistent with complaint and the reported issues were unable to be identified. during testing in danvers we discovered several issues with the purge assembly that needed addressing: broken purge disk retainer and compromised disk detect flag. after purge pressure sensor assembly replacement, console passed functional check. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12858 | 1220648-2024-24874 | 20963846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20963846&pc=OZD | 01/29/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 01/29/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205186 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during a purge cassette change, the blue disc that secured the purge disc came off of the automated impella controller (aic), and the purge fluid would not run. the aic was replaced, and the issue resolved. although there was an issue with the blue disc that secured the purge disc, the aic was exchanged and support was delivered to the patient uninterrupted. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issue has been completed. data analysis: the automated impella controller (console) logs from the complaint date show that the purge pressure was low during the case and purge fluid could not be detected before the user disconnected the purge cassette. device analysis: the console was tested after arriving in field service. the purge disc retainer was confirmed to be broken (broken blue disc retainer). the cause of the issue was a broken purge retainer. this report is being filed as part of a retrospective review of historical records.
12859 | 1220648-2024-24875 | 20964082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964082&pc=OZD | 02/23/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 02/23/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 135608 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | Use of Device Problem | Obstruction of Flow |  |  |  |  | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient was admitted to the hospital after leaving work because they did not feel well, laid down at home and began agonal breathing. cardiopulmonary resuscitation was initiated and the patient was brought to the hospital and then transported to the cardiac catheterization laboratory (ccl). the patient had been administered continuous cardiopulmonary resuscitation and defibrillation, and in the ccl a temporary pacer was placed, and then the impella cp was placed. it was reported that while loading the purge cassette prior to initiating impella support, the staff was unable to line up the silver mechanism of the automated impella controller (aic) with the black mechanism of the back of the purge cassette. it was noted that the staff thought the silver mechanism of the aic was stiff and sticky when they manually rotated it. the purge cassette was inserted and a purge error alarmed on the aic several times, asking the staff to review if the purge bag was inverted. after four or five attempts a new aic was brought in and exchanged with the current aic. there were no issues purging the catheter with the new aic, and impella support was initiated. catheterization of the patient revealed a clotted left anterior descending artery which was addressed with penumbra multiple times. each time the flow was restored, the patient's artery would clot off again. despite the medical team's best efforts the patient expired. the patient's presenting condition may be more likely have impacted the event, however, the aic malfunction cannot be excluded as a possible contributing factor to the patient's demise. | [5 empty] | Automated impella controller (console) logs from the reported day of event are consistent with complaint and show multiple error_rfid messages indicative of purge cassette installation issues, as well as inability to¿ home¿ purge cassette, several ¿purge fluid not detected error¿ message screens, and ¿piston blocked¿ log entry. logs from the backup console (ic1574) confirm successful case start executed with same pump (serial number (b)(6)) and same purge cassette (sn (b)(6)) as those during case start attempt on ic6244. inspection of console ic6244 performed in danvers revealed contamination of the purge motor shaft gasket, which is consistent with the shaft being immobilized by dried-up purge fluid (with dextrose). the cause of the automated impella controller issue was contamination. this report is being filed as part of a retrospective review of historical records.
12860 | 1220648-2024-24876 | 20964258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964258&pc=OZD | 12/02/2024 | 01/20/2025 |  | 12/18/2024 |  |  | 12/02/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576576 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support placed for a surgical patient. it was reported that the purge sidearm was leaking at the sidearm. there was no harm to the patient and the patient remains on support. | [5 empty] | The investigation for the purge leak-pump has been completed. the pump was not returned for investigation. a data log review was not required for this case. according to the clinical report, a purge leak was reported from the sidearm. additionally, a work purge solution was given, and the nurse changed it to the correct one after it was addressed. the cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning states the following: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12861 | 1220648-2024-24878 | 20964291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964291&pc=OZD | 12/05/2024 | 12/31/2024 |  | 12/18/2024 |  |  | 12/05/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562300 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Additionally, during the impella weaning, there was a pump stop alarm on the automated impella controller (aic). the pump was explanted and the patient was stable. | Us complainant reported the patient was on impella 5.5 support when they had changes in their speech with right sided weakness. it was found that the patient had a right intracranial hemorrhage. anticoagulation stopped per neurology request. the patient survived the event. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. b1 updated to both adverse event and product problem. b5 updated with additional information. h6 medical device problem code updated to reflect additional information. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e4 should have been left blank on manufacturer device report 1220648-2024-24878. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12862 | 1220648-2024-24877 | 20964384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964384&pc=OZD | 12/03/2024 | 09/23/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1535515 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was an automated impella controller (aic) issue. a new pump was plugged in and the complainant received a warning saying "optical sensor error" and there was no placement or suction monitoring. the pump was then plugged into a new aic and no further issues were noted. no patient was involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of the event are consistent with the complaint. the console ic8367 was returned to field service for further inspection. the measured "5s" parameters confirmed that the signal-to-noise ratio (snr) was below the threshold. the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to no patient harm as the initial manufacture report stated no patient involvement which is incorrect. d2 common device name revised on manufacturer device report in accordance with updated procedures. d9 device returned to manufacturer date added. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect impact code no patient involvement was erroneously entered on the initial manufacturer device report number.
12863 | 1220648-2024-24880 | 20964775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964775&pc=OZD | 12/03/2024 | 02/06/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570423 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a high motor current and low purge flows with high purge pressure. the pump stopped and it was not possible to restart the pump. the pump was removed and exchanged. a thrombus was noted in the cannula upon explant. the patient outcome was noted as stable. | [5 empty] | The investigation has been completed. the pump was returned for investigation and there was biomaterial observed in the purge gap and in the cannula, around the impellers. the pump was initiated and high purge pressure (hpp) was observed leading to pump stop. the catheter was cut to check for blood ingress and no blood ingress was confirmed. the total case was for 47 minutes. data logs showed rise in the purge pressure leading to elevated motor current and a high motor current pump stop. no suction alarms were observed. the purge pressure was seen rising gradually from 800 mmhg (at case start) to 1080 mmhg in 22 minutes. saturated flows were observed at the time of hpp alarms. therefore, the root cause of the high purge pressure issue was biomaterial. the pump was returned for investigation. there was biomaterial observed in the purge gap and around the impellers. data logs were also investigated. at the time of failure, there was rise observed in the purge pressure leading to elevated motor current and a high motor current pump stop. the purge pressure rise was observed along with drop in purge flow. this elevated motor current leads to this hmc pump stop and highly calculated (saturated) flows. the pump was tried to restart three times but unsuccessful. as per the data log review and product evaluation, the root cause of the pump stop issue was biomaterial. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.9 device return date/information was added. h.6 code 4627 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24880. h.10 some of the instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24880 and were added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smartassist catheter insertion: section: axillary insertion of the impella cp with smartassist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
12864 | 1220648-2024-24881 | 20964797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964797&pc=PYX | 12/03/2024 | 01/21/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025531548 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that there was access site bleeding. manual pressure was held for an hour, but decision was made to place fem stop on site as oozing continued and to delay start of heparin. additionally, the next morning sutures were placed. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
12865 | 1220648-2024-24882 | 20964824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964824&pc=OZD | 12/04/2024 | 01/29/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp developed red urine. lactate dehydrogenase was in the 1000s. an echocardiogram ordered and showed mitral valve interfering with the inlet. the physician repositioned the impella. the patient was taken for percutaneous coronary intervention (pci) later that day and after pci was done the cp was explanted. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation hemolysis has been completed. the impella device was not returned for evaluation.tthe root cause of the hemolysis issue could not be determined since the complaint device was not returned for analysis and insufficient data was provided. no logs were available to review. g.1 revised reporting contact email as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24882. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12866 | 1220648-2024-24879 | 20964863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20964863&pc=OZD | 01/11/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 01/24/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127918 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the automated impella controller (aic) presented with a controller failure alarm and the pump abruptly stopped. it was not reported how the issue was resolved. the patient was supported a total of 50 hours on the impella device, however it was noted that the patient expired on support. given the limited information surrounding the details the patient's condition and cause of death, the aic cannot be excluded as a possible contributing factor. | [5 empty] | Automated impella controller (console) logs from the reported day of event are consistent with complaint and show that 41 minutes after successful start of pump with the serial number (b)(6), a communication breakdown occurred between pmd (sau_motor_1) and 4-in-1/epc, resulting in pump stop. prior to that, pump had been disconnected and reconnected, prior to its initialization, due to unreliable placement signal, however there¿s no evidence that would allow to link these two events (and placement signal issue was resolved upon 2nd connection of the pump). after the console was received in aachen field service (fs) it was tested in the lab using an engineering pump, but the issue was not reproduced. there were however some communication issues between pud/ahp and 4-in-1/epc observed in post-test logs. the suspected components involved in communication issues are 4-in-1 assembly, impellatronic assembly, and their communication cable (which was found during inspection to have a kink in it, which might have contributed to the issue). the cause of the console issue was most likely a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
12867 | 1220648-2024-24884 | 20965671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20965671&pc=OZD | 06/15/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/15/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1255437 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during the purge cassette change, the console was unable to recognize the purge disc. troubleshooting the issue entailed exchanging components (purge cassette, disc and aic) to test the connection. this was unsuccessful. success from changing consoles allowed the procedure (purge cassette change) to continue. no patient harm reported. | [5 empty] | The investigation for the aic purge disc/flag issues has been completed. the console was returned for evaluation. data logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the cause of the issue was the console unable to detect the purge disc due to a broken purge flag. this report is being filed as a part of the retrospective review of historical records.
12868 | 1220648-2024-24887 | 20965737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20965737&pc=OZD | 11/19/2024 | 02/19/2025 |  | 12/18/2024 |  |  | 12/02/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523672 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. pump stopped appeared and could not be restarted for the patient despite three attempts. the device was removed as a result of the failure. the patient outcome was noted as stable. | [5 empty] | The investigation of pump stop has been completed. the pump was inspected and an abnormally large purge gap was observed. no data logs were provided for the case. clinical reported an impella defective alarm was observed. no data logs were provided so it was unable to be confirmed that there were no other abnormal events during the case. based on the dates provided for the case duration, it was determined that the bearing wear occurred within the design life of the impella cp which is 8 days, per design trace matrix. the root cause of the pump stop was most likely due to bearing wear. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12869 | 1220648-2024-24886 | 20965863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20965863&pc=OZD | 06/19/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/19/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663002 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they were unable to power off their automated impella controller ( aic) following the completion of support. it was noted that the aic kept displaying prompts to disconnect the already disconnected catheter. the emergency power off steps were subsequently used to turn off the aic. there was no patient involved. | [5 empty] | Ic logs confirmed the reported failure. aic console (b)(6) started running the pump on 08 jun 2023; no issue was found through 11 jun 2023. the pump was likely disconnected on 12 jun 2023 per log device analysis: the console was tested on an engineering lab bench. tried to reproduce the issue by running a pump for an hour while the p-level was at p-7 and disconnected the pump connector cable from the aic. the console detected the disconnected pump and allowed for normal shutdown. unable to reproduce the issue. the root cause for preventing the proper aic shutdown was most likely due to an impellatronic board malfunction.
12870 | 1220648-2024-24888 | 20966107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966107&pc=OZD | 12/03/2024 | 01/27/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512594 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. following insertion, it was documented that the device was unable to advance past the innominate. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because pump was unable to pass beyond innominate and the delivery was aborted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12871 | 1220648-2024-24889 | 20966174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966174&pc=OZD | 05/11/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/12/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a clinical procedure, and automated impella controller (aic) would not detect a purge cassette. it was reported that the cassette was tried several times, but they had to swap the aic with a new unit to resolve the issue. there was no patient information provided, and no patient harm reported. | [5 empty] | Data review: event logs from the complaint date are not consistent with reported purge cassette detection issues and revealed that the automated impella controller (console) issued the purge fluid not detected error window several times when priming purge cassette (b)(6) due to the inability to advance the stepper motor/ purge piston. device analysis: the console was tested after arriving in field service. the failure mode was reproduced. the purge system would not pass the stepper motor torque test due to build up of glucose on the stepper motor gasket. the cause of the inability to detect purge fluid was due to glucose contamination on the purge stepper motor gasket. this report is being filed as part of a retrospective review of historical records.
12872 | 1220648-2024-24890 | 20966268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966268&pc=OZD | 05/17/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/17/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219576 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) could not be switched on after being received from field service after they performed preventative maintenance. it was additionally reported that it was observed the battery switch was in the "on" position and that the front led stayed off when the power cable was plugged in. there was no patient involvement with the reported issue. | [5 empty] | Console logs were not analyzed as they were not generated due to inability to power-on the console. inspection of the console revealed damaged power button that was "permanently closed", causing inability to power on the console. indentation on the button, suggested it's been actuated with a tip of a pen, which might have caused mechanical damage. after power button was replaced, issue was resolved and console was able to power up. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
12873 | 1220648-2024-24891 | 20966356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966356&pc=OZD | 05/22/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/22/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 4878 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller in use for a clinical procedure had damage. it was reported the housing for the purge disk was broken. there was no patient information provided, there was no patient harm reported. | [5 empty] | Data review: event logs were not reviewed as they were irrelevant of the reported issue. device review: the automated impella controller (console) was tested after arriving in field service. after inspection it was confirmed that purge disk retainer was broken and purge disk detect flag was missing. the cause of the issue was broken purge disc retainer and purge disc detect flag. this report is being filed as part of a retrospective review of historical records.
12874 | 1220648-2024-24892 | 20966492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966492&pc=PYX | 12/20/2023 | 01/23/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  | 08/31/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404286 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Tachycardia | Renal Impairment | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding multiple complications experienced by a patient during impella rp support. it was documented that the patient endured worsening acute renal failure with a possible relationship to the impella device. the patient was placed on continuous renal replacement therapy to treat the condition. it was also documented that the patient experienced sustained ventricular tachycardia with a possible relationship to the impella device. the patient was defibrillated and received an amiodarone bolus to treat the condition. the patient expired the following day. it is unknown if the conditions deemed possibly related to the impella device contributed to the death. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the renal failure and ventricular tachycardia (vt) has been completed. no product was returned for evaluation. data logs were reviewed and lt log of case shows no positioning alarms or persistent suction alarms. stable motor current was observed throughout case. no trends in placement signal. clinical details noted that the patient was experiencing severe rv failure and acute renal and respiratory failure due to sepsis the day before impella implant. after pump implant, patient was placed on crrt for worsening renal failure. additionally, patient experienced sustained vt overnight while on support. the root cause of the renal failure and vt cannot be definitively determined.
12875 | 1220648-2024-24893 | 20966494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966494&pc=OZD | 12/05/2024 | 01/23/2025 |  | 12/18/2024 |  |  | 12/05/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted but when peel away was removed in normal fashion, the repositioning sheath was not long enough to get to the arteriotomy, and the patient experienced bleeding. attempt was done to rewire but unable to advance wire into the artery. pump explanted and closure was attempted with percloses, and one failed. covered stent was placed by vascular surgery. the patient received two units of packed red blood cells. decision was made to go to the operating room for impella 5.5 placement. patient was hemodynamically stable throughout. the doctor stated that not being able to advance the repositioning sheath was related to the depth of tissue track. | [5 empty] | The investigation delivery issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was most likely patient condition related deep tissue tract, based on provided clinical details. the cause of the vascular damage issue was most likely patient condition related to diseased/small vessel condition, based on provided clinical details. instructions for use: delivery issues impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b2 was inadvertently left blank on manufacturer device report 1220648-2024-24893 e4 should have been left blank on manufacturer device report 1220648-2024-24893 h6 medical device problem code 1528 was reported incorrectly on manufacturer device report 1220648-2024-24893. new code was added to medical device problem code h10 incorrect instructions for use were provided in manufacturer device report 1220648-2024-24893. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12876 | 1220648-2024-24894 | 20966506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966506&pc=OZD | 12/05/2024 | 01/29/2025 |  | 12/18/2024 |  |  | 12/05/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed hemolysis. echocardiography confirmed position. it is unknown if hemolysis was resolved. the patient survived the event. | [5 empty] | The investigation for the hemolysis has been completed. the impella was not returned for evaluation. data logs were reviewed finding no motor current (mc) spike outside of p-level changes observed. no positioning issue observed. no indication observed to confirm the rc of reported issue. the root cause of reported hemolysis issue cannot be determined since the complaint device not returned for analysis and insufficient clinical data provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12877 | 1220648-2024-24895 | 20966508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966508&pc=OZD | 02/17/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 02/17/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during patient transfer to another hospital, the automated impella controller (aic) was swapped. it was reported immediately after swapping the aic, there was a purge system failure alarm. to resolve the issue, another aic was exchanged. support was delivered to the patient without interruption and there was no patient harm reported for the reported aic issue. | [5 empty] | D8/d9: date device returned to manufacturer is unknown automated impella controller (console) logs from the reported day of event are consistent with complaint and show that after pump with the serial number (b)(6) and purge cassette with serial number (b)(6) were transferred from ic1226 to ic9455, there were multiple error messages related to ¿piston blocked¿ and ¿home not found¿ events as well as rfid errors, sau_purge communication errors, and purge system failure alarms (#417, due to piston blocked). swapping purge cassette to (b)(6) did not resolve the issue. logs from the first console (b)(6) showed no purge-related issues while the same pump (b)(6) and purge cassette (b)(6) were used. after the pump was switched to ic2410 there were no further issues. console ic9455 was tested in danvers, but the issue was not reproduced. considering there¿s no indication that issue might have been caused by purge cassette, it is most likely that either purge motor shaft was initially immobilized or pud pca had an intermittent malfunction. the cause of the console issue was a defective purge assembly. this report is being filed as part of a retrospective review of historical records.
12878 | 1220648-2024-24883 | 20966511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966511&pc=OZD | 05/09/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/09/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1546472 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) had a broken flag for the purge pressure. this issue was not related to a clinical procedure and was discovered during preventive maintenance. there was no patient involvement. | [5 empty] | Data review: event logs are irrelevant of the reported failure mode. device analysis: a broken purge disk detect flag was confirmed during console inspection. the cause of the issue was a broken purge detect flag.
12879 | 1220648-2024-24896 | 20966558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966558&pc=OZD | 12/03/2024 | 01/16/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548286 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that they encountered positioning issues. the impella cp was at six centimeters and then it was pulled back to three and a half centimeters but appeared to be against the left ventricle wall. an attempt was made to reposition under echocardiogram. it was pulled back to the aortic valve, but the pump was unable to be advanced. due to further concerns of a clot or papillary tangle, the decision was made to remove the pump and replaced using side port. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related.
12880 | 1220648-2024-24897 | 20966582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966582&pc=OZD | 06/20/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/20/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103717 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility¿s biomed reported for an automated impella controller (aic) that a controller failure occurred. there was no report of patient involvement, harm or injury. | [5 empty] | The investigation for the reported kbd/rotary knob issues issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs confirm the issuing of two controller alarms i.e (b)(6) 2023 15:20:56 imcgui: event set: #24 from 80 (0) and (b)(6) 2023 15:22:42 imcgui: event set: #24 from 80 (0). device analysis: visual inspection of the console didn't reveal any loose connections. console was tested using a script to monitor the communication between keyboard and the 4 in 1 over multiple recycles and hours of operation. failure couldn't be reproduced. per the results of the clinical review and investigation, the root cause could not be determined since the issue could not be reproduced during the investigation. this report is being filed as part of a retrospective review of historical records.
12881 | 1220648-2024-24885 | 20966588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966588&pc=OZD | 05/07/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/10/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1404002 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | Obstruction of Flow | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) in use for a clinical procedure alarmed for "purge cassette priming not possible" and "pump failure alarm". the aic was replaced with a back up unit to resolve the issue. no patient information was provided, no patient harm was reported. | [5 empty] | Automated impella controller (console) logs from the reported day of event are partially consistent with complaint and show that during case start the console was unable to prime the line, but - contrary to the complaint description - there were no ¿pump failure¿ alarms, but rather ¿purge system failure¿ alarms (#417, due to piston blocked), as well as multiple error messages related to purge irregularities: piston blocked, home not found, and purge fluid not detected. imc log entries also indicate pud/sau_purge communication issues and pud resets. inspection of the console revealed purge motor shaft being immobilized and dextrose residue was found under the motor shaft gasket. this is consistent with pud communication loss due to increased current draw from motor, resulting in pbm circutry (power regulator on pbm) limiting (or ¿locking out¿) voltage to pud pca. immobilized motor shaft was unable to actuate purge cassette piston, and purge fluid was not flowing, as evidenced by very low purge pressure measurement (below 100mmhg threshold for open line). the cause of the console issue was contamination of the motor shaft gasket which immobilized the shaft, preventing purge operation. this report is being filed as part of a retrospective review of historical records.
12882 | 1220648-2024-24898 | 20966986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20966986&pc=OZD | 06/23/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/23/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the purge disc was not recognized. after changing the purge system, the same error occurred. the console was changed, and the 1st purge system was used again, and the error was resolved. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issues issue has been completed. the product was returned, and the issue was confirmed. data analysis: the reported failure has been confirmed; console (b)(6) received ¿purge disc not detected - alarm¿, event set #402 every time after purge cassette insertion. the following purge cassettes have been inserted several times, and the console received the same ¿purge disc not detected ¿ alarm¿. purge cassette: (s/n (b)(6) lot 1718683 ver 0.3). purge cassette: (s/n (b)(6) lot 1655638 ver 0.3). purge cassette: (s/n (b)(6) lot 1718683 ver 0.3). see the attachments section for ic6073 aic logs from the complaint date.pdf. device analysis: the console was tested after arriving in fs. upon visual inspection, it was confirmed that the purge flag was missing. no dextrose residue was found, and the purge retainer was in good shape. the "purge disc detect flag" was confirmed to be missing. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
12883 | 1220648-2024-24899 | 20967064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967064&pc=OZD | 12/04/2024 | 01/21/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571808 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day two of impella 5.5 support, a high purge pressure system blocked alarm occurred. tissue plasminogen activator was administered. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the cause of the high purge pressure issue was most likely biomaterial, based on given clinical detail and data log review.
12884 | 1220648-2024-24900 | 20967189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967189&pc=OZD | 12/04/2024 | 01/27/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had access site bleeding. surgicel was applied and pressure was held. the pump was explanted due to not being able to anticoagulate the patient. the procedural outcome was successful. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12885 | 1220648-2024-24901 | 20967442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967442&pc=OZD | 07/19/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 07/19/2022 | 12/18/2024 |  |  | 01/31/2024 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174107 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced an infection at the impella access site. due to the catheter infection, the medical team planned on removing both the impella cp and decannulated extracorporeal membrane oxygenation, however before this occurred the infection spontaneously subsided and the impella pump was replaced instead. the impella cp was successfully explanted and replaced with an impella 5.0.. | [5 empty] | The impella automated controller device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
12886 | 1220648-2024-24902 | 20967601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967601&pc=OZD | 06/26/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/27/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1259148 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during case preparation, the aic started to make a strange noise during purge cassette priming. switched to another aic with no issue. the same purge cassette was used, so the hospital assumed the issue was related to aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic received the rfid error continuously for 10 seconds until the cassette (s/n (b)(6)) was removed entirely. upon cassette reinsertion, the ¿rfid error¿ was continuous for 14 seconds, most likely due to the purge cassette not being completely inserted into the purge motor spindle, which might have resulted in the reported complaint - strange noise. no real time log was available to check for "purge pressure" and "purge flow ticks" since there was no pump connected on the case associated with reported failure mode. device analysis: the console was tested on an engineering lab bench. the strange noise was reproduced by placing the purge cassette close to the aic spindle without completely inserting it into the spindle, and during the strange noise, aic received the rfid error continuously until the purge cassette was removed entirely from aic or completely inserted into the spindle. per the results of the clinical review and investigation, the root cause was determined to be most likely due to a use issue. this report is being filed as part of a retrospective review of historical records.
12887 | 1220648-2024-24903 | 20967683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967683&pc=OZD | 12/03/2024 | 01/16/2025 |  | 12/18/2024 |  |  | 12/03/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025523447 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. following insertion, the device was unable to cross the anastomosis. the implant was aborted and an impella cp pump was placed for patient support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12888 | 1220648-2024-24904 | 20967690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967690&pc=PYX | 06/16/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/16/2023 | 12/18/2024 |  |  | 01/31/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024313920 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, there was an active leak from the grey hub of the purge sidearm of an impella rp. no intervention was required, and the decision was made to continue with support using the leaking device. | [5 empty] | The investigation for the reported purge leak - pump issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the purge leak issue was a leak from the sidearm, but the exact location of the leak was unknown since the product was not returned and sufficient clinical information was not provided. impella rp instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records.
12889 | 1220648-2024-24905 | 20967934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967934&pc=OZD | 06/28/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/28/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148257 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the purge disc not recognized. the aic was successfully exchanged. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported urge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: the log shows that after inserting the purge cassette successfully (s/n (b)(6), lot 1707586 ver 0.3) got the purge disc not detected - alarm (event set: #402) even though there is a presence of purge disc in the aic console (ic2855). device analysis: the console was tested after arriving in fs. upon visual inspection, the purge flag was missing. no dextrose residue was found, and purge retainer was in good shape. the "purge disc detect flag" was confirmed to be missing (disc detect flag - see attached image). per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
12890 | 1220648-2024-24906 | 20967982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20967982&pc=OZD | 05/14/2023 | 02/14/2025 |  | 12/18/2024 |  |  | 05/14/2023 | 12/18/2024 |  |  | 02/28/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324697 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and positioning issue and low pump flow that resulted in the device being explanted and replaced. the patient presented to the catheterization lab and the impella cp device was placed via a right femoral artery access. the guidewire was removed, impella cp pump was turned on, and had continuous suction alarms. multiple attempts were made to reposition; however, suction alarms continued and was unable to be resolved. the decision made to replace the pump. there was no report of adverse effects to the patient as a result of this event. long-term plan of medical management and palliative care was noted due to the patient's long-term mental health issues and inability self-care. the patient was supported for a total of five days, the impella was successfully weaned and explanted without incident. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post sterile inspection checks.the pump was not returned for investigation. data logs recorded suction events and low flow while running on p3 followed by the moving to p2 with reported saturated flow for some time. this condition resolved by the repositioning the pump with switching p level to 6. the cause of the positioning issue was use related since issue was resolved by repositioning the pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
12891 | 1220648-2024-24907 | 20968054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968054&pc=OZD | 06/28/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/28/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1551790 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the aic console does not recognize the purge cassette during set up. the aic was successfully exchanged. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs show multiple issuance of the following error [(b)(6) 2023 13:26:24 imc-ahp error: 0xff 0x67 0xf7]. these errors follow piston blocked errors e.g. [(b)(6) 2023 13:26:24 purgeadmin: piston blocked (reset pud) 0]. there were instances of the purge fluid not being detected i.e. [(b)(6) 2023 13:31:16 imcgui: casestart: msg c:90 e:3022 0x825cb90 screen curr:4 prev:4 (3054)] the purge fluid not being detected was due to the purge priming issues. there was no direct evidence of the purge cassette not being detected from the logs. device analysis: visual inspection revealed dextrose residue within the of gasket of the purge motor. contamination of the stepper gasket could cause the spindle to be immobilized and the inability to prime purge fluid. per the results of the clinical review and investigation, the root cause was determined to be due to contamination on the gasket of the purge assembly. this report is being filed as part of a retrospective review of historical records.
12892 | 1220648-2024-24908 | 20968110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968110&pc=OZD | 04/25/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 04/25/2023 | 12/18/2024 |  |  | 03/31/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324145 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an unsuccessful impella 5.5 device implant attempt to a patient with cardiogenic shock from an acute myocardial infarction. the impella 5.5 device was unable to be delivered due to axillary artery calcification, and impella 5.5 implant was aborted. the decision was made to place an impella cp device for mechanical circulatory support, which was successfully completed. there was no report of adverse effects to the patient. | [5 empty] | The impella 5.5 device was not received for analysis. investigation was completed and based on clinical description, the root cause of delivery issue was most likely due to patient condition, specifically the calcification of the patient's axillary artery. this report is being filed as part of a retrospective review of historical records.
12893 | 1220648-2024-24909 | 20968136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968136&pc=OZD | 06/21/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/21/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150570 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) alarmed for purge failures. the team replaced the console and support proceeded without harm or injury from console replacement. | [5 empty] | The investigation for the reported "aic - purge issue - other" issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. aic logs showed that the purge bag change wizard was initiated twice. it was completed once successfully. for an unknown reason, the user attempted the bag change wizard a second time but could not complete the procedure, which was likely associated with inability to progress through the steps of the purge bag change wizard. there was an instance of ahp communication errors which was followed by repetitive aic log entries ¿pumptest: purge_start_bolus. the root cause of the issues during the purge bag change wizard was due to an aic sw issue (v12.1.1). this report is being filed as part of a retrospective review of historical records.
12894 | 1220648-2024-24910 | 20968272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968272&pc=OZD | 05/20/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/26/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Alarm Not Visible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the console presented with sound/alarm issues. there was noise. the patient continued on support with the pump and console despite the issue. there was no patient harm reported. | [5 empty] | The investigation into the reported issue has been completed. the device was received for evaluation. console logs do not contain any information related to the aic speaker sound quality. console was tested, but the issue was not reproduced, and no noise or distortion of the alarm sounds could be heard. this issue, however, did occur previously during manufacturing of this console and replacement of power battery manager did not resolve the issue. no other repairs were performed at that time due to inability to reproduce the issue further. it is most likely that the issue is caused by intermittent malfunction of the console¿s speakers. the console speakers were replaced and functional check of the console was performed. this report is being filed as part of a retrospective review of historical records.
12895 | 1220648-2024-24911 | 20968273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968273&pc=OZD | 12/04/2024 | 01/30/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568785 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported complications were encountered during attempted implant of an impella cp pump for mechanical circulatory support. following insertion, the pump was unable to advance due to the patients tortuous anatomy and an existing stent in the left artery. the implant was aborted and the scheduled percutaneous coronary intervention was completed without impella support. there was no resulting patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the right femoral artery was tortuous and in the left artery was a stent, which was too small, so the can¿t pass it with the impella as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12896 | 1220648-2024-24912 | 20968313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968313&pc=PYX | 12/04/2024 | 01/16/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575871 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, the patient became confused and pulled on their impella. the impella became malpositioned. the patient was stable without the pump support so the impella was explanted. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was patient movement related. g4 pma/510(k)number was reported incorrectly on manufacturer device report 1220648-2024-24912. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12897 | 1220648-2024-24913 | 20968542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968542&pc=OZD | 05/30/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/11/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1221065 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.0. | [18 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that after purge system exchange, the automated impella controller (aic) no longer detected purge disc. aic exchange resolved the issue. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. logs from the complaint date revealed that the console issued the purge disc not detected alarm several times and was unable to complete case wizard. the console was tested. the issue with properly detecting the purge pressure disc was reproduced confirming the failure mode seen in the field, and purge flag was confirmed to be missing. the cause of the issue was a broken/missing purge flag.
12898 | 1220648-2024-24914 | 20968559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968559&pc=PYX | 12/04/2024 | 01/23/2025 |  | 12/18/2024 |  |  | 12/04/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025475545 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp support and during support, there was malpositioning and the patient had hemolysis. there was concern the pump was too deep. the patient had cola colored urine, elevated plasma free hemoglobin and lactate dehydrogenase. the pump was removed and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. the event of positioning issue was removed and captured under hemolysis as the hemolysis was most likely caused by the positioning of the device. the root cause of the hemolysis is most likely due to improper positioning of the device. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12899 | 1220648-2024-24915 | 20968716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968716&pc=OZD | 07/12/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/12/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161394 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | Altered State of Consciousness | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that it failed to recognize an installed purge disc. the console was exchanged to resolve the noted issue. there was no patient harm. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flagissue has been completed. the product was returned, and the issue was confirmed. data review: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set # 402) several times for example [12/07/23 07:45:36 imcgui: event set: # 402 from 13 (1)]. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
12900 | 1220648-2024-24916 | 20968762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968762&pc=OZD | 06/21/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/21/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085769 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was an automated impella controller (aic) controller failure alarm that occurred during preparation. the user switched to another aic console. no patient was involved. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported confirmed the reported controller failure (event set #418), and there were onset purge pressure sensor voltage out-of-range errors. the console was tested on an engineering lab bench, but the failure mode was not reproduced. purge pressure current and voltage were consistently in bounds as defined by alarm 418. the root cause of the controller failure was not determined because the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
12901 | 1220648-2024-24917 | 20968978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20968978&pc=PYX | 06/15/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/15/2023 | 12/18/2024 |  |  | 09/30/2024 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023266237 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, there was low purge pressure with no leaks that would not resolve with cassette change and increased dextrose concentration. the family made the patient a do not resuscitation (dnr) and the patient expired while on support. | [5 empty] | The investigation for the reported purge leak ¿ pump and optical signal issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a placement signal not reliable alarm occurs 4 minutes into the case. the snr goes to 0, lamp steps up, contrast decreases, gain increases and light remains stable. these 5s parameters are consistent with a sensor break. the product was returned and it was confirmed that there was a crack in the sensor face. however, with limited clinical information if there were other devices used, it cannot be determined how the sensor became damaged. the impella has a hmc pump stop but then it turns back on. at the same time the pump stops the placement signal cuts out. due to this, it looks as if another device/material is interacting with the impella. however, due to lack of clinical information, it cannot be confirmed. additionally, the data logs were analyzed and the purge pressure begins to gradually drop and purge flow increases. it does not resolve with a cassette and bag change, and there were no leaks visualized, and the purge fluid exited the purge gap. the alarm did not reproduce and the purge pressure head around 450mmhg for around 10 minutes with both purge cassettes returned. the root cause of the purge leak-pump and the pump stop cannot be determined. the root cause of the optical signal issue is due to a damaged sensor. this report is being filed as part of a retrospective review of historical records.
12902 | 1220648-2024-24918 | 20969013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969013&pc=OZD | 07/04/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/04/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107076 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The automated impella controller (aic) failed to recognize the purge system disc twice, and despite a change in the purge cassette, did not allow to complete the priming. this resulted in an aborted procedure. no patient was involved since the pump was not opened or implanted. | [5 empty] | The investigation for the reported "aic - purge disc/flag issues" issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. logs from the complaint date revealed that the console issued the purge disc not detected alarm. the console was tested. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the root cause for the purge disc not detected alarm was a missing purge flag. this report is being filed as part of a retrospective review of historical records.
12903 | 1220648-2024-24919 | 20969084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969084&pc=OZD | 03/04/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 03/04/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112227 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported automated impella controller (aic) displayed "purge disc not detected". it was noted that the alarms started randomly and then resolved. it was advised to replace the aic to resolve the issue. no patient harm was reported. | [5 empty] | The investigation for the reported "aic - purge disc/flag issues" issues has been completed. the controller device has been returned for evaluation. aic logs confirmed the reported failure. the complaint was reproduced on the engineering bench. the purge flag was missing and purge disc not detected alarms were received when going through case wizard. the cause of the issue was a missing purge flag. this report is being filed as part of a retrospective review of historical records.
12904 | 1220648-2024-24920 | 20969162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969162&pc=OZD | 07/05/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/05/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomed reported for an automated impella controller (aic) that during a customer in-service of an impella cp insertion, the purge disc bracket was detached from the aic. the biomed was notified and provided a loaner until the aic is repaired. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery charging issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs from console (b)(6) show confirm that battery hot temperature alarm was issued by the aic but almost immediately cleared i.e. [(b)(6) 2023 07:22:53 imcgui: event set: #49 from 17 (0)] and [(b)(6) 2023 07:23:03 imcgui: event clear: #49 from 17 (0)]. it should be noted that this all happened right after the console was plugged into ac mains i.e. [(b)(6) 2023 07:22:51 imcgui: event clear: #109 from 17 (0)]. lt power data confirms that there was a small spike in the temperature of the batteries, but it was not above the threshold for battery temperature high (>60 celsius). device analysis: visual inspection of the internal circuitry of the console didn't reveal any issues with the console. the issue could not be reproduced in the lab when a known good pump and purge was run with the console for over an hour. per the results of the clinical review and investigation, the root cause was determined to be a glitch of the power battery management (pbm) communication triggered by the console being plugged into ac mains and doesn't warrant replacement. this report is being filed as part of a retrospective review of historical records.
12905 | 1220648-2024-24921 | 20969189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969189&pc=OZD | 05/29/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/29/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1246562 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was not transmitting. upon investigation, the engineer found that there was not an impella connect issue, however there was a ¿system self-check failure¿ error upon starting up the aic. as the reported issue occurred at start up, there was no patient involvement. | [5 empty] | The investigation into the reported issue has been completed. logs from the complaint date showed communication issues which resulted in console unable to boot up. from the logs available, there was no record showing the console had been used for patient support around the complaint occurred date due to the power on issue. console was unable to boot up, with ¿system self-check failure¿ on the screen. inspection of the power battery management (pbm) found burn mark near the capacitors which caused the syscall error and pbm1 communication issues. once the defective pbm was temporarily replaced with a known-good one, the failure mode was resolved and console could boot up normally. the root cause of syscall error was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. the pbm was replaced and a full functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
12906 | 1220648-2024-24922 | 20969229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969229&pc=OZD | 07/07/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/07/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the patient was transported from the cath lab on an aic battery then upon arrival, the bedside team plugged in the console and a ¿battery hot¿ alarm was noted. the aic console was replaced and was functioning. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery charging issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs from console ic8001 show confirm that battery hot temperature alarm was issued by the aic but almost immediately cleared i.e. [(b)(6) 2023 07:22:53 imcgui: event set: #49 from 17 (0)] and [(b)(6) 2023 07:23:03 imcgui: event clear: #49 from 17 (0)]. it should be noted that this all happened right after the console was plugged into ac mains i.e. [(b)(6) 2023 07:22:51 imcgui: event clear: #109 from 17 (0)]. lt power data confirms that there was a small spike in the temperature of the batteries, but it was not above the threshold for battery temperature high (>60 celsius). device analysis: visual inspection of the internal circuitry of the console didn't reveal any issues with the console. the issue could not be reproduced in the lab when a known good pump and purge was run with the console for over an hour. per the results of the clinical review and investigation, the root cause was determined to be a glitch of the power battery management (pbm) communication triggered by the console being plugged into ac mains and doesn't warrant replacement. this report is being filed as part of a retrospective review of historical records.
12907 | 1220648-2024-24923 | 20969346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969346&pc=PYX | 03/19/2022 | 05/23/2025 |  | 12/18/2024 |  |  | 03/19/2022 | 12/18/2024 |  |  | 01/31/2024 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023170831 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the abs were hemolyzed and two units of packed red blood cells were transfused. it was noted that there was renal failure and tolerated continuous renal replacement therapy with no urine output noted. low dose levophed was required when the treatment was stopped. ultimately the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12908 | 1220648-2024-24924 | 20969399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969399&pc=PYX | 07/13/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/13/2023 | 12/18/2024 |  |  | 06/30/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024359012 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. during support, impella pump had some noticeable audible sound at the repo sheath and some small vibrations were noted as well. there was no harm to the patient reported as a result of this incident. | [5 empty] | The investigation for the reported "introducer damage - mechanical" has been completed. the product was returned, and the reported failure observation was not confirmed. impella rp pump was returned for evaluation for noise and vibration. returned pump was visually inspected and biomaterial ingestion was observed on the outflow of the pump. pump was connected and ran to verify the noise issue and vibration reported. pump was soaked and ran in tergazyme solution water, and it was found that the noise and vibration of the pump was little more across the wall than when running away from the wall. the noise became little better after running the pump in tergazyme solution water for some time. pump was teared down to visually inspect any impeller damages in pump housing, alignment, id surface of pump housing, impeller gap, bearing assembly and magnetic shaft. no abnormalities were found related to any mechanical failure. the noise observed while running the pump is just an additional symptom during support but not relative to product damage. data logs were returned, and low flow was observed relative to motor current spike. the cause of the low pump flow issue was due to biomaterial ingestion observed on the outflow. this report is being filed as part of a retrospective review of historical records.
12909 | 1220648-2024-24925 | 20969508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969508&pc=OZD | 12/05/2024 | 02/07/2025 |  | 12/18/2024 |  |  | 12/06/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu was delivering a cp pump in an overnight emergency and kinked the .018 wire in the process. the wire was replaced. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since the pump kinked and the second was implanted successfully. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12910 | 1220648-2024-24926 | 20969534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969534&pc=OZD | 12/06/2024 | 01/16/2025 |  | 12/18/2024 |  |  | 12/06/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1699285 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that during prep while going through prompts to prime device, once click purge disk then close door prompt displayed, the automated impella controller (aic) would not advance screens. a new console was used with no issues. no patient harm has been reported. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested after and the issue with properly detecting the purge disc was reproduced, and broken purge flag was identified (missing at blue retainer see attached image). the cause of the inability to detect the purge disc was due to a broken purge flag. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-24926. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12911 | 1220648-2024-24927 | 20969544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969544&pc=PYX | 07/15/2022 | 05/28/2025 |  | 12/18/2024 |  |  | 07/15/2022 | 12/18/2024 |  |  | 10/31/2022 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2021013208 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the patient experienced persistent hemolysis. the pump was repositioned, but the condition continued. the patient was given platelets and one unit of blood as treatment for the ongoing hemolysis. after three days, the patient required continuous renal replacement therapy (crrt). the pump was explanted that same day. additional information noted that the care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12912 | 1220648-2024-24928 | 20969938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20969938&pc=OZD | 07/27/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 07/27/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137669 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Tear, Rip or Hole in Device Packaging | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the aic was received from preventive maintenance service in a damaged shipping box. the biomed requested that the aic be evaluated for potential damage. there was no patient involvement or harm. | [5 empty] | The investigation for the reported packaging issue has been completed. the product was returned, and the issue was confirmed. data analysis: upon reviewing the aic logs, no particular records pointed toward any aic-related issue. device analysis: the console was tested on an engineering lab bench. no damage to the console ic2553 hardware was found. per the results of the clinical review and investigation, the root cause was determined to be a damaged package. this report is being filed as part of a retrospective review of historical records.
12913 | 1220648-2024-24929 | 20970095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20970095&pc=OZD | 09/09/2021 | 12/18/2024 |  | 12/18/2024 |  |  | 10/28/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541587 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) with lost placement signal which did not prompt pump or aic replacement. there was no harm or injury. investigation found an opm failure. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. data logs were reviewed. the transient loss of placement signal could not be reproduced.
12914 | 1220648-2024-24930 | 20970251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20970251&pc=OZD | 10/22/2021 | 12/18/2024 |  | 12/18/2024 |  |  | 10/28/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1281138 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Complainant reported a console with a "controller error" alarm. support noted that the catheter was still in operation and that the common screen elements including motor current were viewable; white cable was seated correctly. by the time the replacement console was in the room, support was informed that the failure had resolved and placement signals were once again visible. investigation found an opm failure. no patient harm has been reported. | [5 empty] | The impella device was returned and evaluated. the cause of the controller error (transient loss of opm data) was unable to be determined because the issue was unable to be reproduced.
12915 | 1220648-2024-24931 | 20970453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20970453&pc=OZD | 10/23/2024 | 12/18/2024 |  | 12/18/2024 |  |  | 10/23/2024 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | While investigating the failure of an automated impella controller (aic), aic-purge issue and knob issues were noted. there was no patient impact since no patient was involved. | [5 empty] | The investigation for the reported "aic - purge issue" and "aic - kbd/rotary knob issues" has been completed. the cause of the loss of rpb functionality was due to a malfunctioning 4-in-1. the exact root cause of the malfunctioning 4-in-1 is undetermined. the cause of the repeated communication errors was likely due to malfunctioning pud board. the exact root cause of the malfunctioning pud board is undetermined. this report is being filed as part of a retrospective review of historical records.
12916 | 1220648-2024-24932 | 20970934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20970934&pc=OZD | 06/25/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 06/25/2023 | 12/18/2024 |  |  | 01/31/2025 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during the repositioning of the impella cp, the doctor pulled the impella into the aorta and was unable to re-cross the aortic valve. the doctor removed the impella and another cp was implanted. no patient harm was reported. | [5 empty] | The investigation into the reported positioning issue has been completed. no device was received for evaluation. per the clinical information provided, the root cause of the repositioning issue was most likely use related as the pump was accidently pulled back across the valve during repositioning. this report is being filed as part of a retrospective review of historical records.
12917 | 1220648-2024-24933 | 20970971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20970971&pc=OZD | 03/31/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 03/31/2023 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomed reported for an automated impella controller (aic) that the silver handle on back of device broke off. there was no report of patient involvement, harm or injury. | [5 empty] | The investigation for the reported damage before therapy issue has been completed. the product was returned, and the issue was confirmed. device analysis: console was sent back for further investigation. the gray bar on the back of the console is broken which was confirmed on visual inspection. per the results of the clinical review and investigation, the root cause was determined to be the gray bar on the rear breaking. this report is being filed as part of a retrospective review of historical records.
12918 | 1220648-2024-24934 | 20971486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20971486&pc=PYX | 03/11/2023 | 05/16/2025 |  | 12/18/2024 |  |  | 03/11/2023 | 12/18/2024 |  |  | 09/30/2024 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023271004 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that flows were reduced to 2.3 at p6, causing suction alarms. decision was made to explant rp. no patient harm was reported. | [5 empty] | No product was returned was investigation. logs were returned, upon investigation suction was noticed and low pump flow was observed than expected flow. the root cause of the low pump issue was not determined since product was not returned and insufficient clinical information was provided. | No product was returned was investigation. logs were returned, upon investigation suction was noticed and low pump flow was observed than expected flow. the root cause of the low pump issue was not determined since product was not returned and insufficient clinical information was provided. this report is being filed as part of a retrospective review of historical records.
12919 | 1220648-2024-24935 | 20971574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20971574&pc=OZD | 11/04/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 11/04/2022 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379222 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an optical sensor alarm after connecting an impella cp for scheduled support. a new automated impella controller was used without issue. there was no harm to the patient. | [5 empty] | Console logs (either from the cloud or the device) from the reported day of event were not available for analysis, and reported events could not be confirmed or verified. console ic6596 was not returned to abiomed for investigation.
12920 | 1220648-2024-24936 | 20971656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20971656&pc=PYX | 05/08/2023 | 12/18/2024 |  | 12/18/2024 |  |  | 05/08/2023 | 12/18/2024 |  |  | 12/31/2023 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022154395 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that difficulty was met while attempting to backload the impella rp onto the guidewire during implant. once the pump was in the right ventricle, the wire was accidentally pulled out of position and the pump was removed. however, once explanted, a clot was noted on the outflow cage and the decision was made to replace the rp pump for ongoing support. | [5 empty] | The investigation into the reported issue has been completed. no product was returned for evaluation. per the clinical information provided, the root cause of the delivery issue was patient condition related to patient anatomy. the root cause of the thrombus issue was most likely patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
12921 | 1220648-2024-24937 | 20971678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20971678&pc=PYX | 06/28/2022 | 10/20/2025 |  | 12/18/2024 |  |  | 07/06/2022 | 12/18/2024 |  |  | 02/28/2023 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022091234 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported during the second day of impella rp support, the position of the pump was not accurately sensing the location within the heart when compared to an x-ray image. as a result, the impella rp too deep and was pulled back. the patient was withdrawn from care two days later and outcome at explant was expired. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the positioning issue was unable to be determined as no logs or product were returned for analysis. g1: updated manf. email address.
12922 | 1220648-2024-24938 | 20971852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20971852&pc=OZD | 10/28/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 10/28/2022 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456795 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a possible optical signal failure experienced with this console. there was no patient harm reported. the investigation found that the console was replaced due to the issue. | [5 empty] | The impella device was returned and evaluated. the optical bench with marginal snr was most likely the root cause of the optical placement signal spiking.
12923 | 1220648-2024-24939 | 20972004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20972004&pc=OZD | 10/20/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 10/20/2022 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112876 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the door release on the automated impella controller was jammed. the device was removed from service as a result of the noted issue. a loaner was onsite for patient support as needed. there was no patient injury as a result of the reported failure. | [5 empty] | The impella device was returned and evaluated. the console was confirmed for contamination and jammed at the purge door push button.
12924 | 1220648-2024-24940 | 20972245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20972245&pc=OZD | 10/19/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 10/19/2022 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1460542 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the door release button on the automated impella controller (aic) was stuck. there was no patient involvement or harm as a result of the noted issue. the aic was replaced as a result of the issue. | [5 empty] | The impella device was returned and evaluated. the purge door was stuck due to build-up of purge fluid.
12925 | 1220648-2024-24941 | 20972336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20972336&pc=PYX | 10/07/2021 | 12/18/2024 |  | 12/18/2024 |  |  | 10/07/2021 | 12/18/2024 |  |  | 02/28/2023 |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022056192 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella presented a placement signal failure on the second day of support. pump flow was reported to be 0 lpm. the patient continued to improve so the pump was assumed to be functioning, and the patient continued on support until the device was successfully weaned. | [5 empty] | The product was returned and large amounts of biomaterial were found to have been ingested into the cannula. the pump was run and reproduced low flows until the biomaterial was ejected and then the flows returned to a normal range with an offset. the data logs from the field confirm the placement signal issues and low pump flow. at the start of the case the placement signal was lower than expected and the correct the field performed a manual re-zero. once the re-zero was performed an offset of approximately 175mmhg was added to the placement signal. this offset created a high placement signal reading and low pump flow as a result. the root cause of the initial low placement signal was determined to be ingested biomaterial. the root cause of the false flow reading was determined to be high placement signal as a result of the manual re-zero.
12926 | 1220648-2024-24942 | 20972504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20972504&pc=OZD | 09/30/2022 | 12/18/2024 |  | 12/18/2024 |  |  | 09/30/2022 | 12/18/2024 |  |  |  |  | 12/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1206339 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) does not connect to impella connect. it will on occasion connect and then shortly after the title will disappear. it was also making a loud grinding noise last time it was used to support a patient. successfully switched to a new console and no harm done to the patient. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. data analysis showed that logs and journals are consistent with the complaint. the console was connected to network for long term testing, but the universal serial bus (usb) interruption and data streaming issue was not reproduced. inspection on the microusb revealed the minor microusb compromise. the compromised micro usb (due to a defective backstrap cable) was high likely the root cause of the data streaming issue. during testing, the grinding noise was not reproduced. the console during long term testing did not manifest the issue, and the console was able to pass the functional check without issue. the root cause was unable to be determined.
12927 | 1220648-2024-24832 | 20950735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20950735&pc=OZD | 11/28/2024 | 01/14/2025 |  | 12/17/2024 |  |  | 11/28/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in with acute myocardial infarction / cardiogenic shock. the pump was placed and the patient was in heart block and had multi-vessel coronary artery disease. the patient was on support in the intensive care unit and was seen to have signs of hemolysis. the hemolysis was observed with labs rising, plasma-free hemoglobin level high, and hemolyzing labs. the hemolysis persisted until the family made choice to withdraw care. the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12928 | 1220648-2024-24833 | 20951020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951020&pc=OZD | 11/28/2024 | 04/24/2025 |  | 12/17/2024 |  |  | 11/28/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Failure to Pump |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | It was further reported that a notification was received on feb 17,2025, via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a device deficiency where "the motor current and placement signal of the impella are noted decrease quite dramatically with no obvious cause¿pump not working suddenly". | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and after insertion pump was getting sub optimal flows. the impella cp would be explanted and replaced. the patient had refractory ventricular tachycardia on the unit and return of spontaneous circulation was achieved; patient defibrillated twice. the decision was made to place ecmo on the unit, patient went into ventricular fibrillation and was shocked again. the patient taken to the catheterization lab, in ventricular fibrillation and shocked again. return of spontaneous circulation was achieved. the decision was made now to implant an impella cp device. the impella cp device was successfully placed, access made for percutaneous coronary intervention (pci) which was completed. swan-ganz catheter was placed post pci. patient receiving extracorporeal membrane oxygenation (ecmo). distal perfusion cannula (dpc) was placed post pci as well. however, once the swan and dpc were placed, the team was about to remove drape and noticed reduced impella flows and lack of motor current pulsitility. under fluoroscopy it was noticed that the impella was shallow, the physician attempted to advance, and suction alarms triggered. the physician pulled the pump catheter back and tried to readvance it with the same issues continuing to occur. at one point, the physician pulled the impella cp pump back all the way across valve and decided to place a new impella cp device. the physician thought there may have been a kink in the catheter. however, upon inspection, no issues were observed. the replacement impella cp device was successfully implanted; peel-away sheath removed, slack removed from impella and secured. upon dressing the groin site where impella sheath was located, a hematoma was noticed, managed with compression and resolved by the time the patient was fully transitioned onto neurotrauma intensive care unit where the care team planned for daily ramp up studies to determine if the impella mcs could be increased from p-3. |  |  |  |  | B.5 additional information was recieved. h.6 medical device problem code was added due to new information received in b.5. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Low pump flows: clinical details allege the motor current and ps trended down dramatically with no obvious cause. data logs were reviewed and there were no pump stop alarms or motor current spikes. for these reasons the failure mode was changed to low pump flows. the logs show the pump running on auto and automatically dropping the p-level to adjust to suction conditions several times. there were 5 suction alarms toward the end of the case, which likely correlate with the reported attempts to reposition the pump. there were no positioning alarms, however. clinical details report that the patient lost pulsatility during the case, which likely correlates with the first period of p-level drops that resolved. it is also reported that attempts to reposition the pump did not resolve suction alarms. since product was not returned for investigation and limited details were given on the patient's condition at the time of the issue, the root cause of the low pump flows could not be determined. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique as during attempts to reposition pump pulled back across the valve into aorta. h.6 code 1248 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24833. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this is one of two medical device reports associated with this event; the associated the impella cp replacement device is reported under manufacturing report number: 1220648-2024-24820. | G.2 added study as it was inadvertently omitted from previously submitted manufacturer device report number.
12929 | 1220648-2024-24834 | 20951088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951088&pc=OZD | 11/30/2024 | 01/21/2025 |  | 12/17/2024 |  |  | 11/30/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573235 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced oozing at the access site the following day. the patient was decompensating on the unit and advanced heart failure recommended further escalation to an impella 5.5 with surgery urgently. successful double barrel technique was completed when inserting the impella 5.5 with the impella cp pump in place. the impella 5.5 pump was slowly ramped up to p-8 simultaneously as the impella cp pump was turned down to p-0 with subsequent steps to successfully remove the impella cp. the patient was transferred to the unit on impella 5.5 support. the next day, the patient had oozing from the access site. surgiseal was applied and heparin was withheld due to the bleeding. additionally, it was noted the patient had amber, concentrated urine. continuous renal replacement therapy was started, and the patient's hemolysis labs improved. the patient would remain on support until successful weaning and explant of the impella 5.5 device. it was reported the patient was stable following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis, access site bleed, and renal failure has been completed. based on the clinical details provided, the root cause of the access site bleeding is most likely due to use as there was a very deep pocket from the cut down. the data logs show that there are no motor current spikes or a trend in the placement signal. there are no positioning alarms seen when the patient had hemolyzed labs. there are a few suctions alarms present on the 6th as well as the 10th. towards the end of the case the flow varies at p-4 so the p-level was turned down to stable it out. the patient was on continuous renal replacement therapy (crrt) and was most likely not able to tolerate higher flows. due to no product returned or sufficient evidence in the data logs, the root cause of the hemolysis and renal failure cannot be determined.
12930 | 1220648-2024-24835 | 20951114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951114&pc=OZD | 12/01/2024 | 02/05/2025 |  | 12/17/2024 |  |  | 12/01/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558149 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp being prepped when the priming failed and so the team replaced the pump. no patient harm has been reported. | [5 empty] | The investigation for the priming issue has been completed. the pump was returned with the red lumen still in and the pump was able to be primed. the imc logs show that the case start wizard was started and then step 1 - spike dextrose bag was completed. step 2 - insert the purge cassette and disc shows that there was a "purge disc not detected" alarm but that was cleared within a second. the cassette began deairing but it immediately prompted the user to step 3 - connect the impella catheter and then pump was plugged into the console and then prompted the user to step 6 - enter the purge fluid information. when the serial number was read, it showed that the pump had already been initialized but no ran. when looking at the rt log there is already back pressure in the system when the pump is connected and before the de-airing is done. this infers that the user connected the pump to a console before connected it to ic5681 and potentially started to prime it alluding to the back pressure seen in the rt logs and it already being initialized. however, without confirmation the root cause of the priming issue cannot be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
12931 | 1220648-2024-24836 | 20951193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951193&pc=OZD | 12/01/2024 | 01/22/2025 |  | 12/17/2024 |  |  | 12/01/2024 | 12/17/2024 |  |  | 10/31/2025 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024419860 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that in intensive care unit, the impella cp suddenly had a white curve disappear. the motor curve became flat and the mean flow was 0.3 liters per minute. after this, a suction alarm appeared. repositioning did not resolve the issue. the impella cp was explanted due to this issue. the patient outcome was noted as stable. | [5 empty] | The investigation low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12932 | 1220648-2024-24837 | 20951309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951309&pc=OZD | 12/01/2024 | 01/21/2025 |  | 12/17/2024 |  |  | 12/01/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding during support. this was treated by discussing proper angle of sheath, applying a pressure dressing and giving the patient one unit of packed red blood cells for drop in hemoglobin levels. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
12933 | 1220648-2024-24838 | 20951485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951485&pc=OZD | 12/02/2024 | 01/13/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249363 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomedical engineer reported an automated impella controller (aic) battery failure during "prep." there was no report of harm to a patient. | [5 empty] | The investigation for theautomated impella controller battery failure-red alarm has been completed. after reviewing the console logs, the reported battery failure was confirmed. there were numerous instances of battery failure (transport connection blown/missing) alarms observed the logs from the reported complaint date. the console was returned for evaluation. the reported battery failure issue was able to be reproduced. results of running the batttest utility on telnet revealed that cell 4 in battery 2 had a voltage that was significantly less than the rest of the cells in the battery. the reported battery failure was determined to be caused by a cell imbalance in battery 2. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12934 | 1220648-2024-24839 | 20951739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951739&pc=OZD | 12/04/2024 | 01/31/2025 |  | 12/17/2024 |  |  | 12/04/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417541 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that during the automated impella controller (aic) self-check, the message self-diagnosis failed displayed. troubleshooting was unsuccessful. the aic was replaced. there was no harm to the patient. | [5 empty] | The investigation for the console (aic) system self check error has been completed. data logs were reviewed which showed that during boot up, the console began having pbm communication errors. this failed communication caused the console to have a system self check as described in the complaint. the console was returned for evaluation. the failure mode from the field was reproduced. the pbm board was found to have a corroded trace that carries communication for pbm1. the corrosion was caused by a leaking super capacitor on the pbm. a firmware check confirmed that pbm1 was not communicating correctly. the cause of the pbm communication failure was a contaminated communication pcb trace. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
12935 | 1220648-2024-24840 | 20951971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951971&pc=OZD | 12/02/2024 | 01/13/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572005 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella cp implant and shortly after the implant, there were low pump flows with suction. the pump was repositioned with no resolution of the suction events. the p level was lowered and then raised to no avail. the physician elected to explant the pump due to the patient not needing the impella. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. data logs were reviewed and suction alarms with low flow found upon insertion. auto suction alarms seen with pump p-level reducing from higher to lower p-level. the root cause of low flow was most likely due to patient condition as no problem was found on the pump.
12936 | 1220648-2024-24841 | 20951986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951986&pc=OZD |  | 01/28/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541722 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the impella sterile sleeve detached in the terminal part, close to the red plug. the physician commented this is the first time this has happened and believes it is a manufacturing defect. there was no report of harm to the patient noted to be in stable condition following this event. | [5 empty] | The investigation for the product damage has been completed.the pump was not returned for investigation. clinical information is limited. without the pump, it is not possible to investigated the failure further. therefore, the root cause of the product damage - sterile sleeve issue was not determined since the product was not returned. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24841 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised manufacturing site name and address section as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24841. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. note: the malfunction event date is provided as six days after the explant date of the impella device used. the event date is being verified, and for this report it is indicated as unknown with a blank response.
12937 | 1220648-2024-24842 | 20951996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20951996&pc=OZD | 12/02/2024 | 01/16/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590709 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and one day into the support, the patient developed hemolysis. the condition was verified via elevated hemolysis lab results and pink urine output. the pump was removed in response to the condition and the hemolysis resolved. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions-including catheter position, pre-existing medical conditions, and small left ventricular volumes-may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. data logs showed the pump ran without issue during support. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24842 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 4641 was reported on initial manufacturer device report number 1220648-2024-24842 and should not have been.
12938 | 1220648-2024-24843 | 20952004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952004&pc=OZD | 12/02/2024 | 01/21/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568012 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 device as pre-operative placement for surgery, and during the surgical procedure, the pump stopped. angiogram was completed and showed the patient had "severe" multivessel coronary artery disease. an intra-aortic balloon pump was placed at this time, and the decision was made to schedule the patient for coronary artery bypass graft (cabg) surgery (3 vessels), aortic valve replacement and possible mitral valve replacement. the following decision was to implant the impella 5.5 device pre-operatively three days prior to the planned surgery. the patient came down for planned aortic valve replacement and cabgx3 with the prior placed impella 5.5 pump. during the surgical procedure the impella had a pump stop failure. the decision was made to replace with new impella 5.5 pump. the surgeon decided to recycle previous axillary graft and irrigated it, the new impella 5.5 pump successfully placed without issue, and the rest of the surgery was completed. the patient was reported to be stable following the event. | [5 empty] | The investigation for the pump stop has been completed. no product or logs were returned for evaluation. clinical details noted that pump was stopped and pulled out during surgery for approximately 50 minutes then reinserted. shortly after re-insertion, pump stop occurred. the root cause of the pump stop cannot be definitively determined as no product or logs were returned for evaluation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
12939 | 1220648-2024-24844 | 20952022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952022&pc=OZD | 12/02/2024 | 01/16/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402361 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was an automated impella controller (aic) error while powering on the console. no patient was involved. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. during the data analysis, the console showed a controller failure due to epm1 software corruption which reoccurred for every subsequent boot up. the device analysis was performed and the failure mode consistently was reproduced. the firmware was downloaded and was found to be different from the original firmware. the correct firmware was re-downloaded and the console performed as expected. the cause of the controller failure was corrupted epm firmware. it is undetermined what caused the initial corruption of the firmware. d.9 date device returned/information was added. a.1 patient identifier was inadvertently omitted from the initial manufacturer device report 1220648-2024-24844 and was added. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24844 was submitted in accordance with updated procedures. d.3 e-mail should have been left blank on the initial manufacturer device report number 1220648-2024-24844. e.1 us phone number was inadvertently omitted from the initial manufacturer device report 1220648-2024-24844 and was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24844 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12940 | 1220648-2024-24845 | 20952098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952098&pc=OZD | 12/02/2024 | 01/31/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471887 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion, it was documented that the device was unable to advance beyond the ascending aorta. the ascending aorta was said to be dissected during the failed delivery attempt. as a result, the implant was aborted and an impella cp device was implanted for ongoing support. | [5 empty] | The investigation for the reported vascular damage and delivery issue has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was inserted without the use of catheter. the root cause of vascular damage is determined to be use issue because the excessive force used in pushing the impella forward. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12941 | 1220648-2024-24847 | 20952130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952130&pc=OZD | 12/02/2024 | 01/20/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576581 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma. three units of packed red blood cells were administered to the patient and manual pressure was applied to achieve hemostasis. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12942 | 1220648-2024-24846 | 20952135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952135&pc=OZD | 12/02/2024 | 01/20/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572842 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Vascular Problem | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had hemolysis and gastrointestinal (gi) bleeding. the gi bleeding was treated by holding the heparin as it was found that both hemoccult and gastoccult were both positive. the patient also experienced massive hemolysis and a kink was found in the aorta. as a result, the device was removed and replaced. | [5 empty] | The investigation for the hemolysis and gastrointestinal bleed (gi) bleed has been completed. pump and data log were not provided for the investigation. as per the clinical, on third day of implant on 12/02, hemolysis and the gi bleed was reported. doctor observed kink during hemolysis event. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between gi bleed and pump was unknown. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: h6 clinical code 4441 changed to 4476. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
12943 | 1220648-2024-24848 | 20952154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952154&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1848957 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 93 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller was inadvertently dropped during patient support. it was stated that when the patient was being carried out of an ambulance, the aic was dropped. the pump was reconnected to an aic owned by the hospital and the potentially damaged aic was sent for inspection. there was no patient harm resulting from the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. the console logs from the reported day of event do not contain any evidence of malfunction during transport and while pump was used. minor damage of console housing was confirmed during device inspection and was considered a cosmetic defect which did not affect device functionality. during testing, the console operated within specification. the root cause of the issue was use related. e4 should have been left blank on manufacturer device report 1220648-2024-24848. g2 report source; foreign was missing on manufacturer device report 1220648-2024-24848.
12944 | 1220648-2024-24849 | 20952174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952174&pc=OZD | 12/02/2024 | 01/09/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon initial implant, a large calcification was noted on angiogram in the left subclavian artery proximal to the graft anastomosis. multiple attempts made using balloon angioplasty, without success. the doctor performed an axillary endarterectomy and patch repair of left axillary artery to create an adequate vessel for impella 5.5 insertion. impella 5.5 was placed in laa, in the operating room, with cutdown and arterial repair performed. despite the impella not causing the ischemia of the vessel, the endarterectomy was needed for the impella implant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
12945 | 1220648-2024-24850 | 20952179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952179&pc=PYX | 12/04/2024 | 01/08/2025 |  | 12/17/2024 |  |  | 12/04/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529186 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella rp insertion due to patient body mass index, the doctor considered high risk accessing the right internal jugular (rij) so decision was made to use the right femoral vein (rfv). during initial access, the doctor perforated the vein with 23 fr sheath. perclose applied with manual pressure. sheath removed and re-accessed below, and pump was able to be implanted successfully. the doctor made decision to leave peel away in place for 24 hours to help with hemostasis. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. no product was returned for analysis. clinical details mention right femoral vein access was used due to patient bmi. the root cause of the vascular damage - peripheral vessel was most likely patient condition related. instructions for use: impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ e4 should have been left blank on manufacturer device report 1220648-2024-24850. h10 instructions for use was missing on manufacturer device report 1220648-2024-24850.
12946 | 1220648-2024-24851 | 20952183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952183&pc=OZD | 12/04/2024 | 01/15/2025 |  | 12/17/2024 |  |  | 12/04/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464568 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support ongoing for a non-st elevated myocardial infarction patient. the pump was placed to the left ventricle but there was limb ischemia. the ischemia was treated by the placement of an antegrade sheath / 5fr catheter line. the limb was perfused, and the pump remained on for support. the patient survived the event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24851. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12947 | 1220648-2024-24852 | 20952213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952213&pc=OZD | 12/04/2024 | 01/13/2025 |  | 12/17/2024 |  |  | 12/04/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support there was access site bleeding. heparin was stopped temporarily. purse string sutures were applied and the site was redressed. the procedure outcome was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since hemostasis achieved by adding purse string suture.
12948 | 1220648-2024-24853 | 20952403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952403&pc=OZD | 10/28/2024 | 01/09/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025527518 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cytomegalovirus (CMV) Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification via received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient endured bilateral lower extremity edema and cytomegalovirus infection. the patient was on direct oral anticoagulants due to prior pulmonary embolism but was given small effusion post endometrial biopsy. this was held off. swelling worsened even after starting lasix and stopping amlodipine. the patient continued on treatment dose intravenous ganciclovir. the patient was on support until bridged to transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and infection has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and infection could not be determined.
12949 | 1220648-2024-24854 | 20952558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952558&pc=PYX | 12/02/2024 | 03/14/2025 |  | 12/17/2024 |  |  | 12/03/2024 | 12/17/2024 |  |  | 11/30/2025 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024435398 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Sepsis | Cardiac Perforation | [10 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that two days prior to explant the patient developed pericardial effusion with subsequent tamponade in early morning. ct surgery evacuated effusion and patient subsequently developed right ventricle perforation. perforation repaired at bedside by ct surgery. no change was noted in impella rp flex position and performance. | Additional information was received from the clinical site and there was no allegation against the impella rp flex for ventricle perforation that occurred while on support. | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, the patient experienced hemolysis. the patient was administered four units of red blood cells. also, there was suspected clot ingestion as low pump flow was experienced. tissue plasminogen activator protocol was given via purge and increased partial thromboplastin time therapeutic goal. the interventionist discussed switching purge solution to five percent dextrose (d5) and heparin versus d5 and sodium bicarbonate and the possibility of starting dialysis. additionally, sepsis was suspected. the patient continued to be monitored until successful weaning and explant. |  |  |  | The investigation low pump flow, hemolysis, ventricle perforation, and infection/sepsis has been completed. upon visual inspection, no biomaterial or damage to pump was observed. clinical details noted that clinical team suspected clot ingestion that was causing hemolysis and administered tissue plasminogen activator (tpa) to purge in attempt to resolve hemolysis. additionally, pump position was confirmed on echo. the root cause of the hemolysis was most likely due to biomaterial. data logs were reviewed and lt logs show pump flow is lower than normal range for p-7 at time when ps is shifted up. rt log of time of placement signal shift shows pump flow drops and all pump parameters becoming fully saturated. clinical details noted that low pump flow was observed in the field around the same time as the hemolysis was occurring. the root cause of the low pump flow was most likely due to biomaterial. the root cause of the ventricle perforation and infection/sepsis cannot be definitively determined as insufficient information was provided. d9 date device returned to manufacturer added. instructions for use: impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella rp smartassist system with automated impella controller & rp flex with smart assist system with automated impella controller: section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smartassist system with automated impella controller: section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp smartassist system with automated impella controller: section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ b5 updated with additional information that was inadvertently omitted from manufacturer device report 1220648-2024-24854. h6 health effect - clinical code 2513 and health effect - impact code 4624 were inadvertently omitted from manufacturer device report 1220648-2024-24854. h10 instructions for use regarding ventricle perforation were inadvertently omitted from manufacturer device report 1220648-2024-24854. | B.5 additional information was received. e.1 revised fxa number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24854-1. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24854-1. according to the additional information received from the clinical site regarding ventricle perforation the following have been revised. h.6 code 2513 and 4624 were reported in error on follow up manufacturer device report number 1220648-2024-24854-1. h.10 investigation results and instructions for use regarding ventricle perforation were reported in error on follow up manufacturer device report number 1220648-2024-24854-1. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer: ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12950 | 1220648-2024-24855 | 20952868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20952868&pc=OZD | 07/09/2024 | 01/15/2025 |  | 12/17/2024 |  |  | 12/04/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479145 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured klebsiella variiocola bacteremia with a "possible" relationship to the impella device used. it was reported that roughly 53 days into impella 5.5 support, a patient's blood cultures came back positive for klebsiella variicola. the patient was initially started on vancomycin and merrem with the addition of micafungin two days later. the bacteria was believed to be related to the impella. support was continued with the implanted pump for another seven days until planned explant. the patient was said to have recovered from the condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12951 | 1220648-2024-24856 | 20953742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953742&pc=OZD | 02/06/2021 | 12/17/2024 |  | 12/17/2024 |  |  | 02/06/2021 | 12/17/2024 |  |  | 10/31/2022 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021017443 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being implanted with an impella 5.5 device for mechanical circulatory support to replace an existing impella device and to decannulate venoarterial extracorporeal membrane oxygenation. it was reported that the impella 5.5 implantation through the right axillary artery had to be aborted due to stenosis. an impella cp was instead successfully inserted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. this report is being filed as part of a retrospective review of historical records.
12952 | 1220648-2024-24857 | 20953843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953843&pc=OZD | 12/02/2024 | 02/20/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448731 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the pump flow rate decreased. a suction alarm appeared and then the flow rate completely disappeared. activated clotting time (act) at the time of insertion was 250 seconds or more. the act was kept for 180 seconds or more during management. dual antiplatelet therapy was also being taken orally after stent placement. when the pump was checked after removal, a thrombus-like foreign body was found in the cannula. hemolysis red urine was observed while the patient was receiving impella cp support alone. the patient was escalated to the impella 5.5 for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12953 | 1220648-2024-24858 | 20953845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953845&pc=OZD | 11/30/2024 | 02/27/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570052 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and during the axillary artery implant, the pump kept kinking and would not advance. the device was removed and a new impella was placed directly. the patient remained stable and no reported harm. | [5 empty] | The investigation of delivery issues has been completed. based on the clinical details the root cause of the delivery issue was most likely patient condition due to patient's anatomy. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
12954 | 1220648-2024-24859 | 20953879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953879&pc=OZD | 04/03/2023 | 12/17/2024 |  | 12/17/2024 |  |  | 04/03/2023 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323599 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) displayed a purge disk not detected alarm. to resolve the issue, the console was exchanged for a backup and reportedly the purge cassette was detected. it was reported this issue occurred during a clinical case; however no patient information was provided. there was no patient harm reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issues has been completed. console logs from the reported day of event are mostly consistent with complaint and show that after purge cassette was inserted (prior to case start wizard launch), the ¿purge disk not detected¿ alarm triggered (according to log timestamps 8 seconds after purge cassette was inserted), but self-resolved immediately. during case start, there was a delay until pump was connected (after a prompt), followed by uncharacteristically long priming (over 3.5 minutes). subsequently, purge cassette was briefly disconnected and reinserted (twice), and eventually pump and purge cassette disconnected from ic5645 and presumably transferred to a backup (although we were unable to identify the backup console sn, and whether the issue was fully resolved). console was tested in field service, and no deficiencies of the purge system were discovered, and reported issue (either purge disk or purge cassette not detected) was not reproduced. purge disk retainer and disk detect flag were not compromised, although there was some minor contamination under the retainer, most likely resulting from prior purge fluid leak. case start using service pump and purge cassette was successful, and priming was not unusually prolonged as it was in the field. during repeated testing it¿s been observed that with the purge door/lid open, the purge cassette sometimes dislodges itself during priming, which appears in aic log as purge cassette briefly disconnected, however in the field, purge cassette disconnection occurred after priming was completed. the cause of the issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
12955 | 1220648-2024-24860 | 20953958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953958&pc=OZD | 12/02/2024 | 10/06/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109479 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella console flow detector is not working. there was no patient involvement. | The complainant also reported that the placement signal on the automated impella controller might have disappeared during use. |  |  |  |  | The investigation for the console (aic) purge issue has been completed. after reviewing the imc logs for (b)(6), the reported flow detector/placement signal issues were unable to be confirmed. there were no signs of placement signal issues or alarms in the imc logs around the reported complaint date time frame. the aic was returned for evaluation. the reported flow detector/placement signal issue with (b)(6) was unable to be reproduced. the placement signal issue was unable to be determined due to the inability to reproduce. no problem was found with this aic after testing. the reported issue was likely use related. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d4 unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated.
12956 | 1220648-2024-24861 | 20953965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20953965&pc=OZD | 04/05/2023 | 12/17/2024 |  | 12/17/2024 |  |  | 04/05/2023 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1366750 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller had a battery failed that occurred during a clinical procedure. there was no patient harm reported. | [5 empty] | Data analysis: automated impella controller (aic) logs indicate that this battery failure alarm was triggering since the console was shipped back to the customer after service. routinely, field service (fs) disengages the battery transport connection switch for shipping the console back to the customer after a preventative maintenance (pm) procedure. it is the responsibility of the customer to re-engage the transport connection switch. a reminder note is adhered to the screen of the console to remind them of this duty. device analysis: with the transport connection switch properly engaged, the console was power cycled multiple times while on battery and ac power. no battery failure alarms were triggered throughout testing. however, it was observed the battery failure buzzer triggered erroneously. upon review of a video obtained by (b)(6), the battery failure buzzer on pbm circuitry was sounding even though the batteries were not disengaged, and on the aic¿s ui, there was a 100% battery level with no alarms present. the false sounding of the battery failure alarm could not be determined due to the inability to reproduce. however, this failure is isolated the pbm pca based on the rationale above. it unlikely to be related to the reported complaint. the battery failure alarm in the field was most likely due to a disengaged transport connection switch that was not re-engaged by the customer upon receipt of the console after service. this report is being filed as part of a retrospective review of historical records.
12957 | 1220648-2024-24862 | 20954157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20954157&pc=OZD | 12/02/2024 | 01/16/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563253 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed as bailout to the right femoral artery after left heart catheterization. abdominal distended and large hematoma around impella site was noted. additionally, unable to doppler pedal pulse. hgb dropped from 12.5 to 6 and packed red blood cells were infused. the patient was taken to the operating room to escalate from cp to 5.5 and to repair the right femoral artery. the cp was pulled to descending aorta and 5.5 placed. the cp was removed per vascular surgery, cutdown performed and fem artery repaired. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia and vascular damage has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and vascular damage could not be determined as the device was not returned nor sufficient clinical details. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.6 codes 4641 and 4625 were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24862 and were added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-24862. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
12958 | 1220648-2024-24863 | 20954768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20954768&pc=OZD | 11/29/2024 | 01/27/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590604 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the repo sheath was advanced to groin at skin level and oozing was noted where blue suture hub is connected and white cable part outside of body. the repo sheath was damaged. the doctor attempted to do angle matching, but groin site continued oozing from the connection. the doctor decided to exchange for new pump. no oozing was noted with new pump. | [5 empty] | The investigation for the access site bleed has been completed. during the advancement of the repositioning unit, oozing was noted from the access site. it is reported that all best practices were followed during insertion and that replacing the pump with a new one resolved the complication. data logs are not relevant to the complication and the product was not returned for investigation. due to limited clinical details and the product not being returned for investigation, the root cause of the access site bleed could not be determined. b.1 product problem was selected incorrectly on initial manufacturer device report number. e.1 email was incorrect on initial manufacturer device report number and the correct current email is unknown. us phone number was incorrect on initial manufacturer device report number and the correct number was added. g.2 foreign selection was checked on initial manufacturer device report number and should not have been. h.6 an incorrect medical device problem code was noted on initial manufacturer device report number and a new code has been added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12959 | 1220648-2024-24864 | 20954993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20954993&pc=OZD | 11/25/2024 | 02/05/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564689 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, purge fluid was found leaking inside the side arm retainer and there was also a leak on the outside of the retainer. there was no significant change in the purge flow rate/purge pressure, so staff decided to observe the progress until removal. no patient harm was reported. | [5 empty] | The investigation for the purge leak has been completed. the data logs were returned but there were no alarms triggered. the impella was returned with a non-abiomed luer attached to it. this luer does not fit with a purge cassette. there were droplets of dextrose in the sidearm retainer. the impella was purged for 20 minutes and no leak was reproduced. the sidearm was then cut and a syringe leak test was performed and it was looked under a microscope. there was no leak to be found. the sidearm was then reattached to the pump with a needle and luer connection and ran for 24 hours and there was no leak to be found or any purge alarms in the data logs. purge pressure remained stable around 550 mmhg. there was no problem found as the leak could not be reproduced. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
12960 | 1220648-2024-24865 | 20955147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955147&pc=OZD | 12/02/2024 | 01/09/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555458 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. implant and it was aborted due to patient anatomy. the patient had poor vasculature and balloon inflations were performed. during this time, the patient became unstable; therefore, the patient was placed on extracorporeal membrane oxygenation. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12961 | 1220648-2024-24866 | 20955400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955400&pc=OZD | 12/02/2024 | 01/28/2025 |  | 12/17/2024 |  |  | 12/03/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the nurse mistakenly ran d5 in saline as purge solution instead of d5 and water. this caused the motor current, purge flow, and purge pressure to be out of normal ranges. tissue plasma activator was added to the purge, but the motor current remained elevated. the patient expired after 8.67 hours of impella support. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. no product was returned. the data logs show the pump was turned up to p-9. about 5 minutes after being at a steady p-level, there is a motor current spike, spike in flow and change in the placement signal. the purge pressure continues to rise for around two hours and 20 minutes before "purge pressure high" alarms are triggered. it remains high for four hours, but during this time tpa protocol was started and the purge pressure and purge flows improve. the root cause of the high purge pressure is most likely due to ingested biomaterial. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿.
12962 | 1220648-2024-24867 | 20955470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955470&pc=OZD | 08/24/2024 | 01/13/2025 |  | 12/17/2024 |  |  | 12/03/2024 | 12/17/2024 |  |  | 04/30/2024 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524400 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Two st elevated myocardial infraction door-to-unloading (stemi-dtu) studies were forwarded to the complaints team regarding a patient on impella cp support who had access site bleeding with a possible relationship to the impella device and elevated plasma free hemoglobin with a causal relationship to the impella device and procedure. for the access site bleeding, manual pressure was held. for the hemolysis, no actions were needed. the staff monitored. the patient survived the event(s). | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12963 | 1220648-2024-24868 | 20955520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955520&pc=OZD | 12/03/2024 | 01/27/2025 |  | 12/17/2024 |  |  | 12/03/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547451 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications during impella cp support. following successful implant, the patient showed some signs of bleeding when the peel away sheath was being withdrawn. the decision was made to leave the 14 fr peel away sheath in place to stem the bleeding. however, the patient then developed numbness and lost pulse in their right leg. a reperfusion retrograde sheath was placed in response to the ischemia and the issue resolved. support continued with the implanted pump following the intervention. the patient survived the event. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
12964 | 1220648-2024-24869 | 20955652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955652&pc=OZD | 12/02/2024 | 01/15/2025 |  | 12/17/2024 |  |  | 12/02/2024 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849119 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. it was noted that the automated impella controller (aic) experienced a purge block and other purge related alarms. the aic was replaced. the patient continued on support until the device was successfully weaned and the patient was bridged to a left ventricular assist device. | [5 empty] | The investigation for the console (aic) purge issue has been completed. data logs confirmed the reported complaint. the first purge system blocked (event set# 409) and purge pressure low (event set #413) alarms were triggered after 2 days along with multiple imc-ahp error messages. after some time, the purge cassette was removed and reinserted again; however, the purge system blocked (event set# 409) alarms triggered again. the rep then started the purge change wizard, but the following alarms were issued complete. the console was returned for evaluation and upon inspection, the purge assembly door was sealed, and upon closer inspection of the purge assembly, excessive purge fluid residue was found. this residue indicated a purge fluid leakage during clinical use. further inspection revealed that purge fluid residue had accumulated within the purge motor assembly, preventing the motor from spinning. upon cleaning, attempts to move the motor shaft caused the fluid residue to solidify, which resulted in a loose motor. the cause of the purge system block was identified as glucose ingress. corrections to the initial mfr report: d2 common device name updated. g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
12965 | 1220648-2024-24870 | 20955933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20955933&pc=OZD | 01/31/2024 | 01/10/2025 |  | 12/17/2024 |  |  | 11/26/2024 | 12/17/2024 |  |  | 11/30/2025 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening cardiogenic shock with a "remote (unlikely) related" relationship to the impella device used and would later expire.the impella cp was placed in a patient with cardiogenic shock from an acute myocardial infarction. the patient remained intubated, transported to the unit. the patient was additionally supported by extracorporeal membrane oxygenation (ecmo). vascular consultation confirmed bilateral pedal pulses. impella cp was taken from auto to performance level p-3 with adequate waveforms and flow. two days later, the patient remained intubated without sedation, was responsive to painful stimuli but did not follow commands. neurological function was planned for reassessment throughout the day to determine care going forward. it was noted that the team planned to attempt ecmo wean and removal the following day with impella in place in right femoral artery. vascular physician was at bedside early this day in this morning and dp doppler pulses were heard via doppler. anticoagulation was noted within therapeutic goal range with heparin purge and bivalirudin drip (with a plan to hold bivalirudin during ecmo decannulation). the peel-away sheath sutured in place in right groin and gauze dressing was saturated with frank red blood down the impella catheter ending at the peel-away/blue suture hub which is patient¿s knee. per the nurse, the patient's dressing was saturated since the day prior. the physician requested a dressing changed, replaced in order to determine if bleeding is new. the old gauze dressing was removed, insertion site cleaned and down catheter/leg and replaced gauze and tegaderm. it was noted there is no active bleeding now at such time. the patient had received two units of packed red blood cell on prior shift, and one unit finished earlier this day. the patient's hemoglobin transfusion was 7.1; however, there was no current concern for active bleeding. day three of support, updates noted the patient was not doing great, ecmo was decannulated and saturations were suffering. the following day, now day four of support, as reported by the registry, the patient endured worsening cardiogenic shock. the patient would eventually expire this same day. | [5 empty] | The investigation for the reported arrhythmia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of arrhythmia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24870 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 health effect - impact code 4617 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24870 and was added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
12966 | 1220648-2024-24871 | 20956396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20956396&pc=PYX | 03/30/2022 | 09/16/2025 |  | 12/17/2024 |  |  | 03/30/2022 | 12/17/2024 |  |  | 08/31/2021 |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1484873 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in right ventricular failure was implanted with an impella rp device for mechanical circulatory support. the patient had a previous impella cp implanted and escalated to an impella 5.5 for increased support on (b)(6) 2022. the patient was then bridged to a durable left ventricular assist device on (b)(6) 2022. the patient was later found to be in right ventricular failure and the decision was made to place the impella rp. it was reported that during impella rp placement, there was an alarm for ¿flow calculation disabled¿ and failed a manual zero attempted. the insertion was aborted and a new impella rp was successfully placed. there was no patient harm reported. | [5 empty] | . the returned data logs indicate a raw placement signal (ps) of around 100 mmhg, and a displayed ps fluctuating around -200 mmhg and out of range. since the raw ps appeared in range, the logs did not indicate any mechanical or electrical issues with the sensor voltage reading itself. since the displayed ps was unreliable, it indicates a communication issue between the raw signal input and the readout from the epm. thus, the root cause of the dp sensor failure is due to a communication issue. however, the exact communication issue cannot be pinpointed without the pump returned. this report is being filed as part of a retrospective review of historical records. | D4: udi information and expiration date was provided. h4: device manufacture date was provided.
12967 | 1220648-2024-24872 | 20958702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20958702&pc=OZD | 06/12/2023 | 12/17/2024 |  | 12/17/2024 |  |  | 06/12/2023 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported that the aic had a broken purge clip. the hospital biomed department had the aic pulled from use and a backup controller was used for support. no harm or injury to the patient. | [5 empty] | The investigation for the aic purge disc/flag issues has been completed. the data logs were reviewed but did not contain any relevant data to the failure mode. the console was returned for evaluation. the purge disc retainer was confirmed to be broken. the cause of the issue was a broken purge disc retainer. this report is being filed as a part of the retrospective review of historical records.
12968 | 1220648-2024-24873 | 20959166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20959166&pc=OZD | 05/05/2023 | 12/17/2024 |  | 12/17/2024 |  |  | 05/08/2023 | 12/17/2024 |  |  |  |  | 12/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606055 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported an automated impella controller had an alarm for "battery failure". there was no patient involvement. | [5 empty] | Data review: aic logs confirmed the reported failure. there was a battery failure alarm. the ltpowerdata from the complaint date showed cell imbalance with battery 1 (cell 2 failure) which caused the battery pack¿s internal electronics to shut down. device review: the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to failure in battery 1, due to internal cell imbalance (a known pattern failure). this report is being filed as part of a retrospective review of historical records.
12969 | 1220648-2024-24791 | 20937916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20937916&pc=OZD | 11/29/2024 | 09/28/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527615 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a impella cp pump flow issue and pump stop. the pump flow issue was indicated by a left ventricular placement signal issue with an increasing motor current trend. additionally, shortly afterwards, there was a sudden pump stop with several unsuccessful restarts. the decision was made to explant the pump and replace it with a new one. there was no reported patient harm. | The complainant also reported that the pump triggered optical signal issue. the decision was made to explant the pump and replace it with a new one. there was no reported patient harm. |  |  |  |  | The investigation of pump stop and low pump flow has been completed. the root cause of the low pump flow cannot be determined. due to insufficient data logs and product, the root cause of the pump stop cannot be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. product was not returned for investigation. the data logs were only provided for 11/18. in these data logs there are "placement signal low" alarms which coincide with the positioning of the pump. there were no other issues with the placement signal seen. the 5s parameters were within specification. due to insufficient data logs and product, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding.
12970 | 1220648-2024-24792 | 20938016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938016&pc=OZD | 11/28/2024 | 01/15/2025 |  | 12/16/2024 |  |  | 11/28/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560138 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp being delivered to the left ventricle over the .018 wire to support a patient admitted in acute myocardial infarction and cardiogenic shock. the impella cp pump was inserted into the left ventricle via abiomed .018 wire, on removal of the wire there was difficulty separating the wire from the device causing the pump to reposition back to the descending aorta. at that point decision was made to insert a new device with no complications. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issue has been completed. the complaint device for the delivery issue was changed to impella cp to match the reported product from the customer. the returned product was inspected and there were no physical abnormalities. no guidewire was returned for inspection. the clinical details mention difficulty separating the wire from the device. there were no mentions of patient anatomy issues. the root cause of the delivery issues was most likely a use-related issue. the root cause of the positioning issue was most likely use-related.
12971 | 1220648-2024-24793 | 20938031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938031&pc=OZD | 11/28/2024 | 02/05/2025 |  | 12/16/2024 |  |  | 11/28/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163155 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a self check failure. there was a purge hardware revision issue. there was no reported harm. | [5 empty] | The investigation for the console (aic) system self check issue has been completed. unable to download the logs from the console, so unable to confirm the reported failure mode with the log evidence. the aic was returned for evaluation and upon functional testing, reproduced the reported failure mode. there was no input power at the pud connector j1.1 (pin 4 &5). tracing back to the 4-in-1, zero input power was measured at the connector x25 (pin 1 & 5), which confirms the "powerswitch purge" circuit on the pbm was locked out. the reported failure mode was resolved upon replacing the pbm with the known good. the root cause of system self check failed was the pbm lockout. the pbm lockout was caused by pud pca drawing more current. corrections to the initial mfr report: d2 updated common device name. g1 mdr reporting contact email. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12972 | 1220648-2024-24794 | 20938035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938035&pc=OZD | 12/01/2024 | 02/05/2025 |  | 12/16/2024 |  |  | 12/01/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568253 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported approximately eight days after start of support, the impella cp pump stopped (greater than 15 minutes) after sudden increase of motor current. possibilities were discussed and explained on restart of the pump. a clot was suspected, and the physician did not want to take the risk with a pump restart. the patient was noted as stable, and the physician, therefore, made the decision was made to explant the impella cp device. additional information, physician comments, noted there was suspected interference with cardiac structures after mitral valve surgery. the patient was reported to be stable following the event. | [5 empty] | The investigation of pump stop has been completed. the product was returned and evaluated. there was a large purge gap indicating bearing wear found. no other abnormalities were found. log review showed a motor current spike before the pump stopped with a "impella stopped - motor current high" alarm. the pump was restarted momentarily before stopping again. no purge trends were found. the root cause of the pump stop was bearing wear. the pump stop occurred beyond the eight day impella cp indication for use per design trace matrix. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.1 revised facility name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24794. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24794 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.3 on the initial manufacturer device report number 1220648-2024-24794, it should of been marked as the device was returned and that the device evaluation was anticipated, but not yet begun. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-24794. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
12973 | 1220648-2024-24795 | 20938049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938049&pc=OZD | 12/02/2024 | 02/04/2025 |  | 12/16/2024 |  |  | 12/02/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171604 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a controller error outside of patient support. there was no patient involvement/harm. | [5 empty] | The investigation of automated imeplla controller (aic) - controller error (yellow alarm) has been completed. according to data a and device analysis the root cause of the defective battery was due to single cell failure. e4 should have been left blank on manufacturer device report 1220648-2024-24795. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12974 | 1220648-2024-24796 | 20938050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938050&pc=OZD | 12/02/2024 | 01/30/2025 |  | 12/16/2024 |  |  | 12/02/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505951 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump stopped unexpectedly during patient support. the pump was successfully restarted one minute later with no noted harm to the patient. the pump was later explanted as a result of the failure. the patient survived the event. | [5 empty] | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24796. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist for use during cardiogenic shock and high risk cpi section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
12975 | 1220648-2024-24797 | 20938138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938138&pc=PYX | 11/30/2024 | 12/03/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529521 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient post cardiothoracic surgery was implanted with an impella rp flex for mechanical circulatory support and developed low pump flow. consequently, the impella rp flex was explanted and replaced with a different support device. five days prior to the implant procedure, the patient underwent coronary artery bypass graft surgery. the patient had developed worsening cardiogenic shock with up trending inotrope/pressor requirements, and the decision was made to place an impella rp flex for support. after approximately nine hours on support. there was a concern for clot ingestions due to the rising placement signal and sudden "poor" impella rp flex flows. the decision was made to explant the impella rp flex device and a protek duo ventricular support system was used to provide right ventricular support to the patient. there was no report of adverse effects to the patient. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a device deficiency. "impella rp flex placement on 11/30 complicated by most likely thrombus ingestion with inadequate flows and massive hemolysis markers, hemodynamics much improved after exchange to protek duo cannula.". |  |  |  |  | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the returned product was inspected and clot was found in the inlet cage with blue foreign material threads. the data logs confirm low pump flows for corresponding p-level and show a motor current spike with speed deviation and a rise in placement signal with placement signal trending upwards during case duration. the root cause of the low pump flow was determined to be biomaterial. | B5 additional information for new failure modes added for hemolysis and thrombus h6 clinical codes 1886 and 4440 added.
12976 | 1220648-2024-24798 | 20938336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938336&pc=OZD | 12/02/2024 | 09/19/2025 |  | 12/16/2024 |  |  | 12/02/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | It was further reported that on day 12 of support, patient was on impella cp and ecmo therapies and experienced a pump stop. the pump was able to be restarted but with saturated flows and inverted placement signal (ps) and lv waveforms, a decision was made to explant current cp in use and replace with a new impella cp. | Us complainant reported the patient was on impella cp support. while on support, there were frequent suction alarms that were treated with albumin bolus. the patient survived the event. |  |  |  |  | The investigation for the low pump flow has been completed. clinical details noted that pump was experiencing frequent suction alarms on day 9 of support and was treated with an albumin bolus. low pump flows were not noted after bolus was administered. lt log at time of reported low flow shows repeated suction alarms. pump flows at that time were all within normal range for p-level. the root cause of the low pump flow was most likely patient condition related. clinical details further state that a pump stop occurred at the end of the case and the impella cp was replaced. rt logs at time of pump stop shows a motor current spike with speed deviation and fully saturated pump flow after spike. slight increase in purge pressure with a step down in purge flows also observed at time of motor current spike. 14 minutes after motor current spike, a pump stop occurred with a "impella stopped - motor current high" alarm. pump was attempted to be restarted 3 times with pump stops occurring immediately after. on fourth attempt, pump was able to be restarted successfully and remain on for remainder of case. pump flow becomes less saturated after pump was able to be restarted successfully. clinical details noted that a pump stop occurred but pump was able to be restarted but saturated flows were observed and ps and lv waveforms were inverted. the saturated flow was part of a cascading event caused by the ingested biomaterial that caused a pump stop, the saturated flow reported by the customer was investigated under the failure mode of "temporary pump stop". the root cause of the temporary pump stop was most likely biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ added b5 additional narrative including new failure mode., h6 med dev problem code 1503, h6 impact code 4629. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the product was not returned. upon review of data logs, it is apparent that the console is the potential cause of the issue. no problem was found with the pump via logs during the case. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number.
12977 | 1220648-2024-24799 | 20938468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938468&pc=PYX | 09/12/2024 | 01/09/2025 |  | 12/16/2024 |  |  | 11/26/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517744 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | Cardiac Perforation | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiac tamponade with a "unknown (undetermined)" relationship to the impella device used: ¿(b)(6) 2024 mediastinal exploration and washout for pericardial effusion/tamponade. removed large amount of old blood clot adherent to the heart and blood around the heart and in l chest, drain repositioned. (b)(6) 2024 bleeding around his heart, has a drain but effusion getting worse, stopped heparin, vasodilated, heart under filled. impella rp removed to prevent clotting it on (b)(6) 2024.¿. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of ventricle perforation and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings and cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella rp smartassist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-24799. b3 date of event. g2 report source; study missing. g4 pma/510(k)number was reported incorrectly. h6 health effect - clinical code 2001 was reported incorrectly. new code was added to. health effect - clinical code. h10 instructions for use missing.
12978 | 1220648-2024-24800 | 20938566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938566&pc=OZD | 11/27/2024 | 01/09/2025 |  | 12/16/2024 |  |  | 11/27/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535351 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | Vascular Dissection | [11 empty] | Death | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient with cardiomyopathy on impella cp support had a type a dissection confirmed by trans-esophageal echocardiography. the patient had no option for surgery. extra corporeal membrane oxygenation was decannulated and patient expired. cp was turned to p0. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the dissection has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues dissection could not be determined.
12979 | 1220648-2024-24801 | 20938643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938643&pc=OZD | 11/27/2024 | 01/15/2025 |  | 12/16/2024 |  |  | 11/27/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108090 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a pump support ongoing when there was an automated impella controller (aic) malfunction. the aic purge would not advance to next screen. there was no harm to the patient. | [5 empty] | The investigation for the console (aic) display issues has been completed. data logs showed multiple occurrences of the "purge disc not detected" alarm (event set #402), which would caused the case start to not proceed to the next step 3. the console was returned for evaluation and upon inspection, the issue of purge disc not detected was reproduced by connecting the purge cassette. it was confirmed that the purge flag was broken. while investigating purge system, found some dextrose residue. the root cause for the purge disc not detected alarm (event set #402) was broken purge flag. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
12980 | 1220648-2024-24802 | 20938744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20938744&pc=OZD | 11/27/2024 | 01/16/2025 |  | 12/16/2024 |  |  | 11/27/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546189 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Ischemia Stroke | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | PROTEKDUO | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. following implant, it was immediately noted that the pump was experiencing suction issues above support level p-4. as a result, a protekduo was implanted to assist with patient support. roughly 5 days into impella 5.5 support, the patient's pan-computed tomography results showed a frontal lobe ischemic infarct. no plans for intervention were made and support with the implanted pump continued. the patient remains on impella support. | The patient was unstable throughout requiring increasing vasoactive support. eventually, patient was able to be transferred to the icu where they have started to settle out. the patient condition continued to deteriorate and expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva and low pump flow has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. the cause of the low pump flow issue most likely was patient condition related since the rv was reported to be dilating out on higher than p-4 support. b2 death and date of death added. b5 updated with the patient outcome. h1 type of reportable event updated to death. h6 health effect - impact code added to reflect patient outcome the following have been reported incorrectly on manufacturer device report 1220648-2024-24802. e1 initial reporter name reported incorrectly. h6 code 4445 and 1770 were reported incorrectly. new code was added to health effect - impact code.
12981 | 1220648-2024-24803 | 20939058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939058&pc=OZD | 11/26/2024 | 01/09/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721936 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was being used for a patient case and when turned on it read self-test failure. shut down was attempted and was restarted with the same issues. a new aic was used. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | A regulatory report is no longer necessary for manufacturer device report 1220648-2024-24803 as it was found to be a duplication of manufacturer device report 1220648-2024-24828. should any new information become available, please see manufacturer device report 1220648-2024-24828.
12982 | 1220648-2024-24804 | 20939078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939078&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538036 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had oozing from the right femoral artery overnight. manual pressure was held overnight, and sandbag was placed. heparin was stopped. the patient was taken to the cath lab later that morning and the cp was repositioned, re sutured and redressed. the oozing was controlled. additionally, the patient started having runs of ventricular tachycardia (vt) and eventually coded. patient was cannulated for extracorporeal membrane oxygenation (ecmo) at bedside and taken to the lab for percutaneous coronary intervention (pci). during the pci patient was in sustained vt and had to be shocked 5 times during the procedure. it was believed the vt was due to the lack of perfusion. the patient¿s rhythm had stabilized once back to icu, but some of those runs were back. the team decided to manage that with lido and weaning off the epi. although the impella was a little shallow on echo, the team was good with the position since patient was on ecmo and the impella was utilized for lv offloading. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the root cause of the access site bleeding was most likely due to anticoagulation management since the hematoma resolved after the heparin was held. the root cause of the vt/vf cannot be determined since the product was not returned and insufficient clinical details were provided. d6a and d6b was missing from manufacturer device report 1220648-2024-24804.
12983 | 1220648-2024-24805 | 20939092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939092&pc=OZD | 12/03/2024 | 02/05/2025 |  | 12/16/2024 |  |  | 12/03/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589524 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an impella cp implant and during the implant, the pump was implanted too far deep into the ventricle. this caused the pump to kink and could not be straightened. the pump was never repositioned, and the pump was explanted. the patient¿s outcome is unknown. | [5 empty] | The investigation of positioning issues has been completed. the returned compliant device visually inspected. cannula kink near out let observed. no other abnormality found. the root cause of the positioning issue was most likely use related that led to the cannula kinked. d9 date device returned to manufacturer added. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
12984 | 1220648-2024-24806 | 20939315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939315&pc=OZD | 12/03/2024 | 01/22/2025 |  | 12/16/2024 |  |  | 12/03/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Other | Company representation |  |  |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella support and expired. the specific impella device used during the incident is unknown. no further details are available regarding the cause of death. | [5 empty] | B.2 the exact date of death is unknown. d.1 the exact brand name is unknown. however, it is known that it is an impella device. d.4 the exact model number, catalog number, serial number, lot number, expiration device and unique identifier (udi) # are unknown. however, it is known that it is an impella device. d.6a and d.6b the dates of implant and explant are known. e.1 prefix, first and last name are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported death has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of death could not be determined as the device was not returned nor sufficient clinical details.
12985 | 1220648-2024-24807 | 20939489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939489&pc=OZD | 11/30/2024 | 01/16/2025 |  | 12/16/2024 |  |  | 12/03/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552507 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues. the pump placement was aborted due to the patient's iliac size. the patient's vascular anatomy was not suitable for placement. the patient¿s outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since patient had poor vasculature.
12986 | 1220648-2024-24808 | 20939781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939781&pc=OZD | 11/29/2024 | 01/20/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524246 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported apatient was on impella cp support and during the implant, the patient had ventricular fibrillation. the impella was placed into position and was started in auto while the staff prepared to defibrilate the patient. additionally, during support, the patient had bleeding at the access site. manual pressure was held and perclose strings were clinched. support was continued and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12987 | 1220648-2024-24809 | 20939935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20939935&pc=OZD | 11/28/2024 | 01/16/2025 |  | 12/16/2024 |  |  | 11/28/2024 | 12/16/2024 |  |  | 08/31/2025 |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024396342 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, impella cp had sudden suction alarms when turning the patient. impella cp attempted to be repositioned under echocardiography at bedside and fluoroscopy in catheterization lab. both were unsuccessful, which lead to removal of pump and placement of new cp. patient developed a nose bleed and heparin was temporarily paused. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the positioning issue and vascular damage has been completed. the returned device had no physical abnormalities. the sterile sleeve was dislodged from the tuohy lock. the data logs from the case confirmed suction alarms and later showed a motor current spike with speed deviation, followed by "impella position in ventricle" alarm and a drop in flows. this occurrence aligns with the clinical mention of patient repositioning at the time of positioning issues. the root cause of the positioning issues is most likely due to patient movement. no reports of excessive force or patient anatomy issues were reported. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information related to failure was provided.
12988 | 1220648-2024-24810 | 20940157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20940157&pc=OZD | 12/04/2024 | 06/26/2025 |  | 12/16/2024 |  |  | 12/04/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468462 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support and a purge pressure system blocked alarm started overnight. changing the purge cassette did not resolve the alarm. tissue plasminogen activator was administered and the next day, the purge system was operating in normal parameters. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. data analysis: the data logs were returned. the purge pressure was stable and then about 25 hours into support the purge pressure begins to gradually rise with not motor current spike seen. the purge pressure rises over the span of 5 hours to trigger purge pressure high alarms with some dips during this time which are standard troubleshooting of the cassette. tissue plasminogen activator (tpa) was said to be started shortly after. the purge pressure remained high for about 14 hours before it trended down. there were no trends in motor current seen. device analysis; the returned product had biomaterial in the purge gap. high purge pressure was reproduced when connected to the console. there was no blood seen in the purge line but the catheter was cut near the motor housing and fluid expelled from the catheter indicating there was no blockage from the sidearm all the way to the cut. a syringe of fluid was attempted to be pushed through the motor side of the catheter and there was a blockage. conclusion: based on data log and returned product analysis, the root cause of the high purge pressure is most likely due to a buildup of biomaterial, however the cause of the buildup is unknown. d9 device returned to manufacturer date added.
12989 | 1220648-2024-24812 | 20940288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20940288&pc=OZD | 06/06/2021 | 11/12/2025 |  | 12/16/2024 |  |  | 06/06/2021 | 12/16/2024 |  |  | 03/31/2022 |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1475275 |  | 005082-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella ld device for mechanical circulatory support for a coronary artery bypass graft procedure. it was reported while on impella ld support, the patient experienced bleeding from chest tubes, and as a result heparin was removed from the purge solution. additionally, the impella ld had alarms for high purge pressure with low purge flows. abiomed support staff recommended adding tissue plasminogen activator or sodium bicarbonate to the purge solution to resolve the issue, however the physician declined troubleshooting. the physician decided to explant the impella ld and replace it with an intra-aortic balloon pump (iabp). the pump was successfully removed, and the patient was placed on the iabp without any further complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of high purge pressure and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the data logs were returned and show a gradual increase in purge pressure during the case. there was no indication that there was a kink during support. the device was not returned for review. the root cause of the high purge pressure was most likely biomaterial build-up due to the lack of heparin in the purge. vascular damage peripheral vessel: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
12990 | 1220648-2024-24813 | 20940433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20940433&pc=OZD | 11/30/2024 | 02/03/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569401 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. lysis therapy was initiated to counteract the condition and support was continued without further complication. there was no patient harm. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure is most likely due to biomaterial as the high purge pressure was able to be resolved with lysis therapy. g1 reporting contact email was inadvertently entered incorrectly on manufacturer device report 1220648-2024-24813. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
12991 | 1220648-2024-24811 | 20940676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20940676&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572144 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the impella leg looked mottled and had no pedal pulses. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 4114 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24811. a new code has been added to type of investigation codes.
12992 | 1220648-2024-24814 | 20941045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20941045&pc=OZD | 09/19/2023 | 01/31/2025 |  | 12/16/2024 |  |  | 12/03/2024 | 12/16/2024 |  |  | 03/31/2025 |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024331608 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported through an edc study that a patient experienced some bleeding roughly a month after impella support. although deemed unrelated to the impella device itself, the bleed was considered to be possibly related to the impella procedure. no additional information was provided regarding the bleed or subsequent intervention. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24814.
12993 | 1220648-2024-24815 | 20941206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20941206&pc=OZD | 11/29/2024 | 02/11/2025 |  | 12/16/2024 |  |  | 12/04/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508941 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when escalating from impella cp to 5.5 the tip of the 5.5 could not be advanced from the brachiocephalic artery, so insertion was abandoned. there was no harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. e4 should have been left blank on manufacturer device report 1220648-2024-24815. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
12994 | 1220648-2024-24816 | 20941334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20941334&pc=OZD | 12/04/2024 | 09/24/2025 |  | 12/16/2024 |  |  | 12/04/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454431 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Material Split, Cut or Torn | [5 empty] | Sepsis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient had sepsis, as blood cultures were confirmed positive for staphylococcus epidermis. the sepsis was treated by giving the patient antibiotics. the complainant also reported that there was a sterile sleeve issue, noted by the unsecured distal lock on the sterile sleeve and gap observed between that lock. it was unclear how long it had not been secured or whether the sleeve was torn. the nurse covered the exposed area with a tagederm dressing for preventative measure. the patient remains on impella support. | The complainant also reported that there was a placement signal not reliable alarm triggered and the waveforms were no longer displaying on the console. the patient was stable at the time, so the team decided to leave the pump in and continue to monitor motor current and positioning with echocardiography. |  |  |  |  | The investigation product damage (sterile sleeve damage) and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added b2 was inadvertently left blank on manufacturer device report 1220648-2024-24816 e4 should have been left blank on manufacturer device report 1220648-2024-24816. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock and cp with smartassist: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. section: warnings and cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed on optical signal issue. data logs were reviewed and there were several distinct drops in the pump's signal-to-noise ratio (snr)/contrast through the first couple days of support. additionally, the placement signal is trending downward throughout the case. returned product was investigated and there were no abnormalities noted. the pump passed the laser continuity test. it was then connected to a console and the 5s parameters were all within specification. based on the reported clinical details, data logs, and returned product, the root cause of the optical signal issues was determined to be patient condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24816. d10 concomitant medical products and therapy dates updated.
12995 | 1220648-2024-24817 | 20941349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20941349&pc=OZD | 12/04/2024 | 02/07/2025 |  | 12/16/2024 |  |  | 12/04/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561071 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that one day into support, an implanted impella cp pump experienced positioning issues. it was noted that the device was against the left ventricle wall. attempts to reposition the device were unsuccessful. the pump was subsequently removed and then reinserted to obtain proper positioning. it was additionally documented that a thrombus had developed in the left ventricle. the patient was taken to the intensive care unit where they were cannulated for extra corporeal membrane oxygenation due to the thrombus. the pump was left in place to vent at p-3. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of positioning issues and thrombus has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was use related due to improper technique. as the necessary information was not provided, the root cause of the thrombus was not determined. b2 outcomes attributed to adverse event was inadvertently omitted on manufacturer device report 1220648-2024-24817 e4 should have been left blank on manufacturer device report 1220648-2024-24817.
12996 | 1220648-2024-24818 | 20941946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20941946&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531325 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support placed to support an acute myocardial infarction / cardiogenic shock patient. the patient had been transferred from the community to the (b)(6) center. after transfer the team observed signs of hemolysis and explanted the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
12997 | 1220648-2024-24819 | 20942965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20942965&pc=OZD | 11/11/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546134 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Septic Shock | Respiratory Failure | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure and sepsis both with a "possible" relationship to the impella 5.5 device used. the patient acute on compensated ischemic cardiomyopathy with an ejection fraction of 10-15% was cardioverted for atrial fibrillation and determined to be in cardiogenic shock. the decision was made to implant an impella 5.5 for mechanical circulatory support which was successfully completed. the patient was transferred to the unit intubated and sedated, well support on performance level p-6 at 3.8l/min. the patient was a bridge to recovery as the patient was not a candidate for advanced therapies. patient continued to be intubated until stable enough for extubating which occurred two days after start of support, had endured respiratory failure and recovered with no sequelae. additionally reported, blood cultures were obtained after the impella 5.5 device implant procedure. infectious disease physician was consulted. the reason for the blood cultures was unclear but possibly due to hemodynamic instability. the patient had a fever of 38.2c two days after start of support. blood cultures were positive for methicillin-resistant staphylococcus epidermidis (mrse) and four days later, staphylococcus epidermidis. sputum culture was positive for serratia marcescens, pseudomonas aeruginosa and positive cocci. broad antibiotics vanco and cefepime were administered. the care team believed patient had mix of cardiogenic shock and septic shock though origin was unknown, and per the registry's report, the patient had not recovered from this event. however, it was noted the patient did well on support with a plan to wean and explant which was completed after 10 days of support. the documented outcome noted the patient survived at explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. h.6 code 4617 was inadavertently omitted from initial manufacturer device report number 1220648-2024-24819 and was added.
12998 | 1220648-2024-24820 | 20943962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20943962&pc=OZD | 11/28/2024 | 01/14/2025 |  | 12/16/2024 |  |  | 11/28/2024 | 12/16/2024 |  |  | 07/31/2024 |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when impella cp was implanted a hematoma was noticed upon dressing the groin site. manual compression resolved the hematoma by the time the patient was in intensive care unit. it is unknown if blood products were given, or heparin held. five days later the family made the decision to withdrew care and patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
12999 | 1220648-2024-24822 | 20944856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20944856&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691800 |  | 0042-0040-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) did not detect the purge disc and the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued the purge disc not detected alarm. the issue with properly detecting the purge disc was reproduced, and a broken purge disc retainer was identified. the root cause of this issue was a broken purge disc retainer.
13000 | 1220648-2024-24823 | 20944860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20944860&pc=OZD | 07/23/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/26/2024 | 12/16/2024 |  |  | 11/30/2025 |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425905 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella device used. it was reported that the patient/event has not recovered/not resolved. additional information noted care was withdrawn and the patient expired. there were no reported impella concerns. however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13001 | 1220648-2024-24824 | 20944862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20944862&pc=OZD | 11/27/2024 | 01/17/2025 |  | 12/16/2024 |  |  | 11/27/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568411 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to escalate from the cp pump. the pump was supporting when after 20 days the pump sleeve had damage that was troubleshot with tegaderm. additionally, there was persistent ventricular tachycardia that prompted multiple shocks. the pump remained on for support with a pacer and ablation until explant. the patient survived the event. | [5 empty] | The investigation of product damage (sterile sleeve damage) and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. the root cause of the product damage (sterile sleeve damage) is most likely due to use as there was not enough slack on the sleeve so excessive force was used. as the necessary information was not provided, the root cause of the vt was not determined. on the initial manufacturer device report number 1220648-2024-24824, the incorrect device was reported. the below section fields were revised to show the correct device information d.1 brand name, d.4 model number, catalog number, serial number, lot number, expiration date, and unique identifier (udi), d.6b explant date, h.4 device manufacture date. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13002 | 1220648-2024-24821 | 20944903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20944903&pc=OZD | 11/29/2024 | 09/26/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547762 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support and when the patient returned from the operating room a hematoma developed at surgical site. this was treated through manual pressure and giving the patient two units of blood. the patient remains on impella support. | The complainant also reported that the placement signal was unreliable. flow was otherwise stable on the pump, and no alarms were occurring. there were no changes in purge flow or pressure and no other clinical changes with the patient. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Access site bleeding: the clinical stated that the patient returned from the operating room and a hematoma developed at the surgical site. the pump was returned cut. due to lack of sufficient details, the root cause of the access site bleeding could not be determined. d6b explant date added : d9 device returned to manufacturer date added b3 date of event was erroneously entered on the initial manufacturer device report number. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. due to lack of data logs returned and insufficient product returned as it was cut, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24821. d10 concomitant medical products and therapy dates updated.
13003 | 1220648-2024-24825 | 20945044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20945044&pc=OZD | 11/29/2024 | 01/13/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590576 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an impella cp implant due to needing a bailout. after the implant, distal angio showed present, but decreased flow to left lower extremity, possible baseline due to patient's severe peripheral vascular disease (pvd). bilateral legs were mottled from the knees down. support continued. additionally, the patient had no flows which led to the pump having constant suction alarms. the patient went into pulseless electronic activity and the patient coded. the patient expired after 4.39 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the limb ischemia and low pump flow has been completed. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. placement signal was trending down, suction alarms seen with low flow. at the end, impella position in ventricle alarm was also found. product was not returned for review. the root cause of the low flow was most likely patient condition related per log and clinical analysis. the root cause of the limb ischemia could not be determined without any additional clinical details.
13004 | 1220648-2024-24827 | 20945156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20945156&pc=OZD | 12/01/2024 | 02/26/2025 |  | 12/16/2024 |  |  | 12/01/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263482 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a console failure (controller error - yellow arm) during impella cp pump support. the automated impella controller was swapped/replaced and support was continued until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported days of events are consistent with the complaint and showed multiple occurrences of controller error alarm were triggered during the clinical procedures. real time logs for one of the pumps confirmed that all signals received from the optical bench were represented as -16384 values due to the crc error. signatures of opm signals and log entries were identical for both pumps (serial number (b)(6), proving the same failure mode on both occasions. controller errors and associated with them momentary losses of placement signal were due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. the cause of the aic issue was a software issue. d9. date device returned/information was added.
13005 | 1220648-2024-24828 | 20945316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20945316&pc=OZD | 11/29/2024 | 01/30/2025 |  | 12/16/2024 |  |  | 11/29/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1721936 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller displayed a persistent self test failure message after power up. the power was cycled to troubleshoot the failure, but the issue remained. the aic was subsequently replaced for planned support. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the root cause of this issue was likely a defective impellatronic. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13006 | 1220648-2024-24826 | 20945472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20945472&pc=OZD | 11/30/2024 | 01/16/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474094 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that implant of the impella 5.5 pump was aborted. while the nurse was preparing the pump, it was briefly placed on p1 on accident while priming, but before wiring. there was a red retrograde flow alarm and due to the risk of starting an impella outside of the body before implant, a new impella 5.5 was opened and implanted for safety instead. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported priming issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. based on clinical description, the root cause of priming issue was most likely use related.
13007 | 1220648-2024-24829 | 20945556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20945556&pc=OZD | 11/30/2024 | 01/21/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535666 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump triggered a purge system blocked alarm during patient support. the purge cassette was replaced in an attempt to troubleshoot the failure but the issue remained. the site opted not to initiate a tissue plasminogen activator protocol and instead moved to replace the device. the pump was removed and a new impella cp device was placed for ongoing support. there was no patient harm. | [5 empty] | The investigation for the high purge pressure has been completed. only the console was returned for evaluation. data log review showed a large motor current spike followed by a rise in purge pressure, which increased from 500 to 1100 mmhg in 10 minutes. there were then alternating "purge pressure high" and "purge system blocked" alarms. the purge cassette was replaced with no change in purge pressure. there were initially "purge system failure" alarms and purge pressure was around 800 mmhg before purge flow increased from 0, but the purge pressure decreased down to 500 mmhg and purge flow increased to 8 ml/hr in a few minutes. the pump was unplugged after only 4 minutes of being connected and the users suspected an issue with console. however, based on log review, it appears that the pump reached an adequate purge pressure around 500 mmhg just before it was unplugged. a different pump was inserted in the console and worked without issue. the root cause of the high purge pressure is most likely biomaterial ingestion as a large motor current (mc) spike was seen on the data logs indicating ingestion and the purge pressure rise followed the spike. the pump was swapped out and high purge pressure did not reproduce on the second pump used on the same console. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13008 | 1220648-2024-24830 | 20947486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20947486&pc=PYX | 11/30/2024 | 07/02/2025 |  | 12/16/2024 |  |  | 11/30/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2025569646 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella rp flex support. while on support, pump position was lost and could not be regained. upon post-explant, a clot was visualized in cannula. a decision was made to use a new rp device. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the thrombus and positioning issues has been completed. no product was returned for evaluation. data logs were reviewed and no "pump position unknown" alarms were noted. pump flow and motor current both observed to be pulsatile throughout the case. clinical details noted that pump position was lost and could not be regained. additionally, it was noted that a clot was present in the cannula post-explant. the root cause of the positioning issues cannot be definitively determined as insufficient clinical details were provided. the root cause of the thrombus could not be determined without additional clinical details. the instructions for use for the impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The previous investigation analysis concluded that the cause of the positioning issue could not be determined. however, additional information was received which states that the pump was not fixated. based on the clinical details that the pump was not fixated, the root cause of the positioning issues is most likely due to use. added h6 conclusion code 19
13009 | 1220648-2024-24831 | 20947521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20947521&pc=OZD | 12/01/2024 | 01/09/2025 |  | 12/16/2024 |  |  | 12/01/2024 | 12/16/2024 |  |  |  |  | 12/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a thrombus in their left ventricle during impella cp support. the pump was explanted the following day and an intra-aortic balloon pump was placed in response to the thrombus. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24831.
13010 | 1220648-2024-24790 | 20933065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20933065&pc=OZD | 07/20/2022 | 12/14/2024 |  | 12/14/2024 |  |  | 07/20/2022 | 12/14/2024 |  |  | 01/31/2024 |  | 12/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023177936 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after under 1 hour of support the team determined the placement signal was lost and so the 5.0 was replaced. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. the device was received for evaluation. no data logs were returned. a significant catheter kink was found at the kink protector of the returned pump. the placement signal spiked to 3000 immediately upon connecting the pump to the aic, followed by the "placement signal not reliable" alarm. the red handle was opened and the placement signal fiber appeared to be stretched along its length and damaged near the catheter kink. therefore, it was determined that the cause of the placement signal issue was an open electrical line due to excessive force. this report is being filed as part of a retrospective review of historical records.
13011 | 1220648-2024-24736 | 20921933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20921933&pc=OZD | 11/28/2024 | 01/09/2025 |  | 12/13/2024 |  |  | 11/28/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025568578 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support and that there was arterial damage and that the patient developed limb ischemia. after the pump was inserted and started, a clot was noticed in the left artery which appeared to be same in size as the impella cannula. the patient then lost bilateral radial pulse which was found to be caused by a descending aorta dissection. it was thought that during traumatic insertion through left subclavian artery, the patient developed the dissection and lost left femoral artery flow. it was decided to leave patient's ischemic leg as is because of a reperfusion injury. the patient was then transferred to the intensive care unit (icu). however, the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the vascular damage and limb ischemia has been completed. no product was returned for analysis. clinical mentioned patient anatomy issues: tortuosity and patient¿s recent aortic aneurysm repair as contributing factors. the root cause of the vascular damage - great vessel was most likely due to patient condition. the root cause of the limb ischemia could not be determined without additional information.
13012 | 1220648-2024-24742 | 20923868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20923868&pc=OZD | 11/28/2024 | 01/15/2025 |  | 12/13/2024 |  |  | 11/28/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505132 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the groin site oozing. this was treated by placing sutures on the impella sheath and extracorporeal membrane oxygenation cannulas, holding systemic anticoagulation for four hours and giving the patient two units of packed red blood cells. the procedure was successful and the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by applying the sutures at the bleeding site.
13013 | 1220648-2024-24743 | 20924023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20924023&pc=PYX | 11/25/2024 | 01/10/2025 |  | 12/13/2024 |  |  | 11/25/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025534070 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, the pump had low pump flows. one unit of packed red blood cells was given to the patient. support was continued and eventually the flows decreased again. the flows were at 1.8 liters from 3 plus liters at p7. there were spikes in motor current and the staff has a suspicion of a clot ingestion. the pump was explanted. the clot was never confirmed. the patient remained stable and no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed after the pump started and ran for 417 hours. motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. the product was returned for review. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion (thrombus).
13014 | 1220648-2024-24744 | 20924111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20924111&pc=OZD | 11/27/2024 | 01/10/2025 |  | 12/13/2024 |  |  | 11/27/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558257 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella cp implanted and after the implant, a left ventricle perforation was noted. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system for use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. no product was returned for analysis. limited clinical information was provided. it is unknown the cause of the perforation. the root cause of the ventricle perforation was not determined as limited information related to failure was provided h.1 revised type of reportable event as it was reported incorrectly on initial manufacturer device report number 1220648-2024-24744. h.6 code 2135 was reported incorrectly on initial manufacturer device report number 1220648-2024-24744. a new code was added to health effect - clinical code.
13015 | 1220648-2024-24745 | 20924462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20924462&pc=OZD | 11/27/2024 | 01/16/2025 |  | 12/13/2024 |  |  | 11/27/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered delivery issues as the impella could not be fully inserted due to vessel size being insufficient. the vessel was sized by vascular ultrasound at 6mm inner to inner diameter. however, during the operation, the vessel was exposed and found to be much smaller. the 10mm woven graft was anastomosed to the axillary artery with a 60 degree bevel. a 5fr angled pigtail catheter was advanced to the left ventricle over a standard j wire. the j wire was exchanged for a 0.018¿ impella wire. the wire did not provide enough support and a 0.018¿ glidewire advantage was attempted. again, the tip of the impella became lodged in the axillary artery and would not advance. the physician elected to try an 0.035¿ glide advantage wire at this point and again exchanged the wire. this did not yield any different results. a 0.018¿ steelcore wire was then advanced to the left ventricle as a buddy wire. one last attempt was made to advance the impella and it was not successful. the wires and impella were then removed, the graft was trimmed and the incision was closed. the patient remained stable and no reported patient harm. | [5 empty] | E.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24745 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 correction should have been selected on the first follow up manufacturer device report number 1220648-2024-24745. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella could not be fully inserted due to vessel size being insufficient as per the clinical description provided.
13016 | 1220648-2024-24746 | 20924543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20924543&pc=OZD | 11/27/2024 | 03/27/2025 |  | 12/13/2024 |  |  | 11/27/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551223 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was a delivery issue of the impella cp. they attempted impella cp insertion in right femoral artery (rfa). however, the impella could not pass calcification in rfa and a dissection was noted. therefore, the case was aborted and rfa was stented. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed. the product was not returned for evaluation. based on clinical description, the root cause of delivery issue was most likely due to patient condition. the root cause of vascular damage was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24746. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24746 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-24746 was submitted in accordance with updated procedures. h.6 code 2993 was reported incorrectly on initial manufacturer device report number 1220648-2024-24746 and a new code has been added. h.10 additional instructions for use were added that were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24746.
13017 | 1220648-2024-24747 | 20924703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20924703&pc=OZD | 11/29/2024 | 02/03/2025 |  | 12/13/2024 |  |  | 11/29/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570740 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there were positioning issues. the pump became mispositioned in the aorta. repositioning was not optimally solvable, which is why the decision was made to wean and explant the pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. the pump was returned and was visually inspected. the cannula was noted to be slightly deformed. no cannula kink or catheter kink observed. no other issue found. clinical stated: pump repositioning was not optimally solvable. the root cause of the positioning issue cannot be determined since insufficient clinical information was provided to determined why the pump was unable to remain in the correct position. added d9 device return date.
13018 | 1220648-2024-24748 | 20925097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925097&pc=OZD | 11/26/2024 | 01/14/2025 |  | 12/13/2024 |  |  | 11/26/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple attempts were made to insert impella cp, but it was unsuccessful. there was multiple balloon inflation with 6mm/7mm/8mm, buddy wire and mineral oil. also attempted a new impella cp via axillary which was also unsuccessful. the impella placement was aborted and there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, multiple attempts made to cross right axillary with impella 5.5 & impella cp were unsuccessful as per the clinical description provided. h.4 revised as previously entered incorrectly on initial manufacturer device report 1220648-2024-24748. h.6 medical device problem code 2682 was missing from initial manufacturer device report 1220648-2024-24748 and was added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13019 | 1220648-2024-24749 | 20925227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925227&pc=OZD | 07/23/2022 | 09/25/2025 |  | 12/13/2024 |  |  | 07/23/2022 | 12/13/2024 |  |  | 04/30/2024 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1466562 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while the patient was in the operating room on impella 5.0 support, the automated impella controller experienced purge issues with no resolution specified. the purge issues continued intermittently throughout the patient's support. there was no reported patient harm. the patient's outcome at explant is unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of the purge leak-pump has been completed. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. corrections to the initial mfr report: h10 it was initially reported that a device was returned however a device has not been returned. investigation has been completed. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
13020 | 1220648-2024-24750 | 20925359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925359&pc=OZD | 10/17/2022 | 12/13/2024 |  | 12/13/2024 |  |  | 10/17/2022 | 12/13/2024 |  |  | 04/30/2024 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236424 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Noise, Audible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the pump began making a loud buzzing sound during use. the medical team was concerned there would be a pump stop, and decided to wean and explant the pump. there was no patient harm reported, the patient was successfully weaned and the impella was explanted. | [5 empty] | There was no problem found with this pump after product was returned. all parameters were in specification and it ran in the lab with no issues. this report is being filed as part of a retrospective review of historical records.
13021 | 1220648-2024-24753 | 20925558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925558&pc=OZD | 03/29/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 03/29/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395994 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) presented with a controller error alarm upon boot-up. the aic was rebooted, but the issue did not resolve. there was no patient involvement. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show controller error alarm #(b)(6) due to opm communication loss, that was persistent upon each boot-up. inspection of the console performed in aachen by field service (fs) revealed misaligned communication cable between optical bench (opm) and 4-in-1 carrier. after the cable was re-seated, the console was able to boot-up without further controller error alarms. it is most likely that the cable was inadvertently misaligned during prior preventative maintenance procedure, as only fs would have a reason, and be able to, open the console¿s housing. the cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
13022 | 1220648-2024-24752 | 20925584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925584&pc=OZD | 03/21/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 03/21/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288345 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during preparation and prior to implant, the automated impella controller device for the impella cp displayed a red alarm for battery failure. the battery level displayed 0, despite being charged or hours. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that when automated impella controller (aic) (b)(6) was started on 2023-03-21. it crashed and upon restart it presented with battery failure alarm (#360) and unexpected controller shutdown alarm (#603). the battery alarm remained unresolved after aic power-cycle. the console log also shows that the last time this console was used prior to the day of event, was 2022-05-09. power and battery information embedded in long tem log data shows batteries fully charged on (b)(6) 2022, and fully depleted (internally shut down due to low voltage of all cells) on (b)(6) 2023. it is most likely that batteries were depleted due to aic not plugged into wall ac while in storage for 11 months. the cause of the aic issue was a depleted battery. this report is being filed as part of a retrospective review of historical records.
13023 | 1220648-2024-24754 | 20925649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925649&pc=OZD | 03/29/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 03/29/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1622563 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) presented with a battery failure alarm. there was no patient involvement when the alarm occurred. | [5 empty] | Data analysis: console was not powered on for almost a year and booted up with a battery failure. it is likely the console was not routinely charged which led to depletion of the batteries. device analysis: the batteries were able to be recharge even though the low voltage safety lockout mechanism had activated triggering the battery failure alarm. the root cause of the battery failure alarm was due to depleted batteries. this report is being filed as part of a retrospective review of historical records.
13024 | 1220648-2024-24751 | 20925710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925710&pc=OZD | 11/26/2024 | 12/23/2024 |  | 12/13/2024 |  |  | 11/26/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532989 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during an impella cp device implant for mechanical circulatory support to a patient with cardiogenic shock from an acute myocardial infarction, the support was prematurely started and resulted in explant and replacement of the device. the patient presented to the emergency department as non-st-elevation myocardial infarction and converted to an st-elevation myocardial infarction. the decision was made to implant the impella cp device for bailout. the impella was backloaded over the guidewire and into the peel-away sheath. inadvertently, support was prematurely started prior to the device being properly positioned and with the guidewire in place. the decision was made to explant the impella cp device and re-start with a new impella cp device to prevent possible device failure and/or complications. the new impella cp device was successfully advanced into the left ventricle and support initiated without incident. percutaneous coronary intervention (pci) commenced, but the patient would ultimately expire due to a dissection occurring during the pci. it was reportedly stated while performing the pci, the physician was able to stent the obtuse marginal. however, when the physician attempted pci on the left main coronary artery, it was dissected. the patient did not recover and met his demise. | [5 empty] | D.4 expiration date and unique identifier (udi) were incorrect on the initial manufacturer device report number 1220648-2024-24751 and were revised. | The decision was made by the user facility to dispose of the pump, and therefore, an evaluation and analysis of the device is not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk.percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13025 | 1220648-2024-24756 | 20925814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925814&pc=OZD | 12/01/2024 | 01/31/2025 |  | 12/13/2024 |  |  | 12/01/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480797 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding. bleeding from the access site and retroperitoneal could not be stopped. as a result, red blood cells and fresh frozen plasma were given. however, clotting became an issue. therefore, the physician made the decision to remove the pump. the procedure was unsuccessful and the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, bleeding from the access site and retroperitoneal as per the clinical description provided.
13026 | 1220648-2024-24757 | 20925860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925860&pc=OZD | 12/01/2024 | 02/14/2025 |  | 12/13/2024 |  |  | 12/02/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574318 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the guidewire had a crease and was bent after opening the packaging. as a result, a new guidewire was used. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. no product/images were returned and logs are not applicable. the root cause of the guidewire damage was not determined since product/images were not returned and insufficient clinical details were provided.
13027 | 1220648-2024-24755 | 20925907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20925907&pc=OZD | 11/28/2024 | 07/07/2025 |  | 12/13/2024 |  |  | 11/28/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568401 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient on impella 5.5 support kept getting runs of arrythmia. the patient had tachyarrhythmia. the impella was repositioned to 31 centimeters. the nurse stated that the impella was hitting the anterior wall. inotropes and vasopressors support was increased. the next day, the patient continued with ectopy and atrial fibrillation. an echocardiogram was done to check placement and distance was 4.7 centimeter. the patient continued to have tachyarrhythmia and an amiodarone bolus was given. the patient remained stable and the procedural outcome was successful and the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported arrhythmia has been completed. the device was returned for evaluation. there were no issues with the returned device. clinical details were not sufficient to determine root cause. therefore, the root cause of the arrhythmia could not be determined. added d9 device return date.
13028 | 1220648-2024-24759 | 20926006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926006&pc=OZD | 04/27/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 04/27/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 130624 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported during an active case automated impella controller (aic) was not detecting the purge cassette. consequently, automated impella controller (aic) was switched to a backup aic. the patient¿s outcome was noted as unknown. however, there was no report or indication of any adverse effects to the unknown age patient because of this event. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs from the reported day of event are consistent with complaint and show excessive number of analytical hierarchy process error messages pertaining to radio frequency identification (rfid) communication (between purge cassette and rfid circuitry attached to the pud board). in the result, purge cassette insertion was not registered by the console, while purge disk and pump were detected, and case start wizard could not proceed without functional purge system. inspection of the console did not reveal any damage to the purge area. however, during testing of the console, the issue was reproduced, and the purge cassette was not being detected. the purge motor shaft did not spin when a known-good purge cassette was positioned close to the purge insert. after replacing the rfid board with a known-good one, the issue was resolved, and purge cassette could be detected. in conclusion, the console was unable to detect purge cassette due to malfunction of the rfid board in the purge assembly. this report is being filed as part of a retrospective review of historical records.
13029 | 1220648-2024-24760 | 20926060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926060&pc=OZD | 12/02/2024 | 01/14/2025 |  | 12/13/2024 |  |  | 12/02/2024 | 12/13/2024 |  |  | 04/30/2025 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353627 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and had a significant drop in hemoglobgin/hematocrit (hemoglobin 11 to 6.3). there were no signs of bleeding from any obvious source. a computed tomography scan was done and a spontaneous mesenteric bleed was found. they were unable to intervene per interventional radiology. heparin was stopped and protamine was given. four units of packed red blood cells were given for initial result. hemoglobin was increasing and stabilizing. the patient is stable and survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. a mesenteric bleed was found on the first day of support. the root cause of the vascular damage was not determined since the product was not returned and insufficient clinical detail was provided.
13030 | 1220648-2024-24758 | 20926121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926121&pc=OZD | 04/27/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 04/27/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the purge cassette would not remain in place. consequently, automated impella controller (aic) was switched to a backup aic with no report of interruption in support and no harm to the unknown age patient. | [5 empty] | The automated impella controller (aic) data logs and console were return for evaluation and analysis. data analysis noted console logs from the reported day of event are consistent with complaint and show, what appears to be, purge cassettes being intermittently detected, and/or perceived by aic software as being frequently removed and reinserted. this issue did not re-occur on the backup console where purge cassette was used. the console was tested and confirmed that the unit was able to correctly and consistently detect purge cassettes. no damage to the purge system components (purge cassette retainer, purge disk retained and flag) was observed. during testing in fast priming mode, with purge door open, the reported issue was reproduced as a known-good purge cassette ¿working its way¿ out of the purge insert was observed. a combination of certain less common usage conditions (purge lid open, console tilted forward, incorrect routing of purge tubing, etc.) was shown to exacerbate the situation. a product history review was completed, and no action was required as a result of this review. in conclusion, possible wear-out of purge cassette retainer and relation between mechanical tolerances of purge insert and purge cassettes that were used, might have also contributed to the issue in the field. this report is being filed as part of a retrospective review of historical records.
13031 | 1220648-2024-24761 | 20926367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926367&pc=OZD | 12/02/2024 | 01/29/2025 |  | 12/13/2024 |  |  | 12/02/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576358 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the following day after an impella cp was inserted, the venous sheath was oozing. sutures were added and dressing was changed. it was further reported that upon removing the impella, bleeding was present at the femoral site. manual pressure and femstop was placed. post impella removal from the right common femoral artery, no distal pulses to popliteal artery were found using doppler. duplex ultrasound was performed and no flow distal to popliteal artery. the plan is to consult vascular surgery for intervention. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by adding sutures and changing dressing.
13032 | 1220648-2024-24762 | 20926565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926565&pc=OZD | 04/27/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 04/27/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller (aic) had a broken purge clip. when the complaint was discovered is unknown. however, there was no report or indication of any adverse effects to a patient because of this event. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs confirmed that there purge pressure low alarms issued, which is likely to be an indication of a broken purge retainer. the console was tested; it was found that purge disk retainer was broken, and the purge disk detect flag was missing. product history review was completed, and no action was required as a result of the review. this report is being filed as part of a retrospective review of historical records
13033 | 1220648-2024-24763 | 20926865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926865&pc=OZD | 05/03/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 05/03/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1175804 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller failure was experienced. no further information was provided including no indication of any adverse effects to a patient because of this event. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs was consistent with the complaint. upon boot up, there were aic alarms for a controller error (loss of communication with the keyboard pca). device analysis was completed, and console interior connections were verified. the 5-inch ribbon cable from the kbd to the 4-in-1 was inspected without severe kinks or discoloration. the cable was found slightly bent but not affecting the function of kbd. console was powered on, and the alarm was not present/ reproduced. the console was then powered cycled 5 times with ample time to boot up fully. normal boots were left on, and the keyboard was tested extensively with no issues occurring. product history review was completed, and no action was required as a result of this review. in conclusion, the root cause of the controller error alarm was most likely due to defective kbd assembly. this report is being filed as part of a retrospective review of historical records. note: initial reporter contact name is unavailable and indicated as unknown to section e1 accordingly.
13034 | 1220648-2024-24764 | 20926895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926895&pc=OZD | 05/04/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 05/04/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 160120 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported during impella mechanical circulatory support, the automated impella controller (aic) displayed a battery communication failure. consequently, the aic was switched to a backup aic with no report of interruption in support and no harm to the unknown age patient. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing), a battery communication failure alarm (smbus communication loss or sda short due to low pack voltage, battery level low (50%) and battery failure, low voltage, (v<12v) alarm issued. the ltpowerdata from the complaint date showed cell imbalance with battery 1 which caused the battery pack¿s internal electronics to shut down. the console was tested, and the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery 1. in conclusion, the root cause of the defective battery was due to internal cell imbalance in battery 1. this report is being filed as part of a retrospective review of historical records
13035 | 1220648-2024-24765 | 20926942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20926942&pc=OZD | 05/05/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 05/05/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 171128 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During training when the automated impella controller (aic) was powered up, a received battery failure error and battery communication failure were experienced. there is no patient associated with this event. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing) and a battery 1 communication failure alarm (smbus communication loss or sda short) due to low pack voltage. there was an unexpected controller shutdown alarm was likely because the user had shut off the console via the battery transport connection switch. the ltpowerdata from the complaint date showed cell imbalance with battery 1 which caused battery pack internal electronics to shut down that battery. the console was tested, and the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure. this report is being filed as part of a retrospective review of historical records.
13036 | 1220648-2024-24766 | 20927287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20927287&pc=OZD | 05/28/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 05/28/2023 | 12/13/2024 |  |  | 01/31/2025 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333966 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support (mcs) and had fluctuated purge pressures between the 200's-500's. the patient presented with 100% left anterior descending artery (lad) occlusion and opened with a balloon when patient's blood pressure dramatically decreased. the patient started on levophed and decision made to place the patient on impella mcs. left femoral artery was accessed, and the impella was placed without complication. following, after start of support, a drug-eluting stent was placed to the proximal lad. post procedures, the patient was transported to the unit on impella support with a plan to rest on impella mcs for the next 24 to 48 hours. however, approximately three days after start of support, purge pressures were noted "consistently" changing and the team frequently had to do a mock bag change for low purge pressure. the pump was weaned and successfully explanted. it was reported the patient was as doing well following the event. | [5 empty] | The pump was not returned for investigation. data logs show low purge pressure alarms only during the purge cassette or the purge bag change procedure. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. this report is being filed as part of a retrospective review of historical records
13037 | 1220648-2024-24767 | 20927443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20927443&pc=OZD | 05/30/2023 | 12/13/2024 |  | 12/13/2024 |  |  | 05/30/2023 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 150112 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller had a battery failure alarm. the complaint was discovered during preparation; however, the was no report of harm to the unknown age patient. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the console logs from the reported day of event are consistent with complaint and showed battery failure alarm (due to transport connection blown/missing or 0v output) presenting upon each boot-up on (b)(6) 2023 and (b)(6) 2023. prior to that, the console was undergoing preventive maintenance procedure at abiomed danvers field service on (b)(4) 2023, and there were no battery failure alarms. the console was tested, there were no irregularities observed, and the aic operated within specification. essential note: each time prior to transport, the console batteries must be disengaged via transport switch. if the switch is not placed in the ¿on¿ position after the transport, then when powered on, the console will present with same symptoms as if the batteries were missing or depleted - which is consistent with this complaint description. this report is being filed as part of a retrospective review of historical records
13038 | 1220648-2024-24768 | 20929019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20929019&pc=OZD | 05/06/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 05/06/2021 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1442007 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported that after successfully placing the impella 5.0 via the patient¿s right axillary artery, the physician was removing a balloon pump and the white impella cable became disconnected from the red pump handle. this disconnection was not immediately apparent to the staff. no alarm sounded from the aic and the display showed flat placement and motor signals as well as ¿pump stopped ¿ alarm condition resolved¿. abiomed staff discovered the disconnection from the pump handle and resolved the pump stoppage within 2-3 minutes. there was no harm or injury and the support proceeded on the pump. the case outcome was the patient was successfully weaned and the impella explanted. | [5 empty] | The field real time logs were reviewed and confirmed an impella stopped alarm (alarm 119) at 10:13am on(b)(6) 2021. further review of the logs revealed that impella 5.0 resumed normal operation at 10:16am. the root cause of the pump stop was determined to be the disconnection of the white cable from the red pump handle during the removal of a balloon pump. no product was provided for evaluation. note: the automated impella controller with serial number (b)(6) associated with this impella pump complaint was reported with manufacturing report number 1220648-2024-22320. this report is being filed as part of a retrospective review of historical records.
13039 | 1220648-2024-24770 | 20929536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20929536&pc=OZD | 08/21/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 08/21/2021 | 12/13/2024 |  |  | 03/31/2023 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022060109 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | Tachycardia | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient had been admitted to the hospital for respiratory failure with ekg changes. the patient was initially implanted with an impella cp but was later then escalated to an impella 5.0. it was reported while on impella 5.0 support, the patient experienced episodes of ventricular tachycardia on (b)(6) 2021 that was resolved with defibrillation on two occasions. the patient also had sputum cultures that day that were positive for gram positive cocci and rods with an elevated white blood count of 15,000 cells/ul. the relationship of the infection to the impella is unlikely, however the cause of the infection was unknown or not reported. it was noted that there were no impella concerns. the patient was moved to the covid unit the same day, on (b)(6) 2021. it was reported the night on (b)(6) 2021 the pump had alarms for high purge pressure and low purge flow. the medical team performed the tissue plasminogen activator lysis protocol to resolve the issue. the following day on (b)(6) 2021, the patient experienced two more episodes of ventricular tachycardia that required defibrillation to resolve. the patient was given lidocaine during this time. the patient's sputum was found to be positive for staphylococcus and was given antibiotics. at this time, the medical team began the discussion of palliative care. on (b)(6) 2021 it was noted that the patient's temperate had spiked to 103f overnight and cultures were taken, but this morning on (b)(6) 2021 am today temp normal. it was noted that the patient was blinking, but not responding to verbal commands. on (b)(6) 2024, it was noted the patient had a fever of 103f that spiked overnight and the patient had seizures as a result. cultures and sent to lab. lidocaine was stopped this morning and patient has gone into an agonal rhythm. the family of the patient made the decision to withdraw care, therapy was discontinued and the patient expired. there was no report of association between the impella and the fever and sepsis, however it is unknown and the impella cannot be ruled out as contributing to the patient's sepsis. with impella related adverse events and malfunction occurring throughout the patient's therapy, impella cannot be ruled out as a contributing factor to the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13040 | 1220648-2024-24769 | 20929624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20929624&pc=OZD | 06/01/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 06/01/2021 | 12/13/2024 |  |  | 01/31/2023 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022051472 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the physician attempted to deliver the impella 5.0 multiple times but was unable to cross into the aorta. the decision was made to move forward with an impella cp instead which was delivered with no issues. additional information indicated that nothing remarkable was noted regarding the patient¿s anatomy. | [5 empty] | The reported delivery issues were determined to most likely be related to patient condition. the returned pump was inspected, and no abnormalities were observed. this report is being filed as part of a retrospective review of historical records.
13041 | 1220648-2024-24771 | 20930066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20930066&pc=OZD | 06/07/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 06/07/2021 | 12/13/2024 |  |  | 07/31/2022 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1502978 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. the site reported that impella 5.0 would not make the turn into the aorta from the subclavian artery during insertion. the cannula reportedly kept collapsing when the physician attempted to advance. the physician stated that the patient¿s vessel was too small. insertion was aborted and an impella cp was placed instead. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation; however, there were sufficient clinical details to determine a root cause. the reported delivery issues were determined to be related to the patient¿s anatomy. this report is being filed as part of a retrospective review of historical records.
13042 | 1220648-2024-24772 | 20930139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20930139&pc=OZD | 05/12/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 05/12/2021 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1469637 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was being implanted with an impella 5.0 for mechanical circulatory support. the patient had a dissected aorta prior to impella insertion, that the physicians were attempting to repair when the decision was made to implant an impella for support. it was reported that during the insertion of the impella, the physician was unable to get the guidewire across the aortic valve. the implant was aborted, and a short time later the patient expired. although the impella catheters were not inserted, the guidewire was inserted, and therefore there is an identifiable connection between the impella device and the event. however, the patient's presenting condition may be more likely have impacted the event. | [5 empty] | D4: udi and expiration date are unknown. h4: device manufacture date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13043 | 1220648-2024-24773 | 20930342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20930342&pc=OZD | 02/18/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 02/18/2021 | 12/13/2024 |  |  | 12/31/2022 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022026925 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support after an aborted impella 5.5 insertion. it was reported while on impella cp support, the patient experienced a gastrointestinal bleeding and as a result systemic heparin was discontinued. on the last day of support, on (b)(6) 2021, the pump experienced a sudden stop. the medical team attempted to restart the pump three times, all unsuccessful. the automated impella controller (aic) was exchanged for a back up aic, the issue was not resolved. it was reported the patient expired shortly after the pump stop. it was later reported that up until the time of the pump stop, the patient had been hemodynamically stable. the patient decompensated after the pump stop and died about 45 minutes later. it was reported that, based on this information, the patient's death was likely due to pump stoppage. medical intervention at the time of the pump stop was unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13044 | 1220648-2024-24774 | 20930441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20930441&pc=OZD | 12/13/2024 | 12/13/2024 |  | 12/13/2024 |  |  | 12/13/2024 | 12/13/2024 |  |  | 07/31/2021 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the position detection signal was not stable after priming. an alarm appeared and the physician was asked to try and insert the pump. even though the pump started, the position detection signal was not stable. the alarm continued. as a result, the pump was replaced. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. review of the provided logs showed that the placement signal was unreliable. a 'placement signal not reliable' alarm was triggered at 16:58 on 08jun2021. apart from the placement signal, the pump ran as expected without issues. analysis of the returned pump found open lines of the pressure sensor and voltage lines. no kinks or cuts in the catheter were observed. the pump was run in the lab, and the 'placement signal not reliable' alarm was triggered. the alarm did not resolve, despite manually re-zeroing of the sensor from the console. the root cause of the unreliable placement signal was an open line. this report is being filed as part of a retrospective review of historical records.
13045 | 1220648-2024-24775 | 20931126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931126&pc=OZD | 06/25/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 06/25/2021 | 12/13/2024 |  |  | 10/31/2022 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024095 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was being implanted with a new impella 5.0 in the axillary artery for longer support and ambulation. the patient had an existing impella cp that had been placed in the right femoral artery. it was reported that the impella 5.0 was not able to be advanced past the distal innominate (brachiocephalic) artery due to patient anatomy. impella 5.0 was coated in a lubricant (viperslide) and reinserted again unsuccessfully. the decision was made to abort the 5.0 insertion and insert an impella cp instead. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13046 | 1220648-2024-24776 | 20931289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931289&pc=OZD | 06/28/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 07/12/2021 | 12/13/2024 |  |  | 07/31/2022 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024960 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Unspecified Infection | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0 support, the patient experienced heparin induced thrombocytopenia. as a result, heparin was removed from the purge solution. i was also reported on (b)(6) 2021 that the purge pressure was high. the cassette was replaced, and the catheter was checked for a kink, however no kink was found and the cassette did not resolve the issue. on (b)(6) 2021 it was reported the patient had an infection at the impella access site which was resolved with "debris, cleaning, and additional treatment". at the conclusion of the case, the patient was successfully weaned and the impella was explanted. | [5 empty] | The returned purge cassette and pump were connected to a console and high purge pressure reproduced. the pump was visually inspected and biomaterial was observed in the purge gap. this report is being filed as part of a retrospective review of historical records.
13047 | 1220648-2024-24777 | 20931303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931303&pc=OZD | 12/13/2024 | 12/13/2024 |  | 12/13/2024 |  |  | 12/13/2024 | 12/13/2024 |  |  | 01/31/2023 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022052315 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the clinical team reported that recurrent "high purge pressure" alarms with low purge flow occurred during impella support. attempting to troubleshoot the issue, the team decreased the pump speed to p-0 and then immediately returned it to p-7, which successfully improved the purge flow. however, the purge flow rate trended downwards over the next 24 hours, which led to the use of the same intervention again. it was noted that the purge flow never exceeded 5 ml/hr throughout support. activated clotting time (act) was asserted to be in normal therapeutic range. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. data logs of the incorrect pump were sent. additionally, only the cannula with the motor housing was returned for investigation; the remainder of the impella was cut off at the catheter. a small amount of biomaterial was observed around part of the purge gap; otherwise, no damage or abnormalities were found on the returned product. however, the nature of the returned product was insufficient to definitively determine the root cause. therefore, the cause of the high purge pressure was unable to be determined. this report is being filed as part of a retrospective review of historical records.
13048 | 1220648-2024-24778 | 20931494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931494&pc=OZD | 12/13/2024 | 03/06/2025 |  | 12/13/2024 |  |  | 12/13/2024 | 12/13/2024 |  |  | 07/31/2022 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021019637 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Death | Yes | No | Restricted Flow rate | [6 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a low pump flow issue involving an impella 5.0 pump. during patient support, it was presumed that the cause of suction was a decrease in right heart function and a lack of left ventricular filling. since it was necessary to reintroduce venoarterial extracorporeal membrane oxygenation (va ecmo) in order to recover from this condition, ic was given to the patient's family. additional information noted care was withdrawn and the patient expired. the attending physician commented that the cause of death was ischemic heart failure due to myocardial infarction. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the product and event logs were not returned for review. based on clinical description, the root cause of low flow was most likely patient condition related as the physician confirmed the cause of the issue was due to failure of the right heart function and a lack of left ventricular filling. this report is being filed as part of a retrospective review of historical records.
13049 | 1220648-2024-24779 | 20931665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931665&pc=OZD | 06/18/2021 | 02/07/2025 |  | 12/13/2024 |  |  | 12/13/2024 | 12/13/2024 |  |  | 03/31/2023 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022059107 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was under do not attempt resuscitation (dnar) status with a cerebral herniation and the policy was not to provide aggressive treatment. it was decided to continue the support as long as venoarterial extracorporeal membrane oxygenation and the impella were running, but the flow gradually could not be maintained, pressure could not be generated, and the patient expired. per the surgeon, the death was not related to the use of impella. further information was unable to be obtained as follow-up was not possible. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined b3 date of event was inadvertently entered incorrectly on manufacturer device report 1220648-2024-24779.
13050 | 1220648-2024-24780 | 20931758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931758&pc=OZD | 05/14/2021 | 03/10/2025 |  | 12/13/2024 |  |  | 05/17/2021 | 12/13/2024 |  |  | 05/31/2021 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1428037 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was escalated from an impella cp and implanted with an impella 5.0 device for increased mechanical circulatory support. it was reported while on impella 5.0 support, the patient experienced hemorrhagic cerebrovascular accident. it was reported at the time, the patient had been heparinized for impella use, and the patient's activated clotting time was extended. after the stroke was identified by computed tomography, the patient's care was withdrawn and the patient expired while on impella 5.0 support. the physician determined that due to the heparin use, the relationship of impella use can not be dissociated with the stroke and therefore the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed that the pump passed all post sterile inspection checks no device was received for evaluation. as sufficient clinical information was not provided, the root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
13051 | 1220648-2024-24781 | 20931859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931859&pc=OZD | 05/17/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 05/17/2021 | 12/13/2024 |  |  | 11/30/2021 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037322 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the impella 5.0 was unable to be inserted through the subclavian artery. the vessel was measured upon initial cutdown and was found to be 8mm, so the 5.0 insertion was attempted. the physician felt resistance and suspected stenosis so it was decided to place a cp instead. there was no patient harm reported. | [5 empty] | The root cause of the delivery issues was determined to be patient condition as stenosis was suspected. no product or data was returned for this investigation. this report is being filed as part of a retrospective review of historical records.
13052 | 1220648-2024-24782 | 20931909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931909&pc=PYX | 12/03/2024 | 12/13/2024 |  | 12/13/2024 |  |  | 12/03/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025564021 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella rp flex for a patient admitted with complex underlying conditions of pulmonary embolism, quadriplegia, hypoxia, pulseless electrical activity arrest, and a cardiac arrest. the team placed the impella rp flex over the right ventricular assist device. the team noted that the pump had a malfunction in that the tuohy would not lock to prevent pump malpositioning. the pump remains on for support despite the locking malfunction. the pump has been sutured to prevent malpositioning. the decision was made to explant the pump two days later. the team was happy with the outcome and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13053 | 1220648-2024-24783 | 20931932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931932&pc=OZD | 12/02/2024 | 02/07/2025 |  | 12/13/2024 |  |  | 12/03/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570285 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support, and during support, there was air in the purge system. the staff deaired the purge system and purge fluid came out of the yellow lure lock. the alarm would not clear so the staff replaced the pump. the patient is stable. | [5 empty] | The investigation has been completed. the pump and purge cassette were returned for product evaluation. there were no visual abnormalities with the pump. the purge cassette was loaded onto the console and the de-airing and priming procedure was performed and a leak was noticed a cartridge cap leak from the thf bond of the gen1 purge cassette. the "air in purge" alarm appaeared intermittently when the purge was loaded but the purge ran through the pump with no issues. the logs for the case were provided. the data logs and impella contorller logs confirmed 'air in purge system' alarms. the root cause of the purge leak - cassette was determined to be due to a bond failure of the thf bond of the gen1 purge cassette. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
13054 | 1220648-2024-24784 | 20931946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931946&pc=OZD | 05/29/2021 | 12/13/2024 |  | 12/13/2024 |  |  | 05/29/2021 | 12/13/2024 |  |  | 03/31/2023 |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022064671 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was being implanted with an impella 5.0 for mechanical circulatory support. it was reported the impella 5.0 was unable to be inserted. according to the physicians the patient's vessel size was too small for the 5.0 to advance. an impella cp was used instead and the case continued as planned. there was no reported patient harm. | [5 empty] | The root cause of the delivery issues was determined to be patient condition as the physician stated that patient anatomy was the reason the impella 5.0 was unable to advance. no product or data was returned for this investigation. this report is being filed as part of a retrospective review of historical records.
13055 | 1220648-2024-24785 | 20931968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20931968&pc=OZD | 12/03/2024 | 01/20/2025 |  | 12/13/2024 |  |  | 12/03/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449553 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis one day into impella cp support. it was discovered via echocardiography that the pump was deep and that the pigtail was at the apex. rather than reposition the device, the decision was made to explant the pump. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the positioning issue and hemolysis has been completed. data logs were reviewed which showed pump ran without issue during support. product was not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of positioning issue was most likely patient movement.
13056 | 1220648-2024-24786 | 20932023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20932023&pc=OZD | 12/03/2024 | 01/20/2025 |  | 12/13/2024 |  |  | 12/03/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569553 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, purge system blocked alarm was present. the nurse traced lines and there were no kinks in tubing. tissue plasminogen activator (tpa) protocol was started. tpa purge solution was infusing, and purge alarms were on-going. the impella pump stopped and multiple attempts were made to restart the pump. the patient remained hemodynamically stable and went to the operating room and had the impella removed. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion . ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the high purge pressure and the pump stop has been completed. clinical details report that the high purge pressure alarm was troubleshot with tpa infusion into the purge, but shortly after the pump stopped anyways. the patient's activated clotting time (acts) at the time were unknown, there were no kinks in the purge line, and the purge cassette wasn't replaced. since data logs were not available for review and product was not returned for investigation, the root cause of the high purge pressure could not be determined. clincical details states that shortly after high purge pressure alarms began to trigger, the pump stopped and could not be restarted. no further clinical details were provided. since data logs were not available for review and product was not returned for investigation, the root cause of the pump stop could not be determined.
13057 | 1220648-2024-24787 | 20932054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20932054&pc=PYX | 12/06/2024 | 09/22/2025 |  | 12/13/2024 |  |  | 12/06/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540879 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Insufficient Information | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was implanted with an impella rp flex pump for mechanical circulatory support and had low flows. the patient status post cardiovascular artery bypass graft surgery with mechanical aortic valve replacement and placement on intra-aortic balloon pump support was brought to the catheterization lab for impella rp flex support. however, approximately eight to nine days after start of support, flows were found to be lower than expected at performance level p-5. heparin dosage was increased in response to the low flow. support continued approximately for an additional four days. | Additional information was received. the complainant reported that an implanted impella rp pump experienced a loss of the placement signal and flow display. positioning was subsequently confirmed via chest x-ray. the decision was made to continue support with the implanted pump despite the intermittent placement signal. the device was successfully weaned and explanted. |  |  |  |  | The investigation of low pump flow has been completed. the pump was returned for investigation. data logs were not provided for this case run. the returned product had biomaterial wrapped around the impeller, which was flushed out during the pump testing in a bucket of water. no issues were noted with the optical signal parameters, and the pump was able to start in the bucket of water. no physical damage was observed on the pump or dp sensor. the pump was sent for a flow loop test, where it failed due to sensor drift after three hours of operation, with loss of placement signal data and pump flow. the optical sensor failure mode was changed to a placement signal issue since the dp sensor was found to be malfunctioning and causing the psnr issue. additionally, the low pump flow was removed from the investigated failure modes since the dp sensor issue was disturbing the pump flow calculation during the case run. as per the clinical, loss of the placement signal and the pump flow reported with peak signal-to-noise ratio alarm and these were reappeared after one day. the cause of the placement signal issue was sensor drift, based on returned product investigation. pump flow calculation was based on the placement signal, so in this case, the pump flow was affected due to sensor drift. d.9 date device returned/information was added. d.6b revised explant date as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-24787. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B.5. updated with additional information. h.11. type of reportable event has been corrected from serious injury to malfunction. h.6. medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-24787.
13058 | 1220648-2024-24788 | 20932079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20932079&pc=OZD | 12/03/2024 | 01/27/2025 |  | 12/13/2024 |  |  | 12/03/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568020 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while moving the patient from the bed to the operating table, the impella pump stopped. the impella could not be restarted even with a new console. the doctor pulled the impella back out of the left ventricle and the patient was prepped for explant. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock: section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation of pump stop has been completed. the pump was returned for investigation. no physical damage was noted on the returned pump. it failed electrical testing, and further investigation revealed that all three motor current lines were stretched and broke inside the red handle on the catheter side, which caused the pump to stop. the data log showed the pump stopped with a motor current spike to 3000 and alarm 115 pump stop at 259 hours of run. according to the clinical team, the pump stopped during patient transport and the pump was unable to start afterward. the root cause of the pump stop issue was an open electrical line inside red handle during patient transfer.
13059 | 1220648-2024-24789 | 20932106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20932106&pc=OZD | 11/28/2024 | 01/15/2025 |  | 12/13/2024 |  |  | 11/28/2024 | 12/13/2024 |  |  |  |  | 12/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569900 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding. post impella cp implant same day in the evening, the patient was oozing blood around the blue hub and suture site in right groin. partial thromboplastin time (ptt) was drawn and over 250. the physician decided to pause heparin drip two hours to allow ptt to drop. light pressure was held at the site, and the oozing stopped. the dressing was changed, and a small hematoma was noted at the site. the next morning the dressing was clean, dry and intact with some bruising at the access site. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was returned for evaluation. upon evaluation, there was no issues were found with the returned product. clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since the hematoma resolved after the heparin was held. added d9 device returned for evaluation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13060 | 1220648-2024-24715 | 20912127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912127&pc=OZD | 11/27/2024 | 01/09/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553857 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was inserted during cross-clamp into a 10 millimeter graft and aorta under negative pressure from the aortic root vent. once in the aorta, the pump was placed via direct visualization through aortatomy. incorrect flows coupled with non-pulsatile motor current along with a non-decompressed ventricle led the team to explant the device for failure. the patient remained hemodynamically stable on minimal inotropes throughout triage period. no patient harm was reported. | [5 empty] | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the data logs showed pump position was in aortic area for almost entire of the case that triggered suction and pump position in aorta alarms. no other issue was found on the log. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of low flow was most likely due to pump was not in the proper position.
13061 | 1220648-2024-24716 | 20912335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912335&pc=OZD | 11/28/2024 | 01/13/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had numerous rounds of ventricular tachycardia/ventricular fibrillation. the patient was defibrillated and was given amiodarone and lidocaine. the family decided to withdraw care and the patient expired after being on impella support for 170.46 hours. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13062 | 1220648-2024-24717 | 20912635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912635&pc=OZD | 11/28/2024 | 01/15/2025 |  | 12/12/2024 |  |  | 11/28/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562237 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 was not implanted due to a catheter kink. due to excessive force, the catheter kinked right next to the motor area. a third pump was needed. the patient expired on support (see patient identifier (b)(6)). | [5 empty] | The investigation for the product damage has been completed. the pump was returned and was visually inspected. the catheter kinked near motor housing was observed. no other abnormality was observed. the root cause of the product damage (cannula kink) was use related due to excessive force. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13063 | 1220648-2024-24721 | 20912830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912830&pc=OZD | 11/27/2024 | 01/10/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the impella had low pump flows. the flows were less than 0.5 and there was suctioning. the pump was repositioned with no improvement. the impella was stopped and restarted with no improvement. the decision was made to explant and replace the pump. the new pump had no issues and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs confirmed low pump flow when pump started and remained to the rest of the case. product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
13064 | 1220648-2024-24719 | 20912852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912852&pc=OZD | 11/26/2024 | 01/13/2025 |  | 12/12/2024 |  |  | 11/26/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566631 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the oozing post removal of the impella cp. a constant ooze was noted post closure and despite manual pressure being held, after over 2 hours the groin ooze had not slowed down. as a result, a balloon tamponade was used to achieve hemostasis. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13065 | 1220648-2024-24720 | 20912911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912911&pc=OZD | 11/26/2024 | 01/16/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572104 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had a gastrointestinal (gi) bleed, noted by the bright red stool, and that there was saturated flow. the gi bleed was treated by removing heparin. the patient received two units of packed red blood cells. the saturated flows were noted by p-level 7 flows at 4.0 liters per minute and elevated left ventricle waveforms. as a result, the decision was made to replace the pump. the patient survived the event. | [5 empty] | The investigation of saturated flow and vascular damage - peripheral vessel has been completed. the returned product was inspected and there was an abnormally large purge gap. no data logs were returned for analysis. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided. the root cause of the saturated flow was not determined as no data logs were returned for analysis. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13066 | 1220648-2024-24718 | 20912922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20912922&pc=OZD | 11/16/2024 | 01/13/2025 |  | 12/12/2024 |  |  | 11/16/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568135 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Mitral Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia (vt). the patient was converted out of vt several times. the impella p level was lowered. the patient had arrythmia issues prior to the impella implant, however it is unknown if the impella might have contributed to the patient's vt. furthermore, there was concern that the pump interacted with the mitral and caused flash severe mitral regurgitation. the impella was repositioned successfully. support continued until the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of ventricular tachycardia/arrhythmia and heart valve damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: impella catheter in papillary muscle ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: warnings & cautions: warnings "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: correct impella catheter position ¿catheter angled toward the left ventricular apex away from the heart wall and not curled up or blocking the mitral valve¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24718 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 a portion of the additional manufacturer narrative instructions for use were inadvertently omitted from initial manufacturer device report number 1220648-2024-24718. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13067 | 1220648-2024-24723 | 20913037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20913037&pc=OZD | 11/25/2024 | 01/15/2025 |  | 12/12/2024 |  |  | 11/28/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528392 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | Life Threatening | Disability | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced gastrointestinal bleed and would later expire from multiorgan failure. very little information was provided. however, details available are the following: ten days prior to the event the impella cp device was implanted through the femoral artery (side is unknown). the patient underwent extracorporeal membrane oxygenation (ecmo) cannulation and was on both impella and ecmo support. the patient experienced gastrointestinal bleeding. bleeding occurred from the time of management with ecmo, there was no bleeding at the access site on the impella side, and the amount of bleeding was thought to be from the digestive tract. the physician commented that the bleeding was inevitable because ecmo was inserted, and it was not impella's fault. it was noted that the relationship between the impella cp device and bleeding is unknown due to the use of both impella and ecmo support. the patient would eventually expire noted due to the inability to control bleeding from multiple organs and removal of the impella. the treating physician commented that the patient's demise is not related to impella cp device; there were no problems with the impella operation. however, given that there is also associated bleeding from the removal of the impella cp device, there could be an associated factor with the patient's demise. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24723 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 coded 4641 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24723 and a new code was added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
13068 | 1220648-2024-24724 | 20913040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20913040&pc=OZD | 11/28/2024 | 01/14/2025 |  | 12/12/2024 |  |  | 11/28/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504279 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was product damage. due to excessive force during insertion, the catheter kinked next to motor. as a result the pump was removed and exchanged for a new one. the patient¿s outcome is unknown. | [5 empty] | The investigation of product damage (cannula kink) has been completed. the returned complaint pump was visually inspected and a catheter kink near motor housing was observed. no other abnormality observed. the cause of the product damage (cannula kink) was use related due to excessive force. e4 should have been left blank on manufacturer device report 1220648-2024-24724. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13069 | 1220648-2024-24722 | 20913062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20913062&pc=OZD | 11/27/2024 | 01/14/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571938 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient found right lower extremity thrombus which required thrombectomy on (b)(6) 2024. wound vac applied on site after thrombectomy procedure. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13070 | 1220648-2024-24728 | 20914545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20914545&pc=OZD | 03/07/2023 | 12/12/2024 |  | 12/12/2024 |  |  | 03/07/2023 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1608973 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported upon boot-up of the automated impella controller (aic), there was an alarm for battery failure despite being plugged into wall outlet for hours and verifying the battey transport switch was in the correct position. the impella cp support was successfully started with the battery failure alarms. the medical staff confirmed the battery continued to charge while plugged into an outlet. there was no patient harm reported, the patient remained on impella cp support until successfully weaned and explanted. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show battery failure alarm (#360) as well as battery level low (#23) and battery failure due to low voltage (#350) upon aic boot up and during aic operation (support with pump 422495 was successfully concluded, despite on-going battery alarms). battery and power data embedded in long term logs confirmed cell imbalance of battery 1 (cell 1 failure), which is a known pattern failure affecting certain number of battery packs. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
13071 | 1220648-2024-24727 | 20914640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20914640&pc=PYX | 11/17/2024 | 02/03/2025 |  | 12/12/2024 |  |  | 11/25/2024 | 12/12/2024 |  |  | 08/31/2025 |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025453551 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  | White | Death | Yes | Yes | Fracture | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella rp implanted and during the implant, the guidewire either broke off or twisted. the implant was still successful. the patient was on impella support for 133.00 hours before the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of delivery issues has been completed. the product was returned for investigation. data log review was not required for this case. no physical damage was noted on the returned pump. the returned guidewire was non-abiomed, which was confirmed by the product dimensions. the coil at the tip of the returned guidewire was frayed and damaged. as per the given clinical, guidewire was broken or twisted during impella rp delivery, but the pump implant procedure was successful. the cause of the delivery issue was most likely use related to a non-abiomed guidewire. b.3 revised date of event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-24727.
13072 | 1220648-2024-24726 | 20914662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20914662&pc=OZD | 02/22/2023 | 12/12/2024 |  | 12/12/2024 |  |  | 02/22/2023 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398270 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | An automated impella controller (aic) was returned for annual preventative maintenance. during inspection of the aic unit, it was discovered that the front and rear housing was broken. there was no patient involvement. | [5 empty] | Data analysis: event logs were reviewed but not related to the failure mode. device analysis: field service inspected the console exterior and interior to confirm that the cracks did not introduce more damage or ingress to internal components. the cause of the issue was most likely due to mishandling during use/ transition. this report is being filed as part of a retrospective review of historical records.
13073 | 1220648-2024-24730 | 20914967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20914967&pc=OZD | 11/27/2024 | 10/09/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570931 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Intracranial Hemorrhage | Renal Failure | [10 empty] | Other | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella cp support and during support, the patient had hemolysis. the patient was put on continuous renal replacement therapy (crrt) and support was continued. additionally, while on support a stable subarachnoid hemorrhage with adjacent edema (contusion, stable) and a four millimeter thick subdural hematoma was found. the staff did not believe this to be impella related. however, the cause of the hemorrhage is unknown. the procedural outcome was the patient survived the surgery. | The complainant also reported that there were positioning issues. eventually, the pump was repositioned, and the placement signal alarms that were previously occurring were resolved. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of hemolysis, renal failure, and stroke/cerebrovascular accident (cva) has been completed. impella cp pump was returned for investigation. the pump was visually inspected and no biomaterial, broken blades, or other physical abnormalities were observed. no issues with the pump delo joint or epo-tek was visually seen. the pump was connected to the console and the optical bench 5s parameters were verified. all optical parameters were in normal range within specification. the light continuity test was done and the light passed through the sensor face at pump outflow. the pump was placed in a bucket of water and ran at p-9. no alarms relative to placement signal was seen and the pump ran within specification. no other issues were found with return product. data logs were reviewed and a few suction alarms and several placement signal low alarms were observed. a motor current spike was found with no major impact to the purge, pump flow parameters. the motor current spike was not correlative to hemolysis event as hemolysis was reported a few hours prior to the motor current spike. the optical 5s parameters were all found to be in normal range with no major abnormalities. signal-to-noise ratio and contrast were slightly dropping intermittently. all other light, lamp and gain were stable. the placement signal trend was noisy, low and declining since the start of the case. as per clinical, the patient was very critically ill with multiple comorbidities. the doctor suggested the patient was not a candidate for escalation or advanced therapies due to advanced age and other underlying conditions. the cause of the hemolysis issue was most likely patient condition related based on return product analysis, log and clinical review. the cause of the renal failure issue was not determined due to insufficient clinical information. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. d.9 date device returned/information was added. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the optical signal issue was most likely patient condition related based on return product analysis, log and clinical review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated.
13074 | 1220648-2024-24731 | 20915056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20915056&pc=OZD | 11/29/2024 | 01/09/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542819 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant and due to patient anatomy, the implant was aborted. no known patient harm or injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to the patient's anatomy.
13075 | 1220648-2024-24729 | 20915082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20915082&pc=OZD | 04/16/2023 | 12/12/2024 |  | 12/12/2024 |  |  | 04/19/2023 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1608973 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an unknown age patient [information unavailable] with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support, and on the automated impella controller (aic) there was an impella defective alarm. consequently, the impella cp pump was switched to a backup aic with no report of interruption in support and no harm to the patient. | [5 empty] | The automated impella controller (a(c) data logs and console were returned for evaluation and analysis. data analysis of the aic logs noted on console there was an impella defective alarm that was issued persistently (multiple attempts) due to issues reading the first section of the pump eeprom. the pump was switched to a backup controller with no issues. logs from the backup controller were not obtained to confirm whether there were any underlying issues reading the eeprom. the device analysis noted the failure mode was not reproduced. a known good impella cp was connected to this console and it read without any issues. a product history review was completed, and no action was required as a result of this review. the root cause of the impella defective alarm is most likely due to an intermittent malfunction of the impellatronic given the log characteristics. this report is being filed as part of a retrospective review of historical records.
13076 | 1220648-2024-24732 | 20915271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20915271&pc=OZD | 11/27/2024 | 01/16/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556224 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient underwent a 6-hour ventricular tachycardia (vt) ablation. in vt, the patient had no pulsatility with mean arterial pressure around 65 throughout the case and eight plus vt morphologies. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d2 common device name was reported incorrectly on manufacturer device report 1220648-2024-24732.
13077 | 1220648-2024-24733 | 20915376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20915376&pc=OZD | 11/30/2024 | 01/20/2025 |  | 12/12/2024 |  |  | 11/30/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump triggered a purge system blocked alarm during patient support. the purge cassette was replaced in an attempt to troubleshoot the failure but the issue remained. the site opted not to initiate a tissue plasminogen activator protocol and instead moved to replace the device. the pump was removed and a new impella cp device was placed for ongoing support. there was no patient harm. the automated impella controller was returned for evaluation. | [5 empty] | The investigation of automated impella controller (aic) - purge issue - other has been completed. the issue was most likely caused by (b)(4) as the failure persisted over three purge cassettes and two pumps and was resolved by switching to a new console. however, the failure mode's root cause could not be determined since it was not reproduced. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13078 | 1220648-2024-24725 | 20915478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20915478&pc=OZD | 11/23/2024 | 01/08/2025 |  | 12/12/2024 |  |  | 11/23/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572849 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support had hematoma. hemoglobin was low. heparin held. the hematoma appeared to be improving. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to anticoagulation management since the hematoma resolved after the heparin was held. e4 should have been left blank on manufacturer device report 1220648-2024-24725.
13079 | 1220648-2024-24735 | 20916096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20916096&pc=OZD | 03/05/2023 | 12/12/2024 |  | 12/12/2024 |  |  | 03/08/2023 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127919 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during use with a patient, an automated impella controller (aic) screen failed and was only displaying "white". the aic was exchanged for a back up unit and the issue was resolved. it was noted that the patient was awake, oriented and responsive when the event occurred. no patient harm was reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that on (b)(6) 2023 imc logs end abruptly, only to resume on (b)(6) 2023 when the aic was already at aachen field service (fs). while the last entry in event log is timestamped 09:00:04, the real time log data indicates that the console was still running the pump at least until 09:02:32 (impellatronic board is designed to be capable of running the pump independently of 4-in-1/epc, for a few minutes even after catastrophic failure of the epc). logs from (b)(6) 2023 are consistent with observations made by fs personnel that the console initially crashed and rebooted after first power-on and indicated batteries at 0%. furthermore, the batteries were found to be fully depleted (as confirmed by batttest data), which would be consistent with hospital staff not performing emergency shutdown (or not being able to, if the aic was completely unresponsive) but rather waiting for the aic to shut itself down upon fully depleting batteries. long term event log data from the day of reported event did not show any evidence of prior power or battery issues. the core dump file associated with imcgui process, referenced by imc log file on (b)(6) 2023 was associated with unsuccessful power-up on the same day, rather than the prior issue in the field on (b)(6) 2023. during testing in aachen, the issue was not reproduced, and batteries were able to be recharged. analysis of available data, combined with the fact that console operated within specification during testing, suggests that the original issue (console ic2241 becoming unresponsive and unable to shut down) was most likely the result of either software or hardware glitch that caused malfunction of the failure notification and recovery mechanism (software and hardware watchdogs), but the exact root cause could not be determined. as the most likely component to have affected this functionality, the 4-in-1 assembly (including epc and cf cards with software) is recommended to be replaced as precaution. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
13080 | 1220648-2024-24734 | 20917379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20917379&pc=OZD | 11/27/2024 | 01/08/2025 |  | 12/12/2024 |  |  | 11/27/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523938 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient with cardiac and systemic history. the impella cp was placed but, there were persistent signs of hemolysis. the labs were elevated and hemolyzing with urine color change. the pump was repositioned but issue remained for a few days then hemolysis improved. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the product was not returned for review. based on clinical description and data review, the root cause of hemolysis was most likely due to pump was not in the proper position as hemolysis was improved post repositioning of the pump. e4 should have been left blank on manufacturer device report 1220648-2024-24734.
13081 | 1220648-2024-24737 | 20918718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20918718&pc=OZD | 11/29/2024 | 01/15/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566807 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. leading up to insertion, it was verified via aortic angiogram that a large calcified boulder was present in the mid aorta. the pump was then inserted but would not advance due to that patient's anatomy. the implant was aborted with no documented patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to be inserted due to large calcified boulder in the mid aorta.
13082 | 1220648-2024-24741 | 20919054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20919054&pc=OZD | 03/15/2023 | 12/12/2024 |  | 12/12/2024 |  |  | 03/15/2023 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103170 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 support, the automated impella controller (aic) presented with a controller failure alarm. the aic was exchanged for a backup unit and the issue was resolved. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that after 19 days and 18 hours of support with pump 430446, console suddenly presented with controller failure alarm (#1027, seemingly due to pbm1 software corruption), followed by battery level low (#23) alarm. after console was received in danvers, it would not boot-up properly due to syscall error (fw mismatch detected). by manually initializing pbm communication to pbm1 and pbm2 via telnet commands in service mode, it¿s been observed that hw revision bit for pbm1 had incorrect value: sau_powermod_1 hw revision: -1 version: e 0 pbm_cbh_2. sau_powermod_2 hw revision: 0 version: e 0 pbm_cbh_2. at the same time, pbm firmware content extracted from both controllers on the board was identical, and consistent with the programming code (pbm_cbh_2). after resetting the hw revision bit, the console was able to properly boot-up, and no controller alarms were present (and battery level was correctly reported). although the hw revision bit maintained its value after several aic re-boots, it is recommended to replace pbm as a precaution to minimize chances of the bit resetting itself spontaneously. the cause of the aic issue was a firmware corruption. this report is being filed as part of a retrospective review of historical records.
13083 | 1220648-2024-24739 | 20919092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20919092&pc=OZD | 11/29/2024 | 01/30/2025 |  | 12/12/2024 |  |  | 11/30/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569477 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, there was a report of high purge pressure. the decision was made to start lysis and the issue resolved. the patient survived the event. | [5 empty] | The investigation of high purge pressure has been completed. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. e4 should have been left blank on manufacturer device report 1220648-2024-24739. g2 report source; foreign was missing from manufacturer device report 1220648-2024-24739. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
13084 | 1220648-2024-24738 | 20920559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20920559&pc=OZD | 11/29/2024 | 01/16/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544916 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed bleeding at the impella site with hematoma. the patient complained of pain at site. heparin was held. computed tomography scan revealed 2 cm enhancing lesion. labs are stable. bleeding was monitored and now stable. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13085 | 1220648-2024-24740 | 20920851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20920851&pc=OZD | 11/29/2024 | 01/17/2025 |  | 12/12/2024 |  |  | 11/29/2024 | 12/12/2024 |  |  |  |  | 12/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510420 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Eu complainant reported that while implanting an impella 5.5 and explanting an impella cp, the guidewire entered the inlet cage of the impella cp which was pulled back into the aortic arch. the wire was pulled back, and the wire got stuck in the inlet and kinked. the patient remains on impella support. | Additional information was received that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information was received and the following were updated on follow up 1 manufacturer device report 1220648-2024-24740. b.1 updated to add adverse event along with product problem. b.2 death was selected and date of death was added. b.5 updated with patient outcome. d.6b explant date was added. d.9 device returned date/information was added. h.1 type of reportable event was updated to death. h.3 device evaluation anticipated, but not yet begun was selected. h.6 health effect - impact code for death was added. d.4 revised expiration date and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report 1220648-2024-24740. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24740 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The investigation of product damage (guidewire damage - great vessel) has been completed. the returned guidewire was inspected and kink was confirmed upon inspection. the clinical details mentioned the issue occurred when the pump was pulled. the root cause of the product damage (guidewire damage - great vessel) was determined to be use-related.
13086 | 1220648-2024-24692 | 20901619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20901619&pc=OZD | 11/26/2024 | 01/13/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient was taken to the operating room for impella cp removal as it had been in for longer than a week. an intra-ortic balloon pump was placed. the patient survived the explant. | The complainant reported that a patient on impella cp support had hematuria. urine output was dark red. possible escalation of support was discussed. it is unknown how the hemolysis was resolved and no further information was given for the issue. the patient is in critical condition. |  |  |  |  | The investigation for the hemolysis has been completed. the pump was not returned for investigation. the data log was missing for the first seven hours of case star. the available data log showed no abnormalities related to positioning issues during the case run. the placement signal showed a slightly recovering trend at the start of the case without observing any motor current trend. additionally, a motor current spike was seen on a day after the patient reported red urine. the cause of the hemolysis issue was most likely biomaterial based on data log review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13087 | 1220648-2024-24693 | 20901666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20901666&pc=OZD | 11/26/2024 | 01/07/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573841 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the impella cp was turned on and could not achieve flows any higher than two liters per minute on auto and when switched to manual, p4 had the same flows as p6. another kink in the cannula was noticed and suction events presented as with previous impella cp implants in the catheterization lab. the decision was then made to perform a sternotomy and place an impella 5.5 direct. the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. the impella device was not returned for evaluation. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely cannula kink.
13088 | 1220648-2024-24694 | 20901746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20901746&pc=OZD | 11/26/2024 | 01/31/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505667 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. approximately nine days after start of support, the pump began to have low flows and would only work at support level p-1. biomaterial was observed and the decision was made to escalate to impella 5.5 device support, which was successfully completed. it was reported there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. data logs were not available. the returned complaint impella cp pump was visually inspected. a big chunk of biomaterial was observed around the impeller, purge gap, inflow cage and cannula. no cannula kink was observed. no broken blade was found. the root cause of the low pump flow issue was most likely biomaterial. d.9 device return date/information was added.
13089 | 1220648-2024-24695 | 20901959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20901959&pc=OZD | 11/26/2024 | 01/09/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  | 11/30/2025 |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425074 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had two new clots in the left ventricle. lactic was increased and the clots were monitored until the impella was removed as planned. the patient is stable. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b was missing from manufacturer device report 1220648-2024-24695. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
13090 | 1220648-2024-24696 | 20901992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20901992&pc=OZD | 11/08/2024 | 01/20/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562121 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Failure | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure with a possible relationship to the impella device used: ¿patient intubated after impella procedure and continued to be intubated until death. patient had difficulty weaning from the vent after the impella implant. team had concern for amio induced in toxicity vs. eosinophilic pneumonia vs. other inflammatory process s/p treatment for aspiration pneumonia (ended (b)(6) and bal (b)(6) showing eosinophilic process. empiric steroids started (b)(6) 2024. team wanted to plan for a trach but family did not wish to pursue. patient was transitioned to cmo on (b)(6) 2024 and died (b)(6) 2024.¿. it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation for the respiratory failure has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues respiratory failure could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13091 | 1220648-2024-24697 | 20902078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20902078&pc=OZD | 11/26/2024 | 01/15/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573401 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with multivessel coronary artery disease. the team was unable to place an impella 5.5 pump and so the team attempted to place a impella cp via the axillary site. the impella cp additionally failed delivery and was aborted. there was no harm or injury noted from delivery attempts. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition because the pump was unable to pass through due to the calcification in the vessels despite multiple attempts. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13092 | 1220648-2024-24698 | 20902206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20902206&pc=OZD | 05/31/2024 | 01/13/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510423 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pseudoaneurysm | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured pseudoaneurysm of right groin with a "probable" relationship to the impella device used. the patient was brought down to the catheterization lab to undergo coronary angiogram with probable intervention with impella. coronary angiogram was completed revealing high grade stenosis to arteries. distal abdominal aortogram with right iliac runoff was completed showing aneurysmal descending aorta and iliac but very large caliber vessel. at this time, the physician consulted with the vascular physician about the patient and vascular anatomy. ultimately, the decision was made to proceed with impella percutaneous coronary intervention which was completed. the pump was supporting for just under two hours and was explanted. the access site had perclose issues. the "j" wire was inserted into the sheath and the sheath was removed. while attempting to tighten the perclose sutures, the physician could not get the last deployed perclose sutures to tighten. it appeared the sutures were wrapped. therefore, the decision was made to abort closure, and a new 14 x13 sheath was inserted back over the "j" wire into the access. hemostasis was achieved. perclose sutures were broken and cut. plans were made for sheath removal on the unit when activated clotting time was less than 160 seconds. doppler pulses were present post suture of sheath. following, same day, post procedure, it was noted the patient endured pseudoaneurysm of the groin. action taken to address the pseudoaneurysm is unknown. the patient recovered with no sequelae. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13093 | 1220648-2024-24699 | 20902436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20902436&pc=OZD | 11/27/2024 | 02/03/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482362 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella 5.5 support. while on support, the patient endured an ischemic stroke. the patient is stable. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13094 | 1220648-2024-24700 | 20902669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20902669&pc=OZD | 11/06/2024 | 01/15/2025 |  | 12/11/2024 |  |  | 11/21/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1254276 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications during impella 5.5 support. seven days into support, the pump began to experience flow saturation. the staff was advised that the pump was struggling and could potentially fail. despite the noted issue, the staff opted to continue use of the device. upon follow up, it was verified that the pump later experienced an unexpected stoppage. the pump was removed and an intra-aortic balloon pump was placed for ongoing support. the site requested that the automated impella controller be investigated for potential failures related to the incident. the patient's outcome is unknown. | [5 empty] | The investigation for the console (aic) pump stop has been completed. since the console was not returned, all logs were retrieved from rlfs. real time (rt) logs show the pump stop occurred again and purge pressure was well above 1000 mmhg. imc logs were not available for this case on (b)(4) but the surrounding cases showed no signs of high purge pressure or pump stops seen during this case. the pump stop was unrelated to this console. the console performed as expected. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13095 | 1220648-2024-24701 | 20902761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20902761&pc=OZD | 11/26/2024 | 01/15/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570602 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a pump was observed to have a significant kink in the cannula. as a result, the pump was removed with an intra-aortic balloon pump. the patient expired. however, there was no additional information. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of product damage (cannula kink) has been completed. the data logs confirmed low pump flow when pump started & remained to the rest of the case. visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. the root cause of low flow was kinked cannula. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13096 | 1220648-2024-24702 | 20903265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20903265&pc=OZD | 11/26/2024 | 01/27/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  | 10/31/2025 |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412892 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. after removal, an echocardiogram revealed a blood clot in the left ventricle. since opening the chest and performing a thrombectomy would be highly invasive, the plan was to administer anticoagulant therapy and to wait until the blood clot was dissolved. the patient was managed with an activated clotting time of 160 to 180 seconds. the patient was successfully weaned and survived the explant. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source; foreign was inadvertently omitted from manufacturer device report 1220648-2024-24702. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13097 | 1220648-2024-24703 | 20903324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20903324&pc=OZD | 11/27/2024 | 01/13/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1434576 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a battery failure. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) battery failure-red alarm has been completed. after reviewing the console logs, the reported battery failure issue was confirmed. there were multiple battery failure (transport connection blown/missing) [event set #360] and battery level low (50%) [event set #23] alarms observed in the imc logs from the reported occurrence date. the console was returned for evaluation. the reported battery failure alarm was reproduced. more specifically, batttest results revealed battery 2 had failed. its logical fuse ¿fu¿ indicator was missing in the batttest results, which is indicative of that battery having no voltage across the negative and positive terminals. further testing was done to this aic in collaboration with field service and the issue was isolated to battery 2. the reported battery failure issue was determined to be caused by a defective battery.
13098 | 1220648-2024-24704 | 20903574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20903574&pc=OZD | 11/27/2024 | 01/31/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568790 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis roughly one day into impella cp support. the pump was repositioned, but the condition persisted and the decision was made to explant the device. the patient survived the explant. | [5 empty] | The investigation of hemolysis has been completed. the cause of the hemolysis issue cannot be determined since the complaint device not returned for analysis and lack of clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13099 | 1220648-2024-24705 | 20903807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20903807&pc=OZD | 10/11/2024 | 01/13/2025 |  | 12/11/2024 |  |  | 11/20/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496167 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a suction alarm goes off when you go over p2. the impella cp p-level was unable to increase greater then p2 despite troubleshooting for adequate position (inlet mid-ventricle on echocardiogram) and adequate volume with impella rp flex in place and providing adequate flows. the automated impella controller was exchanged which resolved the issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the impella cp pump was not returned. data log review indicates several suctions alarms at higher p-level and even few events at p-2 (lower p-level) too. low flows observed immediately after pump start at p-9 with 2.0 to 2.5 l/min. flows at p-2 were 1.0 to 1.5 l/min. cp pump usage log data for first 3 hours on (b)(6) 2024 could only be retrieved. complete log data for cp pump data is could not be retrieved to review. no motor current spike, major position alarm issues or other placement signal trends were seen. clinical mentions patient had sufficient volume. all other patient conditions were unknown. the cause of the low pump flow issue was not determined since product was not returned and due to insufficient evidence per log and clinical information.
13100 | 1220648-2024-24706 | 20903975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20903975&pc=OZD | 11/27/2024 | 01/13/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1734564 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller screen was flickering. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. after the console was received, the reported issue were reproduced, and intermittent horizontal lines with varying intensity were observed on liquid-crystal display screen. further testing was conducted to isolate the malfunctioning component was inconclusive. the root cause of the issue was a defective component. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24706 was submitted in accordance with updated procedures. d.9 revised date as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-24706.
13101 | 1220648-2024-24707 | 20904243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20904243&pc=OZD | 11/27/2024 | 03/10/2025 |  | 12/11/2024 |  |  | 11/27/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was not booting up. the aic was removed from service. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) a/c power issues has been completed. console was returned for investigation but console was not able to get through boot up as it would immediately get caught in a loop and keep restarting in perpetuity, and additionally, the console would only boot up on ac power. logs were unavailable and log analysis was unable to be performed for this console. the flash cards were replaced with known-goods and the console was able to boot up as expected, but battery power was shown to be 0%. battery test was performed and both batteries were shown to be depleted. the batteries were replaced with known-good batteries and the malfunction was resolved. one of the causes of the aic's inability to boot up was depleted batteries but the lack of available logs could not determine what caused the batteries to become depleted. one of the causes of the aic's inability to boot up was corrupt compact flash cards but the lack of available logs could not determine what caused the flash cards to become corrupt. corrections to the initial mfr report: d2 corrected common device name.
13102 | 1220648-2024-24708 | 20904275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20904275&pc=OZD | 11/26/2024 | 01/10/2025 |  | 12/11/2024 |  |  | 11/26/2024 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571900 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted into the left ventricle but noted to have a substantial kink in the cannula just distal to the outlet cage. a decision was made to replace the pump with a new one with the integrity of the cannula compromised. the patient expired. however, there is no additional information. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the product kink has been completed. the pump was returned for investigation. cannula kink was reported on returned product. no other physical damage was noted. the data log shows low pump flow from the start of the case, with an active suction alarm. the pump only ran for six minutes. according to the clinical details, a kink in the cannula was seen after insertion of the pump, and the doctor placed a new pump. additionally, the patient was reported to have a small aortic arch. the failure mode of the cannula kink was changed to low pump flow, as the pump damage resulted in low pump flow. the cause of the low pump flow issue was a kinked cannula. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13103 | 1220648-2024-24709 | 20904336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20904336&pc=OZD | 02/06/2023 | 12/11/2024 |  | 12/11/2024 |  |  | 02/13/2023 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094049 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) issued a controller failure alarm during a clinical procedure. no information was provided on the patient information, however there were no adverse events related to the product malfunction reported. | [5 empty] | The investigation for the reported issue has been completed. console logs from the reported day of event are consistent with complaint and show controller error alarm (due to loss of communication to kbd) on two consecutive boot-ups. there¿s also unusually high number of error messages seemingly associated with ¿high guard level¿ which, in normal operating conditions, is indicative of touch detected outside of the soft-key area on the aic¿s screen but might also be evidence of equipment malfunction. there¿s no evidence of any attempts to initiate support using (b)(6) on that day (no pump connected, no purge cassette connected). it is unknown whether the kbd was unresponsive at that time. after console was received in aachen fs and tested, the issue has not been reproduced. inspection of the console revealed some minor irregularities (kinked inpellatronic cable, contaminated optical pump connector) but unrelated to the complaint or the failure mode. it¿s been determined that most likely component causing kbd communication errors was the 4-in-1 carrier, that contains sub-circuitry responsible for processing data from kbd assembly. the cause of the aic issue was a defective 4in1 board. this report is being filed as part of a retrospective review of historical records.
13104 | 1220648-2024-24711 | 20904561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20904561&pc=OZD | 04/17/2023 | 12/11/2024 |  | 12/11/2024 |  |  | 04/17/2023 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1608973 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller sounded an alarm tone for battery failure. the issue was discovered in preparation, and it was noted there was no harm to a patient (patient-specific information unknown). | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the aic logs confirmed the reported failure. there was a battery failure alarm: transport connection blown/ missing and battery level low (50%) and battery failure alarm: low voltage [v<12v]. the aic was analyzed and confirmed that the battery alarm was due to a defective battery 1 (cell 1). the cause the defective battery was due to cell imbalance. the batteries were replaced, and a functional check and power system testing were performed before returning the device back to service. this report is being filed as part of a retrospective review of historical records. note: the specific contact's name is unavailable and therefore, reflected as "unknown" in section e1 accordingly.
13105 | 1220648-2024-24710 | 20904572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20904572&pc=OZD | 02/17/2023 | 12/11/2024 |  | 12/11/2024 |  |  | 02/17/2023 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) in use for a clinical procedure presented with issues not sensing the purge cassette and disc when priming. the medical team replaced the purge cassette, however this did not resolve the issue. the aic was then replaced, and the issue was resolved. there were no adverse events related to the product malfunction reported. no patient information was provided. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show multiple and persistant rfid errors that interfered with proper detection of the purge cassette which prevented purge and priming from operating during case start. after the console was received in danvers and tested, the issue was reproduced, although intermittently (manifested by rfid error messages and malfunction of the purge cassette detection). by detaching and slightly repositioning the purge assembly, it¿s been determined that the location of the rfid pca, with respect to the other components of the console (4-in-1 carrier and pbm), was affecting its operation - likely the defective rfid pca became susceptible to emi (electro-magnetic interference). temporarily replacing the rfid pca with a known-good one resolved the issue and the rfid error could no longer be reproduced. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
13106 | 1220648-2024-24713 | 20908273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20908273&pc=OZD | 03/01/2023 | 12/11/2024 |  | 12/11/2024 |  |  | 03/01/2023 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support and during a cassette change, the automated impella controller (aic) would not recognize the purge disk. a new cassette was used, however the issue did not resolve. the aic was exchanged, and the issue was resolved. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event were inconsistent with complaint and did not show any entries indicating active support. prior logs were obtained from the cloud, but were incomplete and did not show the moment of purge disk not being detected. inspection of the console revealed broken disk detect flag, which was consistent with reported events. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
13107 | 1220648-2024-24712 | 20908446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20908446&pc=OZD | 02/23/2023 | 12/11/2024 |  | 12/11/2024 |  |  | 02/23/2023 | 12/11/2024 |  |  |  |  | 12/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support, the automated impella controller's (aic) purge disk detect flag broke and the alarms could not be silenced. the medical team was informed on how to change out the aic to resolve the alarm. it is unknown how the issue was resolved. there was no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that during several purge changes, the purge disk was not immediately detected which triggered an alarm (#402, purge disk not detected). during first few purge changes the issue was resolved, but at some point the condition (purge disk not detected) became permanent, due to broken disk detect flag. because fault condition could not be resolved, pressing mute key only silenced alarm for a short period of time. console inspection performed in danvers confirmed broken disk detect flag, but also revealed compromised purge pressure sensor due to purge fluid ingress inside the purge assembly, as well as contamination of the purge insert and purge disk retainer. it is most likely that purge fluid ingress under the purge disk retainer immobilized the flag causing its breakage. this report is being filed as part of a retrospective review of historical records.
13108 | 1220648-2024-24714 | 20909168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20909168&pc=OZD | 04/21/2021 | 05/20/2025 |  | 12/11/2024 |  |  | 04/21/2021 | 12/11/2024 |  |  | 02/28/2023 |  | 12/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding from their chest following an coronary artery bypass graft procedure. the patient had been heparinized prior to the procedure for impella use, however the day and time of discontinuation of the heparin is unknown. the patient was transfused with 2 units of packed red blood cells and was given methylene blue due to the bleeding. three days later, on (b)(6) 2021, the patient's status was noted as critical. the patient expired the following day on (b)(6) 2021. it was note that the medical team suspected the patient to be in right ventricular failure. the medical team was informed of the importance of correctly managing right heart failure with impella support. a plan had been made for further revascularization and a surgical consult, however the patient expired prior to any further action. the use of the impella device cannot be excluded as a possible contributing factor in the patient's outcome. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13109 | 1220648-2024-24627 | 20888806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20888806&pc=PYX | 11/22/2024 | 01/10/2025 |  | 12/10/2024 |  |  | 11/22/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025569634 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding. pump insertion was successful, but once the peel away sheath was removed, there was more bleeding than anticipated. despite manual pressure and securing a purse-string suture, the bleeding persisted. separate attempts to hold pressure and apply more sutures did not achieve hemostasis so cardiothoracic surgery was consulted. two additional ¿woo¿ stitches (purse-string sutures with an additional plastic piece) were placed allowing for hemostasis. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, two additional ¿woo¿ stitches (purse-string sutures with an additional plastic piece) were placed allowing for hemostasis as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
13110 | 1220648-2024-24632 | 20891169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891169&pc=OZD | 11/23/2024 | 12/30/2024 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella support when the automated impella controller (aic) power cord was damaged and the aic would not charge. the aic was replaced and support continued. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) battery issues has been completed. console logs did not contain any data relevant to this complaint. the aic was returned for evaluation and was tested. the reported charging issue was able to be reproduced. the reported damage on the power cord was confirmed after inspection. the copper wire on the inside of the ac power cord plug was detached. the root cause of the charging issue was confirmed to be the damaged power cord. added h6 impact code 4629. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13111 | 1220648-2024-24633 | 20891343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891343&pc=OZD | 08/24/2022 | 12/10/2024 |  | 12/10/2024 |  |  | 08/29/2022 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1453610 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, the white alarm "controller failure" sounded, and the position waveform was not showing. alarm cancellation and controller failure alarm occurred repeatedly, and the following morning the alarm was still ringing. the controller was replaced. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported optical signal issue has been completed. data analysis: aic logs were not available for analysis because a software upgrade was performed, erasing all available logs. device analysis: the console was tested by japan field service (fs) in collaboration with post-market engineering. no issues were observed while running a known-good optical pump. 5s parameters were within specification. however, the snr value was 4032 (still > than snr spec of 2000). as a precaution, fs was instructed to replace the optical bench (0042-1012) and reperform the pm procedure (0042-7309). the product was returned, and the optical signal issue was not confirmed. per the results of the clinical review and investigation: the root cause was not determined due to the inability to confirm and reproduce the reported placement signal issues. this report is being filed as part of a retrospective review of historical records.
13112 | 1220648-2024-24634 | 20891351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891351&pc=OZD | 11/23/2024 | 01/10/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp support and there was access site bleeding. there was continuous bleeding at the right femoral artery access site after the peel away removal and repositioning sheath placed despite the fact that perclose devices were in place. due to the unresolved bleeding, the pump was explanted and 2 perclose were used to achieve successful hemostasis. a new pump was placed in the opposite left femoral artery. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source; study missing from manufacturer device report 1220648-2024-24634. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13113 | 1220648-2024-24635 | 20891519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891519&pc=OZD | 11/23/2024 | 01/07/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing for a post transcatheter aortic valve replacement patient. the cp was placed via the 14fr sheath, which the doctor did not peel away. when there was limb ischemia the team removed the peel away, and finally had to explant the cp pump itself. femostop was placed for hemostasis. the support was for just over 9 hours. the patient expired after care was withdrawn. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13114 | 1220648-2024-24636 | 20891615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891615&pc=OZD | 11/23/2024 | 01/10/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528300 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, there was drainage from the jackson-pratt (jp) drain. heparin was discontinued. a hematoma was noted around jp/impella site. jp drain clotted overnight. the nurse stated she stripped the drain and received more output. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13115 | 1220648-2024-24637 | 20891694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891694&pc=OZD | 11/23/2024 | 01/09/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568058 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had multiple episodes of ventricular tachycardia/torsades requiring shocking/defibrillation and medication adjustments. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added.
13116 | 1220648-2024-24638 | 20891794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891794&pc=OZD | 11/23/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the cannula kinked and flow was lost. suction alarms were constant, most likely due to kink in cannula. additionally, oozing was noted at the access site. the pump was repositioned to control bleeding. heparin was held and a stitch was placed. the bleeding was resolved. the patient was bridged to a left ventricular assist device two days later and survived the explant. | [5 empty] | The investigation of product damage (cannula kink) and access site bleeding has been completed. the data log showed low pump flow from the start of the case, with an active suction alarm. according to the clinical details, the pump cannula was kinked during placement, but the desired flow was eventually achieved after the kink was relieved, and the doctor kept the pump running at a lower level for needed support. additionally, the doctor redressed the access site twice, which resolved the bleeding. after investigation data log, cannula kink failure mode was changed to low pump flow and addressed under low pump flow failure mode as the potential root cause. the cause of the low pump flow issue was most likely kinked cannula, based on given clinical details. the cause of the access site bleeding issue was most likely use related since the bleeding was resolved after redressing the access site. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13117 | 1220648-2024-24639 | 20891809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891809&pc=OZD | 11/23/2024 | 12/04/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606424 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support when the automated impella controller (aic) controller error alarms that self-resolved. the aic was not replaced and support continued. the aic was sent back for investigation. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: d2 device name has been update. h6 type of investigation code added.
13118 | 1220648-2024-24640 | 20891892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891892&pc=OZD | 11/23/2024 | 01/27/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535453 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for a patient admitted in acute myocardial infarction/cardiogenic shock. the pump was placed and supported for over 16 hours when the pump came out of position. the loss of left ventricle positioning could not be troubleshot and so the pump was successfully weaned. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
13119 | 1220648-2024-24641 | 20891969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20891969&pc=OZD | 11/23/2024 | 01/15/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557232 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male | 143 KG | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion of the device, it was documented that the pump was unable to advance through the patient's vasculature. after multiple failed attempts, the implant was aborted and an impella cp was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
13120 | 1220648-2024-24642 | 20892004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892004&pc=OZD | 11/24/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump supporting for a patient admitted in after transfer from community to tertiary center. purge pressure low alarms occurred following a bag change. no leasks were noted. troubleshooting was attempted by mock bag changes and de-airing process without any resolution in alarm. no patient harm was reported. | [5 empty] | The investigation for the purge pressure low has been completed. no product was returned for evaluation. data logs were reviewed and the lt logs showed seven instances of "low purge pressure" alarms occurring during the case. each of the instances of "low purge pressure" alarms showed that a bag or cassette change occurred right before alarm was triggered. all alarms were able to resolve on their own or after multiple troubleshooting attempts including bag changes or de-air procedures. the customer reported the issue against the purge cassette so the product investigated was changed to the purge cassette. additionally, as the "low purge pressure" alarms occurred directly after a bag or cassette change. the root cause of the priming issue was most likely use related as all low purge pressure alarms occurred directly after a bag or cassette change and is able to resolve. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13121 | 1220648-2024-24643 | 20892163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892163&pc=OZD | 11/23/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had coronary artery bypass grafting surgery. when they came back from surgery, staff were having trouble finding a pulse. the pump was explanted. the patient survived the event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13122 | 1220648-2024-24644 | 20892186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892186&pc=OZD | 11/24/2024 | 01/16/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551722 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the nurse was concerned about placement signal. also, the aorta (ao) pressure on the monitor was double what the patient¿s actual pressure and mean arterial pressure was and there was no left ventricle signal noted at level p-8. performance level and pulsatile motor current, flows and pressure were all visible and appropriate. no ao placement signal was noted. additionally, there was some bleeding noted around dressing site. surgical gauze was applied with new dressing that remained clear and dry. the patient received one unit of packed red blood cells due to low hemoglobin, hematocrit, and platelets. no harm was reported. | [5 empty] | The investigation for the optical issue and access site bleed has been completed. no product was returned for evaluation. log review showed a temporary drop in snr and contrast. there was also a slight increase in gain to 1.4 and a decrease in light under 2.7. lamp was stable at 94%. the elevated placement signal lasted for approximately 2.5 hours. there were no p-level changes around this time. log review did not show any "placement signal not reliable" alarms. there were also several suction alarms throughout the case. there was also a slight rise in motor current that occurred before the optical issue began, and pump flow became variable shortly after. the motor current and pump flow changes resolved at the same time as the optical signal issue. the pump was not removed due to the optical signal issue. the root cause of the optical signal issue was not determined since no product was returned and insufficient clinical detail was provided. ¿the root cause of the access site bleeding was not determined since product was not returned and insufficient clinical details were provided. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13123 | 1220648-2024-24645 | 20892261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892261&pc=OZD | 11/23/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/23/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in coding in pulseless electrical activity and having cardiopulmonary resuscitation. the team placed a femoral impella cp to support for a possible high-risk percutaneous coronary intervention. the team placed the pump and there was oozing/bleeding at the site. three units of red blood cells were given. after 19 hours, the family made the choice to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to anticoagulation management since the activated clotting time (act) was above recommended range. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13124 | 1220648-2024-24646 | 20892296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892296&pc=OZD | 11/24/2024 | 01/28/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1185157 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was an issue with the automated impella controler (aic). a frustrating attempt to change the purge cassette because it was not recognized was reported. the holder for the purge disk was broken and the recommendation was given to get another aic console. no patient harm has been reported. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the user removed the purge cassette and re-inserted multiple times without the purge disc ever being recognized confirming the complaint. the console was returned and it was confirmed that it was missing the purge retainer and flag which would have prevented the console from recognizing a purge disc. the root cause of this issue was a broken purge disc flag. d.9 date device returned/information was added. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24646 was submitted in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13125 | 1220648-2024-24647 | 20892423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892423&pc=OZD | 11/24/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient transferred into the tertiary center from the community hospital for heart failure management. the pump had poor positioning which affected the support, and the patient expired when the family decided to withdraw care. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the positioning issue has been completed. no product was returned. positioning alarms confirmed the pump's change in positioning to the aorta. the positioning alarm did resolve. it was noted that the pump moved out of position while moving the patient. the root cause of the positioning issue was patient movement. b.7 additional information was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-24647. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
13126 | 1220648-2024-24648 | 20892506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892506&pc=OZD | 11/24/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  | 03/31/2025 |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328793 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male | 143 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a pediatric patient supported by an impella cp placed via the axillary due to impella 5.5 being too large. the impella cp was placed due to cardiogenic shock and along with venoarterial extracorporeal membrane oxygenation with cardiac family history. the pump was placed while the team considered heartmate 3 or protek duo. after 29 hours of support, ischemia developed. the decision was made by the team to explant. the patient survived the explant. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13127 | 1220648-2024-24649 | 20892688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892688&pc=OZD | 11/24/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558978 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Abscess | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed some swelling at the access site. the care team believed it was a small abscess and started the patient on an antibiotic. the patient is stable. | [5 empty] | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. added h6 clinical code 1690. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13128 | 1220648-2024-24650 | 20892867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892867&pc=OZD | 11/24/2024 | 04/11/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563473 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella cp support. while on support, a pump stop occurred. the pump was replaced. the procedure was outcome was successful with the new pump. the patient survived the event. | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately 11 days later, the motor current was high with saturated flows 4l/min and following the pump stopped. the pump was restarted again. however, the device was explanted and replaced with another impella cp pump 2. however, this device too would temporarily stop. the patient was noted to have subsequently expire. medical safety review of the event noted there is limited information. there is not enough evidence to exclude impella cp pump 1 (this report) as an associated factor in the patient's outcome. there is no information on ultrasound to diagnose a possible clot (as noted for impella cp pump 2 event story). |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the impella device was not returned for evaluation. as per the data log review and clinical information, the evidence suggests biomaterial ingestion. therefore, the root cause of the temporary pump stop issue was most likely biomaterial. the following have been reported incorrectly on manufacturer device report 1220648-2024-24650. | This is one of two devices associated with the reported event. refer to manufacturing report number 1220648-2025-27059 for report on the impella cp device pump 2. correction was made to type of reported event to reflect death. note: device investigation previously reported remains applicable.
13129 | 1220648-2024-24651 | 20892915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20892915&pc=PYX | 11/24/2024 | 01/02/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  | 11/30/2024 |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024286677 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp and impella rp for mechanical circulatory support and following had high purge pressure and a gastrointestinal (gi) bleed. the patient was admitted to the intensive care unit where the patient was intubated upon induction pulseless electrical activity arrest with 30 minutes of cardiopulmonary resuscitation. after successful return of spontaneous circulation, the physician successfully placed the impella cp to the right femoral artery without complications. left heart catheterization was completed and diffused disease with severe lesion in the right coronary artery (rca) was identified. three stents were placed to the rca with no intervention to the left system. after the physician proceeded with implanting an impella rp. the same day, post procedure, a critical purge alert triggered and dextrose five percent in water was switched to tissue plasminogen activator three milligrams in 50 cubic centimeters of sterile water. on day 4 of support, the site was oozy with dressing change. a gi bleed was present over the night. heparin was taken out of the impella rp purge and replaced with sodium bicarbonate. it was noted that systematic heparin was started which was lowered that morning to 550 from 850. overnight, the impella cp was repositioned and the patient was given two units of blood and one unit of platelets. the next day, the impella rp was weaned to performance level p-4. the gi bleed slowed down, and sodium bicarbonate of the impella rp purge was switched back to heparin due to purge flow dropping the past 12 hours. the patient had received two units of blood and one unit of platelets overnight with a possibility of another unit of blood that day. an echocardiogram was completed to see if the patient could be weaned from both devices. the next day, now day 5 of support, both devices were explanted, impella rp at the bedside and the impella cp in the catheterization lab. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella cp serial number (b)(6) with date of event 21-nov-2024 and identifier ending in (b)(6). | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella rp system with automated impella controller states the following: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
13130 | 1220648-2024-24652 | 20893125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893125&pc=OZD | 11/24/2024 | 01/09/2025 |  | 12/10/2024 |  |  | 11/24/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302864 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller (aic) had a battery communication failure. field service stated this error could not be found in the logs. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. the failure mode could not be confirmed in logs nor testing. logs showed battery performance as expected with no related alarms or errors and testing mirrored that. root cause: there was no issue found during the case. the device functioned/operated to specification. d3 manufacturer us zip code was missing on manufacturer device report 1220648-2024-24652.
13131 | 1220648-2024-24653 | 20893152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893152&pc=OZD | 10/01/2024 | 09/25/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559572 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Per investigation results, a complainant had impella cp pump support initiated for a patient admitted in cardiogenic shock/acute myocardial infarction. the pump remained on until the team observed a purge leak. the pump was explanted two days later. the patient survived without any harm or injury. | The complainant also reported that the optical signal was lost, but the pump issue caused no harm or injury or pump explant. |  |  |  |  | The investigation has been completed. the purge sidearm and a purge cassette were returned. the severed sidearm was inspected and a crack along the length of the yellow luer was observed. the sidearm was connected to the purge fluid via the purge cassette and the leak was reproduced. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer. no data logs were returned for analysis. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. insufficient product was returned for investigation. no data logs were returned for analysis. the root cause of the optical signal issue was not determined as no data logs and insufficient product were returned for analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated.
13132 | 1220648-2024-24655 | 20893541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893541&pc=OZD | 10/09/2024 | 12/10/2024 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526529 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had impella cp support ongoing for a patient admitted in for a high-risk percutaneous coronary intervention. access site bleeding was noted after intervention. the pump support was for over one hour and then was explanted. the protect iv documentation noted a perclose failure. no blood products or vascular repair was needed. femostop and a new closure device, an angioseal, were needed. the patient survived the explant. | [5 empty] | The investigation has been completed. the root cause of the access site bleeding was most likely use issue since the device closure proglide failed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13133 | 1220648-2024-24657 | 20893573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893573&pc=OZD | 11/25/2024 | 01/10/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443375 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during repositioning of an impella 5.5 pump, the sterile sleeve came off of the repositioning sheath. the nurse wrapped the exposed catheter in sterile gauze and transparent film dressings. support with the implanted pump continued following the troubleshooting steps. there was no patient harm. the patient remains on impella support. | [5 empty] | The investigation for the product damage has been completed. product was not returned for investigation and data logs are not relevant to the complication. it was reported that the sterile sleeve ripped during pump repositioning and that excessive force was used. based on clinical details, the root cause of the sterile sleeve rip was determined to most likely be a use issue. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13134 | 1220648-2024-24654 | 20893768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893768&pc=OZD | 11/25/2024 | 01/30/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518422 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | White | Malfunction | No | Yes | Packaging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient with high-risk percutaneous coronary intervention. however, the team observed that there was a packaging issue with the sheath that prompted the team to proceed without all the sheath lengths. the team noted the package contained only two short sheaths and no longer sheath length. there was no harm reported. | [5 empty] | The investigation for the package damage has been completed. the introducer kit with the dilators were returned in a different packaging. a picture from the field was provided but it was not in the original packaging. the kit was on the table without the original packaging. without that it is not possible to investigate the failure. therefore, the root cause of the packaging issue was not determined since sufficient proof was not provided. | The introducer and guidewire were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13135 | 1220648-2024-24656 | 20893808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893808&pc=OZD | 11/25/2024 | 01/28/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515621 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, there was a decrease of impella flow 0.1 l/min (p9). correct position was verified and there was no impella flow on p9. the surgeon decided to explant impella and switch to a new pump. the patient survived the event. | [5 empty] | The investigation of low pump flow has been completed. as per data review and device analysis the root cause of the low pump flow issue was cannula kink due to improper technique. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13136 | 1220648-2024-24658 | 20893824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893824&pc=OZD | 11/25/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567253 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the companion sheath was damaged and unable to use out of the box. a replacement sheath was used and there was no patient harm. | [5 empty] | The investigation for the introducer damage has been completed. the companion sheath was returned for evaluation. two kinks were present on body of companion sheath. the returned product label for the 7fr companion sheath was also returned. clinical details noted that sheath was not used and damaged straight out of box. as the customer reported that the sheath was found to be damaged out of box. the root cause of the packaging issues was most likely a damaged component as the sheath was found to be kinked out of box. a trend chart was pulled for all 7fr low profile companion sheaths used with impella cp pumps from november 2023 to november 2024 with a failure mode of "packaging issues" and a root cause of "defective component", this is the only occurrence therefore no trend could be established. | An impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13137 | 1220648-2024-24660 | 20893990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20893990&pc=OZD | 11/23/2024 | 01/02/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571124 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a left ventricle (lv) perforation. a washout was done at bedside after a massive transfusion protocol, since bleeding was found coming from the left internal mammary artery graft. the surgeon manipulated the heart and rotated after fixing the graft and then closed the chest. however, the patient continued to have bleeding and was taken back to the operating room where a lv perforation was found. the lv was patched and another washout was done. the patient remained stable. the patient survived the event. | [5 empty] | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13138 | 1220648-2024-24659 | 20894355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20894355&pc=OZD | 11/25/2024 | 01/17/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559154 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding requiring intervention. the impella cp device was implanted and the physician noted that the patient was in general condition with extracorporeal membrane oxygenation (ecmo) implant planned for later that day and stabilization phase. ecmo was placed. following, it was noted that the patient was lying on the right side and begin "heavily" bleeding from the right femoral artery access site requiring five units of blood. it was noted that the bleeding was stopped. the physician planned to escalate to an impella 5.5 and the impella cp was explanted the following day. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13139 | 1220648-2024-24662 | 20894606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20894606&pc=OZD | 11/25/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535604 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was bleeding between the blue suture hub and the t lock. the doctor determined that there was likely a tissue track bleeding issue. one unit of packed red blood cells was given to the patient. the pump was exchanged. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, the doctor determined that this was likely a tissue track bleeding issue as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13140 | 1220648-2024-24661 | 20894706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20894706&pc=OZD | 11/25/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the hematoma, and limb ischemia. the patient lost pulses in feet, for which a reperfusion line was inserted, originating from the left femoral artery and supplied the right femoral artery. external femoral to femoral bypass was placed. the access site bleeding was treated by holding pressure held and giving two units of blood to the patient. the patient remains on impella support. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13141 | 1220648-2024-24666 | 20895164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895164&pc=OZD | 11/25/2024 | 01/08/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574187 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had a hematoma at the access site. pressure dressing was applied and heparin was withheld. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13142 | 1220648-2024-24665 | 20895192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895192&pc=OZD | 11/25/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported an unexpected controller shutdown alarm on the automated impella controller (aic) when turned on. the aic was replaced and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console was booting up, the logs rebooted with a gap of 51 seconds. the console then booted up again and an unexpected controller shutdown alarm was triggered. the console was then shutdown by the user. the console was returned for investigation and power cycled 200 times but the failure mode was not reproduced. the console performed as expected. the cause of the unexpected reboot was not determined since the failure mode was not reproduced and there was no evidence in the logs to clarify the cause of the reboot. a.1 should have been left blank on initial manufacturer device report number 1220648-2024-24665. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24665 in accordance with updated procedures. h.6 code 2199 was previously entered incorrectly on initial manufacturer device report number 1220648-2024-24665. a new code has been added. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13143 | 1220648-2024-24667 | 20895201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895201&pc=OZD | 11/25/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was vascular damage and suspected thrombocytopenia. the suspected thrombocytopenia was treated by changing to bivalirudin. additionally, surgical wound closure was needed during explant to achieve homeostasis and there was a right femoral cutdown or surgical repair of the right femoral arteriotomy. the patient survived the event. | [5 empty] | The investigation for the vascular damage and thrombocytopenia has been completed. the product was not returned. there were limited clinical details provided in the case. due to a lack of information, the root cause of the vascular damage - peripheral vessel cannot be determined. the root cause of the thrombocytopenia cannot be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13144 | 1220648-2024-24664 | 20895209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895209&pc=OZD | 11/25/2024 | 01/20/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524304 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella cp. the impella was deployed but found to have kinked on itself prior to crossing the aortic valve. the decision was made to remove the impella and place a new impella pump. care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be use issue because second pump was successfully inserted and no issues were noted. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13145 | 1220648-2024-24668 | 20895284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895284&pc=OZD | 11/25/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025459767 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was a guidewire issue. when trying to initially place the impella, the 0.18 wire kept poking out of the side hole of the pig tail. the wire got kinked so the decision was made to replace with a new one. the patient is stable. | [5 empty] | The investigation for the introducer damage has been completed. no product was returned. log review is not applicable. according to clinical details, it was stated that delivering the guidewire through the pump. the root cause of the delivery issue was not determined since product was not returned and insufficient clinical detail was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13146 | 1220648-2024-24663 | 20895338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895338&pc=PYX | 11/25/2024 | 01/10/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025569706 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support. while on support, the impella rp flex was found to have migrated below the pulmonary valve into the right ventricle. several attempts were made to re-advance but were not successful. the decision was made to replace the impella rp flex. case (patient identifier) (b)(6) addresses the second impella rp flex that was implanted. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
13147 | 1220648-2024-24669 | 20895474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895474&pc=OZD | 11/20/2024 | 03/10/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1697069 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Obstruction of Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during startup, the console got stuck during the priming process. it was noted that the console kept showing a screen asking the customer to check the purge fluid bag. there was no report of harm to the patient as a result of the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. on the complaint date, after console boot up, a cassette was inserted before initiating the ¿case start¿ wizard. upon navigating through the 'case start' wizard, step 2 (insert purge cassette and disc) was skipped automatically, and the wizard moved to step 3, as the cassette was already inserted. before inserting the pump, the console recorded ¿purgeadmin: piston blocked (reset pud) 0,¿ indicating that the pud was reset due to the piston being blocked. due to the persistent piston blockage, the pud was reset continuously, and the console displayed ¿case start: purge fluid not detected error.¿ the cassette was then removed and reinserted, and the case start wizard was reinitiated. however, the failure mode was never resolved, confirming the reported failure mode. the console was tested. the reported failure mode ¿purge system failure (piston blocked) alarm¿ was reproduced. additionally, there was a ¿controller error (power battery manager voltage out-of-spec). the piston blockage caused the purge motor to draw more current, resulting in a drop in the vpurge voltage and triggering an event. visual confirmation of dextrose residue was observed behind the purge cassette holder and the purge motor gasket. after cleaning the residue, a full functional check was performed, and there was no longer a controller error event. the root cause of the inability to proceed through the case start wizard while priming was the dextrose deposit between the purge cassette holder and the purge motor gasket. a.1 should of reflected blank on the initial manufacturer device report number 1220648-2024-24669 as the patient identifier is unknown. d.2 revise common device name since the initial manufacturer device report number 1220648-2024-24669 was submitted in accordance with updated procedures.
13148 | 1220648-2024-24670 | 20895483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895483&pc=OZD | 11/26/2024 | 02/11/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559593 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient was hemolyzing and there was a thrombus. the patient survived the event. | [5 empty] | The investigation for the hemolysis and the thrombus has been completed. no product was returned for analysis. the data logs do not show any motor current spikes or any positioning issues before the reported hemolysis. there was no clinical mention of low blood volume or any patient condition related issues. the root cause of the hemolysis and thrombus was not determined as no product was returned for analysis and limited clinical information was provided. g.3 date received by manufacturer was reported incorrectly on the initial manufacturer device report number 1220648-2024-24670. the date should of been 26-nov-2024. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13149 | 1220648-2024-24671 | 20895550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895550&pc=OZD | 11/25/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565635 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue due to the patient's anatomy. the pump was removed as a result. no patient harm has been reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. he root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13150 | 1220648-2024-24672 | 20895774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895774&pc=OZD | 08/02/2021 | 12/10/2024 |  | 12/10/2024 |  |  | 08/03/2021 | 12/10/2024 |  |  | 08/31/2021 |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1444039 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced a hematoma at the impella access site. as a result of the formation of the hematoma, the impella 2.5 was explanted. manual pressure was held to resolve the hematoma after explant. the patient¿s outcome was provided as ¿stable¿. | [5 empty] | The investigation for the reported hematoma has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the patient had been managed with an act (activated clotting time) of over 300 seconds for the percutaneous cardiac intervention. therefore, the root cause of the reported access site bleeding was most likely related to anticoagulation management.
13151 | 1220648-2024-24675 | 20895868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895868&pc=OZD | 09/14/2021 | 12/10/2024 |  | 12/10/2024 |  |  | 09/14/2021 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085509 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a training session when, the automated impeller controller displayed an alarm for "controller error switch to back up controller " after it was booted up. there was no patient involvement. | [5 empty] | The investigation for the reported controller alarm has been completed. data analysis: complaint is consistent with occurrence date. on 9/14/21, imclogs show the following: 14/09/21 08:08:16 imc-kbd error: 0xa4 0x02 0x00 - keyboard: read failure. 14/09/21 08:08:57 imcgui: event set: #24 from 80 (0) - controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1). device analysis: console failure mode could not be reproduced by power cycling console. conclusion: failure mode exhibited in the field is most likely due to the age of the components in ic1207. the root cause of the one-off kbd read failure and subsequent controller error are most likely due to missed communication from the kbd to 4-in-1 and/or vice versa. however, the exact root cause is undetermined due to being unable to reproduce the failure mode.
13152 | 1220648-2024-24673 | 20895880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895880&pc=OZD | 11/24/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/25/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573298 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had ventricular tachycardia (vt) when he coughs and strains. the patient was given hydralazine with a goal of impella flows to be 5.0 or greater. placement was verified. additionally, the patient had a hematoma at the access site. the doctor was not concerned and the impella was explanted the following day. heparin was also withheld, but it is unknown if the heparin was withheld due to the hematoma or if the pump explant was related to the hematoma. the patient survived the event. | [5 empty] | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13153 | 1220648-2024-24677 | 20895883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20895883&pc=OZD | 03/18/2022 | 12/10/2024 |  | 12/10/2024 |  |  | 03/18/2022 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183119 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in non-st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was that the purge cassette pressure disk was unable to sense the pressure during initial pump prime after it was attached to the yellow sidearm of the automated impella controller (aic). the medical team attempted to insert a replacement purge cassette, however this did not resolve the issue. the aic was replaced, and the issue was resolved with the new aic. there was no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. data analysis: imc logs are consistent with the reported failure mode. during case start wizard for console ic3620, purge disk not detected alarm issued and resolved repetitively. the intermittent alarm did not only occur for purge cassette sn (b)(6), but for another replacement pc sn (b)(6) as well. device analysis: console arrived in danvers. fs technician confirmed the broken purge clip and blue retainer via visual inspection of purge unit. the cause of the issue was a defective component.
13154 | 1220648-2024-24676 | 20896002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896002&pc=OZD | 11/25/2024 | 01/30/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565846 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered positioning issues. positioning issues were encountered continuously and ultimately the decision was made to remove the pump. the patient eventually expired after care was withdrawn. | [5 empty] | The investigation for the positioning issue has been completed. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13155 | 1220648-2024-24681 | 20896135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896135&pc=OZD | 10/18/2024 | 09/26/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella 5.5 support. the automated impella controller (aic) was returned for investigation. during review of the logs, oscillating contrast was found. no patient harm has been reported. | The customer reported that the device exhibited an optical sensor error. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) loss of opm data has been completed. data logs were reviewed which showed the pump was inserted into console and had a placement signal unreliable alarm confirming the complaint. this alarm condition was immediately silenced by the user and continued until resolution. the lt log shows this was the only time during the case that the contrast oscillates between +/-3276.8%. and the rtlog at this time shows as the contrast oscillates, the placement signal went to 3002 mmhg as the raw optical sensor was -3276.8 mmhg and snr dropped to close to 1. the console was returned for evaluation. the failure mode was not reproduced during testing. 5s test was run on the console with contrast below specifications which could have influenced the placement signal abnormalities. the cause of the oscillating contrast was not determined since failure mode was not reproduced. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. data analysis: on 10/16, 5.5 pump 524453 was inserted into (b)(6) and on 10/18 at 15:59 had a placement signal not reliable (psnr) alarm ((b)(6) imc psnr.png) confirming the complaint. this alarm condition was immediately silenced by the user and continued until resolution on 10/19 at 7:19 ((b)(6) imc psnr resolve.png). the lt log shows this was the only time during the case that the contrast oscillates between +/-3276.8% ((b)(6) ltlog osc contrast.png). and the rtlog at this time shows as the contrast oscillates, the placement signal went to 3002 mmhg as the raw optical sensor was -3276.8 mmhg and snr dropped to close to 1. device analysis: the failure mode was not reproduced during testing. 5s test was run on the console with contrast below specifications in 0042-7309 section 23 which could have influenced the placement signal abnormalities ((b)(6) 5s.png). root cause: the cause of the oscillating contrast was not determined since failure mode was not reproduced. repair: please replace the optical bench (0042-3015p) as a precaution, then return the old optical bench to pme. then update the lamp runtime, perform section 23 of pm (0042-7309) and functional check (0042-7316). a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue 2917 as part of a retrospective review of optical signal issues in accordance with fda recommendation.
13156 | 1220648-2024-24679 | 20896159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896159&pc=OZD | 09/30/2022 | 12/10/2024 |  | 12/10/2024 |  |  | 09/30/2022 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142970 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an issue with the automated impella controller (aic) that occurred after the aic was turned on. it was reported there was a message at the bottom of the screen that there was a problem with the start-up calibration. this complaint was not related to a clinical procedure. | [5 empty] | The investigation for the reported issue has been completed. console logs from the reported day of event were not available for analysis due to file system corruption. partial logs that were recovered, dated 2022-07-25, were consistent with complaint and showed multiple and persistent communication errors with powermod_1 (pbm). when console was tested in danvers, the issue was reproduced and syscall error prevented the aic from fully booting up. in addition, the console would only start on ac power, due to batteries being fully depleted (batteries were internally shut down - communication has seized). inspection of the console¿s electronics revealed corrosion of the pbm (due to electrolyte leak from supercaps c69, c70). despite pbms sn (b)(6) placing it in the pool of boards susceptible to corrosion, the fcn 71 (designed as corrective action for this failure mode) has not been performed during prior pm of this console on 2021-12-13. the cause of the aic issue was contamination.
13157 | 1220648-2024-24674 | 20896180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896180&pc=OZD | 11/26/2024 | 01/13/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379222 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella mechanical circulatory support, approximately eight days after start of support, the automated impella controller (aic) soft touch buttons became inoperable. the aic was exchanged with no report of interruption in support and no harm to the patient. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. the console logs from the reported day of event are consistent with complaint and show that after 6 days of support, there were no more keypresses recorded in logs for the next 34 hours. during testing the reported complaint was not reproduced, it is most likely that the complaint issue was caused by kbd malfunction, where prolonged data communication anomalies resulted in loss of keypress readings. the root cause of the aic issue (unresponsive keyboard) was a defective kbd/display. d9 date device returned to manufacturer added.
13158 | 1220648-2024-24680 | 20896212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896212&pc=OZD | 11/26/2024 | 01/15/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849988 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller keeps shutting off. the complaint is not related to a clinical procedure. there was no report or indication of patient involvement. | [5 empty] | The automated impella controller was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the automated impella controller shutdown has been completed. console logs from the reported event showed two unexpected shutdowns with the console's software automatically restarting after each shutdown. the first reboot happened after 22 days of pump operation, following a partial reset of the watchdog and a missed acknowledgment by process sau_kbd. the second reboot occurred due to a failed post system (post: system restart not successful). after the system was rebooted, the "unexpected shutdown" alarm was triggered, and the watchdog started again. the pump resumed at the previous p-level. the console was returned for evaluation and the issue could not be reproduced, and no unexpected shutdowns or performance irregularities were observed. additionally, the console was manually power-cycled 10 times without issue. no power supply interruptions were detected between the power battery manager and the 4-in-1 assembly. the most likely cause of the shutdown was a software issue. the root cause of the intermittent, unexpected controller shutdown was most likely a software issue, due to the failure of two software sub-processes (ig4 and kbd). d.9 device return date/information was added. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24680 was submitted in accordance with updated procedures.
13159 | 1220648-2024-24682 | 20896269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896269&pc=OZD | 04/11/2023 | 12/10/2024 |  | 12/10/2024 |  |  | 04/11/2023 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542953 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support as a bridge to transplant. it was reported while on impella cp support, there was a placement signal issue that presented as a placement signal not reliable alarm. to resolve the issue, the automated impella controller (aic) was replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. console logs from the reported day of event are consistent with complaint and show that after 2 days of support, placement signal became unreliable for about 4 hours, only to come back after mump was stopped (p-level deliberately set to 0). the support was continued using same pump (391908) and second console, without any placement signal issues. the console was returned and evaluated. the cause of the issue was a defective lamp assembly component.
13160 | 1220648-2024-24678 | 20896280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896280&pc=OZD | 11/26/2024 | 01/15/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448696 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had a gastrointestinal (gi) bleed and that there was a pump stop. the gi bleed was treated by changing the purge to half strength of heparin and giving the patient albumin. then, no anticoagulation peripherally was given to the patient to further treat the gi bleed. additionally, the pump stopped after 12 days of support. a new impella cp pump was used in exchange. the patient is stable. | [5 empty] | The investigation for the gastrointestinal bleed and pump stop has been completed. the product was not returned. due to limited clinical details, the root cause of the vascular damage - peripheral vessel cannot be determined. the data logs show a pump stop on the 12th day of support. there are no motor current spikes or any purge related events leading up to the pump stop. the design life of a cp is 8 days. the pump was run for 12 days. the root cause of the pump stop is most likely due to bearing wear. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13161 | 1220648-2024-24683 | 20896321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896321&pc=OZD | 11/26/2024 | 02/05/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137157 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a console was having trouble powering up and had a black screen. there was no patient harm or involvement at the time of the failure. | [5 empty] | The investigation for the console (aic) power on issues has been completed. data logs were reviewed which showed that during the boot-up, the log recorded ¿impellalib: found core dump file /userflash/sh.core with size: 573440 and last modified on: (b)(6) 2024. then, system self check failed per the log ¿post: syscall10 checkfirmware failed¿. immediately, the console rebooted automatically. again, during boot-up, the log recorded ¿impellalib: found core dump file /userflash/sh.core with size: 831488 and last modified on: (b)(6) 2024. most likely, during boot-up, the console might have displayed a black screen, which aligns with the claim that ¿when it does, it has a black screen.¿ after boot-up, though the cassette and pump were connected, the console was manually shutdown and restarted. then, the pump was used for 1 hour and 17 minutes. again, the console was manually shut down and restarted. the pump was turned on for 8 seconds, then removed. unable to confirm the ¿console not stating up¿ claim with log evidence, the system self check failed screen was most likely reported as the console not starting up. the console was returned for evaluation. upon functional testing, was unable to reproduce the reported failure mode. upon checking the jama, found a defect (gid-dft-3548323) that addresses an issue similar to the reported failure mode. no hardware issue related to the reported failure mode was found on the console. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). corrections to the initial mfr report: d2 common device name updated | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13162 | 1220648-2024-24684 | 20896434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896434&pc=OZD | 01/11/2023 | 12/10/2024 |  | 12/10/2024 |  |  | 01/11/2023 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203518 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) issued a controller failure alarm. the aic was replaced to resolve the issue. this complaint was not related to a clinical procedure, and there were no adverse events related to this product malfunction. | [5 empty] | The investigation for the reported issue has been completed. data analysis: imc logs did not confirm the reported controller failure alarm on the complaint date. an impella defective (error reading eeprom) alarm was issued after a pump was connected. device analysis: the console was tested with a known good pump, and the failure mode was reproduced. replacing the impellatronics pca resolved the issue. the cause of the impella defective alarm was due to a malfunction of the impellatronics epm circuitry.
13163 | 1220648-2024-24686 | 20896507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896507&pc=OZD | 11/26/2024 | 01/09/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572807 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and had low pump flow seven days after start of support. the care team called the call center during rounds to discuss ongoing suction. it was noted that preload and positioning were correct, and the patient was adequately supported on performance level p-8 at 4.2 l/min with cardiac output/cardiac index. the decision was made to disable audio on suction alarm. the next day noted suction audio alarm continued to remain disabled, although patient was not having active suction alarms. the patient was reported to be stable following the event. | [5 empty] | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
13164 | 1220648-2024-24687 | 20896512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896512&pc=OZD | 11/26/2024 | 01/24/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570438 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, a massive hematoma had formed at the impella access site. repositioning attempts were also made, but these led to even more bleeding.the impella was explanted. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13165 | 1220648-2024-24688 | 20896594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896594&pc=OZD | 11/26/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590786 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and it was aborted due to patient anatomy. the pump was unsuccessful in traversing innominate due to anatomy/calcium. the implant was aborted. no known patient harm or injury. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through due to tortuous condition and calcification of the vessels. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-24688 in accordance with updated procedures.
13166 | 1220648-2024-24689 | 20896635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896635&pc=OZD | 11/26/2024 | 01/31/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553927 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was planned for an impella cp device implant for mechanical circulatory support and the optical sensor system error was displayed. it was confirmed that the connection cable was properly inserted. as there was no problem with the insertion of the connection cable, the automated impella controller was powered on again and the priming process was reperformed. however, the same alarm displayed. the aic was then replaced. again, the error alarm triggered, and message was displayed. a pump catheter malfunction was suspected. the impella 5.5 pump was explanted and another impella 5.5 device was implanted with no said error message. there was no report of adverse effects to the patient. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical issue has been completed. impella cp returned for evaluation. when pump was connected to console, psnr alarm was reproduced. laser light continuity testing was performed and pump did not pass. light was observed from the red handle on the catheter side. upon further inspection, a broken optical fiber was found inside the red handle just before the kink protector on the catheter side. data logs were reviewed and optical parameters out of box are consistent with a fiber break. psnr alarm also noted in rt logs. clinical details noted that during priming process an "optical senor system error" was noted. attempted to restart console and restart priming, then console was changed and did not resolve the issue. no excessive force was applied when handling the pump. a new pump was connected to the console and alarm was resolved. the root cause of the optical signal issue was a damaged fiber. corrections to the initial mfr report: added d6a/d6b.
13167 | 1220648-2024-24691 | 20896691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896691&pc=OZD | 11/26/2024 | 01/14/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571876 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs); however, the attempt to start the pump was unsuccessful. the automated impella controller (aic) reflected "failure." three additional attempts were made to start the pump, and these were unsuccessful. the aic was exchanged, and the same issue occurred. a new pump was successfully placed with start of mcs. the patient was reported to be stable following the event. post event, it was noted there was a concern that the wire was not out of the impella before starting. | [5 empty] | The investigation of pump did not start has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 code 1503 was reported incorrectly on manufacturer device report 1220648-2024-24691. new code was added to medical device problem code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
13168 | 1220648-2024-24690 | 20896782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20896782&pc=OZD | 11/26/2024 | 01/20/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568415 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant, and the pump was unable to traverse into the ascending aorta. the implant was aborted. no known patient harm or injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to the patient's anatomy.
13169 | 1220648-2024-24685 | 20897022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20897022&pc=OZD | 11/26/2024 | 01/28/2025 |  | 12/10/2024 |  |  | 11/26/2024 | 12/10/2024 |  |  |  |  | 12/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1740558 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella automated impella controller went black and restarted on its own. the aic was exchanged and there were no further issues. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the automated impella controller (aic) pump stop has been completed. data logs were reviewed finding multiple "purge pressure high" and "purge system blocked" alarms. however, no alarm or abnormal behavior was seen in the device logs on the mentioned date of complaint. additionally, no issues or alarms related to the reported complaint of the "aic going black and restarting on its own" were found in the device logs. an email was sent to the coordinator and complaint creator to clarify the exact date of the reported malfunction, the specific issue observed, and the aic involved. no response was received. the console was returned for further investigation. a thorough visual inspection of the internal circuitry did not reveal any issues. the console was then tested with a known good pump and purge system, during which no anomalies were observed. the reported issue of the aic going black and restarting on its own was unrelated to any faults within the aic. additionally, the "purge pressure high" and "purge system blocked" alarms were also not caused by a malfunction in the aic. in conclusion, no issues were found within the aic. d.9 date device returned/information was added. d.2 revised common device name since the initial manufacturer device report number was submitted in accordance with updated procedures.
13170 | 1220648-2024-24607 | 20877808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20877808&pc=OZD | 11/19/2024 | 01/28/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536837 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was leaking from the pressure reservoir. the amount was small. it was reported to be less than one milliliter per hour. it is suspected that the incident may have been caused by the lot of the pressure reservoir. the health hazards were reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. data logs were reviewed and there were no symptoms of leak or low purge pressure observed. the cause of the purge leak - pump issue was use related since no leak reproduced with return product testing.
13171 | 1220648-2024-24608 | 20877873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20877873&pc=OZD | 11/21/2024 | 03/10/2025 |  | 12/09/2024 |  |  | 11/21/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1847282 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following support, the nurse unplugged the automated impella controller (aic) from the wall by pulling on the cable. the wires subsequently pulled out of the plug. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - cabling issue has been completed. upon reviewing the data log it showed that on the complaint date, during support, the console log recorded ¿ac power disconnected ¿ alarm. most likely, at this point, while unplugging ac power from the wall, the wire could have pulled out of the plug. device analysis: this console was tested after arriving, the damaged ac power cable and missing plug were confirmed. the root cause of the wires being pulled out of the ac plug was most likely due to a defective cable. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13172 | 1220648-2024-24609 | 20877997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20877997&pc=OZD | 11/21/2024 | 01/09/2025 |  | 12/09/2024 |  |  | 11/21/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524241 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | INTRA-AORTIC BALLOON PUMP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported vascular damage. vascular repair was required for bleeding. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the vascular damage has been completed. the product was not returned for investigation, and a data log review was not required for this case. the physician believed that the vascular damage occurred due to the iabp placement prior to the impella insertion. the cause of the vascular damage issue was not determined without sufficient clinical information. the relation between the vascular damage and the impella device was unknown, as the injury was attributed to the use of a different device, specifically the iabp, which was placed prior to the impella insertion. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added d10-concomitant medical products.
13173 | 1220648-2024-24610 | 20878497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20878497&pc=OZD | 10/14/2024 | 12/09/2024 |  | 12/09/2024 |  |  | 11/23/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546180 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Per investigation results, the complainant reported that a patient on impella 5.5 support had pink/red urine with increased lactate dehydrogenase since insertion of the impella 5.5. the p-level was reduced and the urine had cleared. the hemolysis signs of suction and lab levels persisted until wean and explantation. the patient survived the explant. | [5 empty] | The investigation has been completed. the product was not returned. data logs were reviewed and suction alarms were observed. no motor current spike or flow issue outside of p-levels occurred. there was no indication of reported hemolysis issue found. the cause of the hemolysis issue could not be determined since the complaint device not returned for analysis and lack of clinical details. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13174 | 1220648-2024-24611 | 20878592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20878592&pc=OZD | 11/20/2024 | 12/30/2024 |  | 12/09/2024 |  |  | 11/20/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533262 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump support ongoing for a patient admitted with st elevated myocardial infarction. the pump support was ongoing while the team planned for a family decision as the care could not be escalated. the team had heparin in the purge and the patient was suffering from a gastrointestinal bleed. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13175 | 1220648-2024-24612 | 20878669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20878669&pc=OZD | 12/31/2023 | 12/09/2024 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  | 12/31/2024 |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298175 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Per investigation results, the complainant reported that a patients urine output was low and cranberry color. the medical doctor might reposition the impella to resolve the possible hemolysis, but first had the nurse give volume and keep an eye on it and ordered haptoglobin. the family made the choice to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation has been completed. no product was returned. data logs showed suction alarms throughout the case. placement signal and flows both were trending down. motor current also was a bit noisy and trending down. the patient was decompensating. the cause of the hemolysis was most likely patient condition related. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13176 | 1220648-2024-24613 | 20879092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20879092&pc=OZD | 11/21/2024 | 10/27/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553447 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella 5.5 support. it was reported that there was a small amount of purge fluid leaking from a location that should not leak. it was suspected that the incident may have been caused by the lot of the pressure reservoir. there was no harm to the patient and the patient remains on support. | [5 empty] | The investigation for the purge leak-pump has been completed. pump was not returned for investigation. no data logs were returned. as per the clinical information, the leak was observed from inside the plastic covering of the sidearm but no images of the leak were provided. clinical also mentions of a similar case where similar leak was observed but without the images or product of the complaint pump, the failure cannot be further investigated. therefore, the root cause of the purge leak issue was not determined since product was not returned and clinical information was not sufficient. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was returned for evaluation. pump was returned for investigation. the pump side-arm was isolated to reproduce leaking. the sidearm was pressure purged for over 90 minutes but no leaks were visible. the purge filter and reservoir was checked under keyance microscope for any minor leaks but no leak was found. the plastic covering of the sidearm was spotless with no traces of fluid leak. purge leak was not reproduced. no data logs were returned. as per the clinical information, the leak was observed from inside the plastic covering of the sidearm but no images of the leak were provided. the failure cannot be further investigated without data logs. therefore, the root cause of the purge leak issue was not determined since purge leak failure could not be reproduced and data logs were not returned. d9 device returned to manufacturer date added.
13177 | 1220648-2024-24614 | 20879271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20879271&pc=OZD | 09/15/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539099 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification from abiomed¿s long-term outcome and quality indicator (loqi) impella registry was received and noted the patient endured non-sustained ventricular tachycardia (nsvt) with a "possible" relationship to the impella device used. the patient with non-ischemic cardiomyopathy was successfully implanted with an impella 5.5 device via the right axillary artery access for mechanical circulatory support. approximately two days after the start of support, the patient experienced occasional nsvt runs when aggressively moving around, self-limiting. electrolyte replacement was decreased. adenosine was administered, but there was no response with sustained termination of the ventricular tachycardia. however, the patient was noted to be hemodynamically stable, and following, it was noted the patient did recover with no sequelae. the patient, denied a transplant at such time and remained on support as a bridge to left ventricular assist device (lvad). later, 24 days after start of support, the impella 5.5 was successfully explanted, and the lvad was placed. another notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding that the patient that endured acute blood loss anemia with a "probable" relationship to the impella device used: ¿h/h has persistently been dropping slowly, most likely in setting of impella insertion and hematoma formation at site. h/h 8.5/25.5 this am. platelets remain stable at 55. patient is asymptomatic, great bp off pressor support. secondary to open heart surgery, expected. hemoglobin on admit 8.3, lowest 7.3 (b)(6) 2024, 10.2 discharge.¿. the patient recovered with no sequelae. additionally, another notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient that endured thrombocytopenia with a "possible" relationship to the impella device used: ¿most likely consumptive secondary to impella placement and hematoma present at insertion site. platelet drop is within ~3-4 days, which is not c/w hit picture (and no prior heparin exposure).¿. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis, ventricular tachycardia (vt), and thrombocytopenia has been completed. data logs showed pump ran without issue during support. there were suction positioning alarms triggered during the case but resolved. product was not returned for review. the root cause of hemolysis, vt, and thrombocytopenia was unable to be determined as limited clinical details were provided and no product was returned for review.
13178 | 1220648-2024-24615 | 20879559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20879559&pc=OZD | 11/22/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532528 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, platelets were down from 90 to 60. there was a concern for thrombocytopenia. the patient was stable and the procedural outcome was successful. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the thrombocytopenia has been completed. impella cp was returned for evaluation. no damage to the pump was observed on visual inspection. data logs were reviewed and no alarms noted at time of reported issue. clinical details noted that on second day of impella support, patient's platelets went from 90 to 60 and there was a concern for thrombocytopenia. the cause of the issue was not determined due to insufficient clinical details.
13179 | 1220648-2024-24616 | 20880001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20880001&pc=OZD | 06/29/2024 | 01/08/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489363 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a "possible" relationship to the impella device used: the patient had acute loss anemia and was indicated per the medical doctor that the patient received twi units of blood due to hemoglobin dropping below 9 with hemodynamic instability. hemoglobin was found to be 7.5 the patient's hemoglobin rebounded sufficiently to 9.3, and additional units were not needed at that time. the patient was deemed back to baseline on lab draw a day later. the patient recovered with sequelae. the pump was explanted and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13180 | 1220648-2024-24618 | 20880084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20880084&pc=OZD | 11/22/2024 | 01/13/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient with a cardiac history suffering from st segment elevation myocardial infarction. the pump had supported for just about six hours when the pump was explanted. the bleeding was treated prior to explant by changing the dressing and placement of a mattress suture. the patient had a very large pannus which caused pressure on the insertion site. the explant was due to access site bleeding that could not be controlled. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site issue was most likely patient movement. following tilting of bed, dressing saturated in blood again and interventional cardiology fellows were unable to improve oozing with troubleshooting.
13181 | 1220648-2024-24617 | 20880153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20880153&pc=OZD | 11/22/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210713 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a damaged shipping carton. the automated impella controller (aic) was found to be damaged to the rear housing. the aic was pulled aside for investigation prior to service and repair. no patient was involved. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. the damage to the rear housing was confirmed. the root cause was likely due to mishandling during shipping. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13182 | 1220648-2024-24619 | 20880550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20880550&pc=OZD | 11/22/2024 | 01/09/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493011 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was noted that resistance was met due to extreme calcification and the tortuosity of the patient's vasculature. multiple peripheral balloon angioplasties were attempted coinciding with shockwave treatment, but the pump still failed to advance. both 25 centimeter and 13 centimeter sheaths were attempted. the implant was subsequently aborted and an intra-aortic balloon pump was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since there was extreme calcification and tortuosity.
13183 | 1220648-2024-24620 | 20880680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20880680&pc=OZD | 11/22/2024 | 01/06/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513188 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump being delivered to support a patient admitted in with cardiogenic shock and history of heart failure. the native anatomy prevented the impella 5.5 delivery and so the team placed an intra-aortic balloon pump for support. no harm or injury has been reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". section: warnings & cautions: warnings: "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." section: warnings & cautions: cautions: "dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve." section: warnings & cautions: warnings: "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance." impella 5.5 with smartassist. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." section: warnings & cautions: warnings: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance."
13184 | 1220648-2024-24621 | 20881075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20881075&pc=OZD | 11/22/2024 | 01/09/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552641 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the sterile sleeve of an implanted impella 5.5 pump tore during repositioning of the device. transparent film dressing was placed over the sleeve to maintain the occlusive covering. support was continued following troubleshooting. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. data logs are not relevant to the complication and the product was not returned for investigation. clinical narrative reports that the sterile sleeve ripped due to excessive force during pump repositioning. based on clinical details, the root cause of the product damage was determined to most likely be a use issue. h.3 if the device was not evaluated selection was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24621 and was revised.
13185 | 1220648-2024-24623 | 20881995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20881995&pc=OZD | 11/22/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548289 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a high risk percutaneous coronary intervention patient. the team made the choice to use the 7 french via the single access in the 14 french at the femoral artery. the 7 french sheath interacted/had resistance and the team observed the sheath had collapsed upon itself. the sheath was replaced without harm or injury and the support was proceeded. no patient harm was reported. | [5 empty] | The investigation for the introducer damage has been completed. the device was not returned for evaluation. without additional clinical details, the root cause of the introducer damage has been completed. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13186 | 1220648-2024-24622 | 20882487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20882487&pc=OZD | 11/22/2024 | 01/09/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569094 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fracture | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported complications were encountered during the initial delivery attempt of an impella cp device for mechanical circulatory support. upon attempting to gain access to the left femoral artery for impella placement, the wire broke inside the vessel. a hematoma formed and surgical intervention was required to repair the artery. one unit of packed red blood cells was transfused before the pump was successfully implanted for support. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the product damage and vascular damage has been completed. the product was not returned for investigation. a data log review was not required for this case. as per the given clinical details, the guidewire was break inside the vessel during placement, and a large hematoma was noted at the access site. no picture was provided for the guidewire damage, and the exact location of the fracture was unknown. the doctor took a surgical approach to remove the guidewire and repair the access site. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13187 | 1220648-2024-24624 | 20882594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20882594&pc=OZD | 11/22/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563031 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for a patient admitted in for a high risk percutaneous coronary intervention. the pump was placed but the team observed saturated flows. the team explanted the pump and replaced with an impella 5.5 pump. no patient harm was reported. | [5 empty] | The investigation for the saturated flow has been completed. no product was returned. log review showed a large motor current spike after insertion followed by saturated flows (5l/min during systole and diastole). there were some "impella position unknown" and "impella in aorta" alarms, but it was noted that the pump was in good position during the saturated flows. the root cause of the saturated flow was most likely biomaterial. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13188 | 1220648-2024-24626 | 20883080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20883080&pc=OZD | 11/22/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  | 09/30/2025 |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394602 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after one day of patient support, the implanted impella 5.5 pump began to experience high purge pressure and low purge flow. tissue plasminogen activator was added to the purge to troubleshoot the issue but one day later, the pump began to experience saturated flow. due to the ongoing issues, the decision was made to replace the pump. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the high purge pressure and saturated flow has been completed. no product was returned for evaluation. data logs from field were returned and rt log shows increase in purge pressure over approximately 21 minutes to a purge pressure of just under 1000mmhg. purge remains high for approximately 3 hours before additional purge increase to trigger "high purge pressure" alarms. no motor current spikes or speed deviations were noted at time of purge pressure rise. purge remained high for remainder of case. additionally, towards end of pump run, pump flow becomes fully saturated with motor current rising without a p-level change. clinical details noted that physician started tissue plasminogen activator (tpa) in purge solution after high purge pressure was noted. the root cause of the high purge pressure was most likely due to biomaterial build-up. the root cause of the saturated flow was most likely due to biomaterial. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
13189 | 1220648-2024-24631 | 20883096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20883096&pc=OZD | 11/23/2024 | 01/15/2025 |  | 12/09/2024 |  |  | 11/23/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591517 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in with multivessel coronary artery disease but declined coronary artery bypass grafting. it was reported that the access site had an ooze that prompted the team to place a femstop and two units of red blood cells given. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13190 | 1220648-2024-24625 | 20883539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20883539&pc=OZD | 11/20/2024 | 10/03/2025 |  | 12/09/2024 |  |  | 11/22/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025526737 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Tachycardia | Thrombocytopenia | [10 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient remained off anticoagulation due to having persistent thrombocytopenia. it is unknown if the thrombocytopenia was impella related. support was continued. furthermore, the patient had ventricular tachycardia (vt) throughout support which required cardioversion and medicines were needed. it is unknown if the vt was impella related. furthermore, the patient expired after 433.20 hours of impella support the relationship between the impella and cause of death is unknown. | The complainant also reported that the placement signal monitoring was disabled due to frequent "impella position unknown" alarms. |  |  |  |  | The investigation of thrombocytopenia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the cause of the thrombocytopenia and vf/vt/af/at was not determined due to insufficient clinical details. hemolysis was added as a failure mode since ldh levels spiked on 11/23 and there were concerns for hemolysis as a result. the log review showed no motor current spikes during the case that would indicate biomaterial ingestion. placement was confirmed with echo and the pump was in good position. the cause of the hemolysis was most likely patient condition related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. several "impella position unknown" alarms were reported. log review showed normal 5s optical parameters during the "impella position unknown" alarms. no other optical parameters were affected. the cause of the optical signal issue was not determined since product was not returned and insufficient clinical detail was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24625. d10 concomitant medical products and therapy dates updated. h4 device manufacture date corrected.
13191 | 1220648-2024-24630 | 20886410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20886410&pc=OZD | 11/23/2024 | 01/10/2025 |  | 12/09/2024 |  |  | 11/23/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025566138 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the patient had some non-sustained ventricular tachycardia (nsvt) runs over night. the impella 5.5 decreased from p7 to p6. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13192 | 1220648-2024-24629 | 20887880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20887880&pc=OZD | 11/23/2024 | 01/15/2025 |  | 12/09/2024 |  |  | 11/23/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562298 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and had a purge leak of the pump. the patient had a ventricular tachycardia event at home with one shock and was brought to the emergency department and admitted to the intensive care unit for inpatient ventricular assist device/transplant work up. after admission to the unit, the patient acutely decompensated over 24 hours prompting the heart failure team to consult with the cardiothoracic surgery for mechanical support with an impella 5.5 device. the patient was urgently taken to the operating room for successful axillary placement. the patient was then transferred to the cardiovascular intensive care unit in stable condition. approximately nine days after start of support, a leak was noticed around the purge sidearm. it was reported that the nurse found a small amount of purge fluid leaking from around the purge filter assembly. the purge pressure had fallen to lower 300¿s range and purge rate had increased by roughly 3 milliliters with no alarms. upon inspecting the assembly, the nurse squeezed the outer housing and it "snapped back together" dried off the outer housing, and tightly wrapped it with silk tape. the purge pressure was back up to upper 400¿s with a flow rate of 14 milliliters. due to the patient getting listed for a transplant, the decision was made to do an impella 5.5 exchange as support would be needed for the foreseeable future. the explant was successful, and the additional impella 5.5 insertion was also successful with no report of adverse effects to the patient. | [5 empty] | The investigation of purge leak ¿ pump has been completed. since product was not returned for investigation to attempt to reproduce the leak, the root cause of the purge leak could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." "clean the connector cable with 70% isopropyl alcohol."
13193 | 1220648-2024-24628 | 20888234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20888234&pc=OZD | 11/23/2024 | 01/15/2025 |  | 12/09/2024 |  |  | 11/23/2024 | 12/09/2024 |  |  |  |  | 12/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570726 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed hemolysis following implant of an impella cp pump. the condition was diagnosed via dropping haptoglobin and hemoglobin levels coinciding with a rising lactate dehydrogenase level. the support level was reduced, positioning was adjusted, and blood products were given, but the issue persisted. support was continued with the pump until explanted the following day, at which point signs of improvement in the condition were noted. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data was not returned for review. pump was returned for analysis. visual inspection found no abnormalities on the return pump. sufficient information related to patient condition is unknown and log data is not available to derive at a conclusive evidence for the failure. the root cause of hemolysis was unable to be determined as limited clinical details were provided and since no issues observed on return product.
13194 | 1220648-2024-24576 | 20864374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20864374&pc=OZD | 11/19/2024 | 01/02/2025 |  | 12/06/2024 |  |  | 11/19/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557411 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | Mitral Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had impella 5.5 support and during support, the staff attempted to reposition the pump due to it causing some mitral regurgitation. after the repositioning, a kink was noted as well as the plastic bunched up disrupting integrity of catheter shaft. additionally, it was reported that there was sterile sleeve damage. the pump could not be repositioned due to the damage to the catheter shaft, it will not advance through the touhy valve. support continued and the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (unites states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of positioning issues, heart valve damage, and product damage (sterile sleeve damage) has been completed. the returned product was investigated and the severe kinking/bunching in the catheter was observed. the cath lock had no issues advancing back and forth over a different section of the catheter. this confirms the suspected caused of the issue stated in the clinical narrative. based on clinical details and returned product, the root cause of the positioning issues was determined to be a cath anchor use issue. there are no mentions of the patient anatomy being abnormal or diseased. based on clinical details and data logs, the root cause of the heart valve damage was determined to most likely be improper positioning. based on clinical details and returned product, the root cause of the product damage was determined to most likely be a use issue.
13195 | 1220648-2024-24577 | 20864565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20864565&pc=OZD | 11/20/2024 | 01/22/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561334 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, the pump could not be repositioned well. the pump was successfully weaned and explanted. the patient survived the explant. | [5 empty] | The investigation positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. a5 race; white was inadvertently omitted from manufacturer device report 1220648-2024-24577. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13196 | 1220648-2024-24578 | 20864626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20864626&pc=OZD | 11/20/2024 | 01/03/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505068 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by minimal ooze and a hematoma. the patient underwent a washout of the axillary graft site and the hematoma was evacuated. some bleeding was also located and it was stopped. additionally, heparin was paused for that procedure and planned to be restarted afterwards. the patient continued on support and was reportedly stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13197 | 1220648-2024-24579 | 20864692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20864692&pc=OZD | 11/20/2024 | 01/14/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569324 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient experienced ventricular tachycardia. impella in ventricle alarm occurred and was accompanied by runs of ventricular tachycardia and a subsequent loss of pulsatility with associated hypotension. a representative assisted with repositioning via facetime and the pump was pulled back from 6.2 to 5 centimeters into the left ventricle. no further alarms or ectopy were noted and waveforms were appropriate. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined.
13198 | 1220648-2024-24580 | 20864974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20864974&pc=OZD | 11/20/2024 | 01/07/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566707 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding. additional information noted the 14x13 peel away sheath was placed, and once companion sheath was placed, bleeding was noted from hemostatic valve in peel away sheath. physician repositioned the companion sheath and bleeding stopped. no hematoma was noted. the patient was able to successfully finish the high-risk percutaneous coronary intervention and maintained stability throughout the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, the physician repositioned the companion sheath to a 2 o¿clock position on peel away and the bleeding stopped as per the clinical description provided.
13199 | 1220648-2024-24582 | 20865139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865139&pc=OZD | 11/15/2024 | 12/06/2024 |  | 12/06/2024 |  |  | 11/15/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The company investigator informed that a complainant had a patient that was on impella cp and automated impella controller (aic) support when there were purge issues on the aic with the heparin concentration. the aic was not removed/replaced. there were no problems with impella operation, but the patient's cardiac function did not improve and the died of multiple organ failure. therefore, the patient's peri-procedural condition was critical. there is no associated bleeding and the event was unrelated to the aic. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. the console was tested. while running the optical test pump, the bag change procedure was performed and the heparin value was updated from 10 iu/ml to 40iu/ml. during the second bag change, the console displayed the immediate previously updated heparin value of 40 iu/ml. also, the infusion history screen shows the same value as the one updated during the bag change procedure. unable to reproduce the reported failure mode. the console performed as expected. the root cause of the change in the heparin unit was not determined due to the inability to reproduce. since there was no evidence of automatic change in the heparin value, the reported issue could most likely be use-related.
13200 | 1220648-2024-24581 | 20865281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865281&pc=OZD | 11/20/2024 | 01/10/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523490 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, oozing was noted at the access site. suture was placed and site was redressed. additionally, the patient had 11 beats of ventricular tachycardia. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and tachycardia has been completed. the impella was not returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by adding the sutures and redressing the site. the root cause of the tachycardia could not be determined without additional.
13201 | 1220648-2024-24583 | 20865295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865295&pc=OZD | 11/20/2024 | 01/07/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565524 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, there were suction events all night despite recommendation to optimize position. alarms were silenced. trans-esophageal echo was done which showed impella was malpositioned. care team opted to not optimize patient at that time. chose to leave device suctioning through the evening and night. patient's condition did improve which allowed the device to be removed. additionally, the patient had major bleeding as a result of the coronary artery bypass graft (cabg) surgery. no heparin or anticoagulation given. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical details was not provided. the cause of the positioning issue was most likely use related, based on given clinical details.
13202 | 1220648-2024-24584 | 20865452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865452&pc=OZD | 11/20/2024 | 01/27/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella support when the automated impella controller (aic) failed. a controller failure (red alarm) appeared and the aic was replaced. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller failure red alarm has been completed. data logs were reviewed and found that after 15 days since the pump was started, the console displayed ¿purge pressure high ¿ alarm,¿ event #408 & ¿purge system blocked ¿ alarm,¿ event #409. rt log confirmed the purge pressure was more than 1000 mmhg, and there were no purge flow ticks. after 22 minutes, the console displayed ¿impella stopped (motor current high) [detected by epc] ¿ alarm,¿ event #13. this symptom aligns with the reported failure mode. then there was a pmd reset, and the console displayed the following events repeatedly: #13, #408, and #409 in addition to ¿pmd has detected a pump speed deviation ¿ alarm.¿ event #138. this console was returned for evaluation and upon functional testing, it was unable to reproduce the pump stop. no issue with the purge pressure accuracy. no evidence of dextrose deposits or residue around the purge motor piston. no problem was found with the console hardware. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). the root cause for the high purge pressure was most likely biomaterial, and the cause of the pump stop was most likely high purge pressure.
13203 | 1220648-2024-24586 | 20865504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865504&pc=PYX | 11/20/2024 | 01/06/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025503980 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp and an impella cp to support a patient admitted in acute myocardial infarction/cardiogenic shock. the team had issues in placing the impella rp which was replaced by a second impella rp. the impella rp was placed after multiple access attempts/sticks to the femoral. the impella rp access site was oozing during the support and was treated by a mattress suture, manual pressure, and pressure dressing. the support was for 72 hours and was then weaned and explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13204 | 1220648-2024-24585 | 20865649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865649&pc=OZD | 11/20/2024 | 01/10/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569289 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that two days into impella 5.5 support, it was discovered that the patient experienced a stroke. a computed tomography angiography identified an m2/m3 occlusion coinciding with a subarachnoid hemorrhage. support was continued with the implanted pump until planned explant six days later and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. h.6 code 4627 was reported in error on the initial manufacturer device report number 1220648-2024-24585. the device was explanted as planned and not a result of the event. a new code has been added.
13205 | 1220648-2024-24587 | 20865699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865699&pc=OZD | 11/21/2024 | 10/02/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1419816 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | Per failure investigator, the automated impella controller (aic) logs showed a controller error alarm. the aic has been requested for evaluation. | The complainant also reported a placement signal not reliable alarm. |  |  |  |  | The investigation for the console (aic) yellow alarm has been completed. console logs from the reported day of event are consistent with complaint and show about 40 alarms placement signal not reliable (#1036, due to invalid optical signal) and over 20,000 of event 1036 (psnr without alarm) throughout the case, over 100 alarms placement signal not reliable (#1048, signal out of range) and a controller error alarm (#1044, opm comm invalid). rtlog and ltlog data shows erratic characteristics of opm signals: noisy snr and placement, barcoding, occasional periodicity of snr, and low snr. the cause of the issue was not determined since product was not returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The automated impella controller (aic) was received for investigation and after evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-24587 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24587. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. console logs from the reported day of event are consistent with complaint and show about 40 alarms placement signal not reliable. console was returned to abiomed where it was subject to inspection end testing. some contamination and debris were found on the optical pump connector fiber, as well as some localized anomalies of the optical bench detector (ccd). reported placement signal issues were most likely due to combination of factors: defective optical bench (¿ccd defect¿), compromised blue receptacle (debris creating air gap), and either pump-related issue or pump-patient interaction. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24587. d6a updated with additional information d8/d9 device returned to manufacturer updated
13206 | 1220648-2024-24592 | 20865738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865738&pc=OZD | 11/21/2024 | 01/13/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569605 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp escalated to an impella 5.5 pump. the pump was placed as the patient was critical and planning/waiting for a transplant. the patient had a cerebrovascular accident/stroke on day 10 of impella 5.5 support. the etiology is unknown. the patient continued support and was reportedly stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
13207 | 1220648-2024-24588 | 20865795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865795&pc=PYX | 10/13/2024 | 09/30/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551264 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Th engineer reviewed data logs and found that purge pressure high alarms occurred on impella rp flex soon after arrival of the unit from the cardiac catheterization lab. after troubleshooting the line for any kinks, the cassette was changed by the nurse. the problem was not resolved. tissue plasminogen activator (tpa) protocol was begun. purge flows and pressure with slow improvement. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | Us complainant reported the patient was on impella rp support. while on support, nurse reported ongoing placement signal not reliable alarms. |  |  |  |  | The investigation for the high purge pressure has been completed. no product was returned. there were small spikes in motor current and placement signal right before the hpp and also during the event. flows became saturated. the high purge pressure lasts for about 13 hours. tpa protocol was started and the clinical stated it slowly resolved. the purge pressure remained stable for the rest of the duration of the case. the root cause of the high purge pressure is most likely due to biomaterial. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned. there were small spikes in motor current and placement signal right before the hpp and also during the event. flows became saturated. the high purge pressure lasts for about 13 hours. tpa protocol was started and the clinical stated it slowly resolved. the purge pressure remained stable for the rest of the duration of the case. the root cause of the high purge pressure is most likely due to biomaterial. additionally, 19 hours in the placement signal started to drift upwards with the flow drifted downwards. the drift lasted for 59 hours before the placement signal went out. alarm 1052 was triggered as well indicating the dp sensor was out of range. the root cause of the placement signal issue is due to dp sensor drift. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
13208 | 1220648-2024-24591 | 20865834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865834&pc=OZD | 11/21/2024 | 01/07/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167411 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, there was an issue with the purge cassette. purge system was properly connected but purge solution was not priming in the purge tubing. impella consoles were changed but issue did not resolve. hospital had to open a new purge cassette which worked without any issues. no patient harm has been reported. | [5 empty] | D.2 revised common device name since the initial manufacturer device report number 1220648-2024-24591 was submitted in accordance with updated procedures. h.6 health effect - impact code 4627 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24591. upon review of the provided information, the noted issue was resolved by changing the cassette. there is currently no allegation of a malfunction or serious injury associated with the automated impella controller (aic). therefore, a regulatory report against the aic is not necessary. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13209 | 1220648-2024-24589 | 20865877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20865877&pc=OZD | 11/21/2024 | 12/01/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the automated impella controller was replaced due a yellow controller alarm. there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported issue has been completed. data analysis: since the console was started, after 1 day, the console displayed ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm¿ event # 1045. which resolved automatically after a minute. this confirms the reported failure. real-time (rt) logs confirmed that all signals received from the optical bench (snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. device analysis: this console was tested after arriving in fs. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in the misalignment of data communication packets received by epc. the aic sw operated as designed. it was found that the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). h6 investigation type, findings and conclusions.
13210 | 1220648-2024-24590 | 20866090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20866090&pc=OZD | 11/20/2024 | 12/01/2025 |  | 12/06/2024 |  |  | 11/20/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1569441 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the automated impella controller (aic) had an abnormality during a self-check when the aic was launched. this issue occurred during preparation and was replaced with another aic. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13211 | 1220648-2024-24593 | 20866230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20866230&pc=OZD | 11/21/2024 | 01/08/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464671 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while removing the 5.5 graft, anastomosis tore. the doctor performed repair to the artery. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. no product was returned and logs are not applicable. an 8mm graft was utilized instead of the ifu (document no. 10003049 ri) recommended 10mm graft. a smaller diameter graft than 10 mm will not fit over the introducer. the root cause of the vascular damage was most likely use-related.
13212 | 1220648-2024-24594 | 20866315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20866315&pc=OZD | 11/21/2024 | 01/13/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552280 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella cp. 14frx13 peel away was placed and left ventricle (lv) was accessed but wire and pigtail noted to be shallow. the doctor repositioned wire and exchanged to 0.018 wire and impella placed into lv. the wire was removed, and the impella started. upon starting impella, immediate suction was noted due to impella being too shallow and the doctor tried to reposition but unfortunately pulled it out of the ventricle into aorta. impella was taken out and re-accessed lv and replaced impella but the doctor was still unable to place wire in apex. unfortunately, impella unable to be in optimal position. upon placing on support, suction noted again. the doctor thought it was due to just patient anatomy of ventricle. the doctor opted to just take out impella and proceed with percutaneous coronary intervention. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the pump was unable to placed in the right position because of resistance caused at innominate artery.
13213 | 1220648-2024-24595 | 20866633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20866633&pc=OZD | 11/21/2024 | 01/10/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548288 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient went into ventricular tachycardia, at the time no a-line in place, no pulses, cardiopulmonary resuscitation initiated. after return of spontaneous circulation, impella placement alarm in aorta. bedside echo was done, inlet at aortic valve with pigtail still in ventricle. impella was 6 cm, was then pulled back to 3.5 but appeared to be against the left ventricular wall. repositioning was unsuccessful. the impella pulled back to the aortic valve and attempted to advance with no luck. pump was removed and replaced using side port. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and tachycardia has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the tachycardia was not determined. the root cause of the positioning issue was most likely use related to improper technique.
13214 | 1220648-2024-24596 | 20866713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20866713&pc=OZD | 11/21/2024 | 12/27/2024 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025572831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, while advancing the pump, the left ventricle (lv) perforated by pump. pump retracted and the perforation was repaired. while in the operating room, chest tubes poured out about two liters of blood. the chest was re-opened to look for surgical bleed. no direct source was identified. clotting factors were given, surgicel was placed on chest wall, and chest wall was cauterized to control bleeding. the patient required multiple units of blood products to replace lost fluids. at end of the procedure, the impella measured 3.5cm in lv. due to the patient¿s condition, the doctor did not want to reposition. the patient survived the event. | [5 empty] | The investigation for the ventricle perforation and positioning issues has been completed. the ventricle was perforated during the first time attempting to advance the pump. based on the clinical details, the root cause of the ventricle perforation was most likely due to use as it occurred during insertion. at the end of the procedure the impella measured 3.5 cm in the lv but the md did not want to reposition due to the condition of the patient. the root cause of the positioning issue cannot be determined as no other information was provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist: system with automated impella controller: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13215 | 1220648-2024-24599 | 20867066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20867066&pc=OZD | 11/21/2024 | 01/06/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551720 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during the implant procedure positioning difficulty was encountered that appeared to have contributed to the patient's ventricular tachycardia. the patient was turned down for coronary artery bypass graft surgery. the patient was implanted with an impella 5.5 for planned percutaneous coronary intervention (pci) a couple of days following the impella implant. during the implant, the positioning of the impella 5.5 was a challenge as it kept getting snagged in the mitral apparatus with a smaller left ventricle. eventually, the impella 5.5 device was pulled back, freed, and repositioned appropriately. however, during repositioning, patient went into ventricular tachycardia requiring shocks and a few chest compressions. the ventricular tachycardia improved with proper positioning. the impella 5.5 device ramped up to performance level p-6, and the patient was transferred to the intensive care unit without further issue. the patient was reported to be stable following this event. four days later, the patient underwent the planned pci and remained stable throughout. the patient remained on the impella support post-pci for an additional five days before being weaned and device explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of positioning issue was most likely patient condition as the patient has a small left ventricle.
13216 | 1220648-2024-24598 | 20867332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20867332&pc=OZD | 11/21/2024 | 01/13/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591514 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp and during support there was access site bleeding, noted by the oozing. this was treated by applying pressure dressing and a femostop. additionally, the intensive care unit team tied a purse string suture around insertion site. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, icu team tied a purse string suture around insertion site, oozing decreased as per the clinical description provided.
13217 | 1220648-2024-24602 | 20867398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20867398&pc=OZD | 11/21/2024 | 02/13/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589311 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support due to st elevated myocardial infarction with left main and right coronary artery stenosis. while on support, distal pulses were not palpable and the color of the impella leg was different. heparin was paused. partial thromboplastin time was greater than 120. the patient was overall in poor condition and expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13218 | 1220648-2024-24600 | 20867939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20867939&pc=OZD | 11/21/2024 | 01/29/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  | 06/30/2025 |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369171 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the oozing at the groin site. the bleeding was resolved by giving the patient one blood bag per day as hemoglobin levels significantly decreased. additionally, the patient was not able to be weaned from support and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 manufacturing site facility name was reported incorrectly on manufacturer device report 1220648-2024-24600. g1 manufacturing site fax number was reported incorrectly on manufacturer device report 1220648-2024-24600.
13219 | 1220648-2024-24603 | 20868029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20868029&pc=OZD | 11/21/2024 | 01/08/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560057 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced complications during impella cp support. the patient began to show signs of hemolysis with low haptoglobin and hemolyzed labs. the patient was given three units of blood and one unit of fresh frozen plasma around this time and plans were made to wean the device. upon eventual explant, complications were encountered during vascular closure and a perforation of the iliac artery occurred. it was documented that a wire being used for dilator access kinked and perforated the artery. surgical intervention was performed to repair the site. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage and hemolysis has been completed. the pump was not returned for investigation. data logs were not available as well. without these it is not possible to investigate the failure further. therefore, the root cause of the hemolysis issue was not determined since the product and data logs were not returned. as per the clinical information provided, the team lost control of the access site leading to the wire getting kinked and harming the iliac. therefore, the root cause of the vascular damage issue was use related to improper technique. a.5 ethnicity and race are were reported on the initial manufacturer device report number 1220648-2024-24603 but are unknown.
13220 | 1220648-2024-24606 | 20868262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20868262&pc=OZD | 11/21/2024 | 12/30/2024 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support ongoing for a patient admitted in with acute myocardial infarction / cardiogenic shock. the pump was placed but the left ventricle (lv) morphology was boot-shaped. the lv prevented appropriate positioning, and the team was unable to properly place the pump. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the product was not returned. based on the clinical details provided, the root cause of the positioning issues is most likely due to the patients condition as they had a decompressed, boot shaped lv.
13221 | 1220648-2024-24601 | 20868331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20868331&pc=OZD | 11/21/2024 | 01/03/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569710 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that in the catheterization lab for an impella cp procedure, the impella catheter was plugged in and the automated impella controller showed an impella catheter defective alarm. troubleshooting was done but was unsuccessful. a new impella cp was opened and placed. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the pump did not start has been completed. clinical details state that the "impella defective" triggered when the pump was connected to the console. data logs were not returned for review. returned product was investigated and the pump was returned in out of box condition. the red lumen was still in the cannula and there were no visual abnormalities. the pump was connected to a programmer and the software recognized that a pump was connected. then, the pump was connected to the console and it recognized the pump's serial number and the first step of case start displayed on the screen. the pump also passed electrical testing. since data logs were not available and the complication did not reproduce, the root cause of the pump not detected could not be determined.
13222 | 1220648-2024-24597 | 20868444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20868444&pc=OZD | 11/21/2024 | 06/16/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567135 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella cp and during support had a gastrointestinal bleed. this was treated by administering the patient subtherapeutic levels of heparin. the patient survived the event. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured persistent hematuria with a "possible" relationship to the impella device used: ¿patient presented to ed with gross hematuria from foley and pelvic pain. noted to have clots in urine as facility did not have orders to flush foley. patient had cystoscopy with clot evacuation on (b)(6). foley removed on (b)(6). additionally, patient was treated for cauti (e. faecalis, e. coli) with 14 day course of ceftriaxone. he was also on vancomycin with eot (b)(6)." the patient recovered with no sequelae. |  |  |  |  | The investigation for the vascular damage has been completed. no product was returned for analysis. the clinical details do not mention excess force or any underlying patient condition. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | B5 revised to include additional information received from the investigation site. h6 revised to include additional information received from the investigation site. the investigation results for infection/sepsis and hemolysis are not complete. should the device or any new information be received, a supplemental mdr will be filed. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ infection: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ g2 report source; study was missing from manufacturer device report 1220648-2024-24597.
13223 | 1220648-2024-24605 | 20868892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20868892&pc=OZD | 11/21/2024 | 01/08/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563965 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the attempted implant of an impella 5.5 device for mechanical circulatory support (mcs) to a patient with cardiogenic shock from an acute myocardial infarction was unsuccessful. the patient was a transfer with an intra-aortic balloon pump (iabp). the decision was made to escalate from the iabp to an impella 5.5 and work up for coronary artery bypass graft surgery. the right axillary artery was accessed, graft sewn and kit in place. the impella 5.5 was inserted but met with heavy calcium at the subclavian artery. the decision was made to abort the impella 5.5 implant. the patient's status following the event is unknown. however, there was no report or indication of adverse effects to the patient given this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of the heavy calcification in the vessel and the unable to pass through the vasculature. d6a and d6b were missing from manufacturer device report 1220648-2024-24605.
13224 | 1220648-2024-24604 | 20871158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20871158&pc=OZD | 11/21/2024 | 01/03/2025 |  | 12/06/2024 |  |  | 11/21/2024 | 12/06/2024 |  |  |  |  | 12/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp and impella rp for mechanical circulatory support and following had high purge pressure and a gastrointestinal bleed. the patient was admitted to the intensive care unit where the patient was intubated upon induction pulseless electrical activity arrest with 30 minutes of cardiopulmonary resuscitation. after successful return of spontaneous circulation, the physician successfully placed the impella cp to the right femoral artery without complications. left heart catheterization was completed and diffused disease with severe lesion in the right coronary artery (rca) was identified. three stents were placed to the rca with no intervention to the left system. after next steps, the physician proceeded with the impella rp implant. same day, post procedure, a critical purge alert triggered and dextrose 5% in water was switched to tpa 3mg in 50cc of sterile water. day 4 of support the site was oozy with dressing change. a gastrointestinal (gi) bleed was present over the night. heparin taken out of the impella rp purge replaced with sodium bicarb. it was noted systematic heparin had started was lowered that morning to 550 from 850. overnight, the impella cp was repositioned, and the patient was given two units of blood and one unit of platelets. the next day, impella rp was weaned to performance level p-4. the gi bleed slowed down, and sodium bicarb of the impella rp purge was switched back to heparin due to purge flow dropping past 12 hours. the patient had received two units of blood and one unit of platelets overnight with a possibility of another unit of blood this day. an echocardiogram was completed to see if the patient could be weaned from both devices. next day, now day 5 of support, both devices were explanted, impella rp at the bedside and the impella cp in the catheterization lab. the patient was noted to be stable following the event. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella rp serial number: (b)(6) with date of event 24-nov-2024 and identifier ending in (b)(4). | The investigation for the high purge pressure, vascular damage, and access site bleed has been completed. data logs were reviewed and the only purge pressure high/system blocked alarm triggered on 11/24. at this time, there is a 36 second rise in purge pressure until the alarm triggers. it resolves in 1.5 minutes on its own. that being said, on the reported date of the complication, purge pressure does gradually rise from case start over a 3 hour period as high as 1113 mmhg. there are no motor current spikes before the rise and it starts trending up as soon as the pump starts and the motor current and ps are variable during this period. the complication resolves over a 12 hour period. returned product was investigated and there was no clear abnormalities upon visual inspection. there were signs of potential biomaterial in the purge gap, but it was too tight to confirm. next, the pump was connected to a console in the lab and run without any high purge pressure alarms for 15 minutes. a window was cut in the catheter to visualize the purge lumen just proximal to the motor housing and there was no blood ingress. then, the catheter was cut just distal to the motor housing and the purge pressure immediately dropped to zero. this all suggests that issues with the purge lumen were not the cause of the complication. finally, the motor housing was torn down and there was potential signs of biomaterial residue on the rotor, but no clear evidence. all of these findings align with the clinical information provided that tpa was infused into the purge and resolved the complication. based on all the above details, the root cause of the high purge pressure was determined to most likely be biomaterial. due to lacking clinical details, the root cause of the access site bleed and vascular damage could not be determined.
13225 | 1220648-2024-24556 | 20852390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20852390&pc=OZD | 11/19/2024 | 10/02/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554599 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella cp support. upon placement with the wire was removed and the impella started, the placement signal disappeared, and the placement signal unreliable message occurred. patient had calcification on their valve. the patient was managed by flow throughout the case but was unable to verify placement signal during the entire case. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Upon further review of this event, the reported complaint of placement signal issues is not a reportable malfunction and does not require medical device reporting. the initial manufacture mdr report was submitted in error. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. due to lack of product or data logs returned, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h3 device evaluated by manufacturer updated.
13226 | 1220648-2024-24557 | 20852639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20852639&pc=OZD | 11/18/2024 | 01/27/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540050 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella cp. the pump could not be placed due to stenosis in the abdominal aorta. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to be delivered due to stenosis in abdominal artery. corrections to the initial mfr report: d4 corrected udi number.
13227 | 1220648-2024-24558 | 20852807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20852807&pc=OZD | 11/19/2024 | 01/07/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there were delivery issues as the pump was unable to get across the patient's valve even after trying with multiple catheters and wires and after decreasing extra corporeal membrane oxygenation flow. the decision was ultimately made to abort the procedure. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy.
13228 | 1220648-2024-24559 | 20852979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20852979&pc=OZD | 11/19/2024 | 01/07/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569082 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, an echocardiogram (echo) revealed clot on end of distal end of impella pump. field rep reported that secondary echo showed no clot on impella. it is unclear if the first echo was incorrect. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13229 | 1220648-2024-24560 | 20853821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20853821&pc=OZD | 11/19/2024 | 01/06/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139667 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complaintant reported a console was received damaged. there was no patient use. biomedical engineer reported that the automated impella controller was boxed per instructions and stated that there was no damage when it was packaged. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. the console was returned with two hairline fractures along the top of the console that did not affect console performance the console otherwise performed as expected.
13230 | 1220648-2024-24562 | 20854236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20854236&pc=OZD | 11/19/2024 | 01/08/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560867 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the attempted implant of an impella 5.5 device for mechanical circulatory support (mcs) to a patient with cardiogenic shock from an acute myocardial infarction was unsuccessful. the patient arrived at the hospital with shortness of breath, was admitted, and continued to worsen. the team decided to implant an impella 5.5 device. however, the case was aborted. it was indicated patient support incomplete, and the patient was stable following the event. no further information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition since the pump could not be implanted due to patient's anatomy. b.7 additional information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24562. h.6 code 4627 should of been reported on initial manufacturer device report number 1220648-2024-24562 instead of code 2199.
13231 | 1220648-2024-24561 | 20854318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20854318&pc=OZD | 11/19/2024 | 01/09/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563136 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a patient admitted in with cardiac and systemic comorbidities. on day 2 of the support a bleed was observed at the aortotomy and graft anastomosis site. the bleed was treated by suture and blood. the pump was supporting when after 9 days the team explanted the pump with concerns for thrombus. the pump had purge and saturated flows which the team believed to be due to clot ingestion. when explanted the team did observe thrombosis. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by applying sutures at the site. as the necessary information was not provided, the root cause of the thrombus was not determined. h6 health effect - clinical code 1888 was inadvertently not included in initial manufacturer device report 1220648-2024-24561.
13232 | 1220648-2024-24565 | 20854627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20854627&pc=OZD | 05/28/2024 | 01/07/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468771 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Renal Failure | Sepsis | Heart Failure/Congestive Heart Failure | [9 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured end-stage renal disease, chest pain, sepsis, acute on chronic congestive heart failure with an "unknown (undetermined)" relationship to the impella device used. the patient was admitted for episodic ventricular tachycardia and heart failure with reduced ejection fraction. the patient was stabilized in intensive care unit initially but subsequently presented with worsening heart failure symptoms including elevated central venous pressure, decreased cardiac output/cardiac index, and hypotension. following in the evening, the patient was implanted with an impella cp device support for stabilization of cardiogenic shock. next day after start of support the patient was noted to be hemodynamically stable on impella support at p-8 with 3.5 l/min of flow. two days later, the patient was taken back to the catheterization lab for impella cp explant (reason unknown). the patient tolerated wean of impella down to performance level p-2. re-access port used to place wire and retain access, impella turned off, and explanted with good bleed back. access site hemostasis was obtained, and the patient was transferred back to the unit. sometime thereafter, according to the registry, the patient endured end-stage renal disease, chest pain, sepsis, acute on chronic congestive heart failure. the patient went into cardiac arrest and despite best efforts expired (actual date of death is unknown). | [5 empty] | The investigation for the reported renal failure, heart failure, and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of renal failure, heart failure, and sepsis could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13233 | 1220648-2024-24564 | 20855026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20855026&pc=OZD | 08/12/2024 | 02/13/2025 |  | 12/05/2024 |  |  | 08/12/2024 | 12/05/2024 |  |  | 12/31/2024 |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292105 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intracranial Hemorrhage | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered a brain bleed during impella 5.5 support. heparin was being administered within the purge and an extensive subdural hematoma was noted at the time of the event. the patient passed away. it is unknown if the complications were exacerbated by impella 5.5 support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cerebrovascular accident (cva) has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
13234 | 1220648-2024-24571 | 20855383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20855383&pc=OZD | 11/20/2024 | 01/08/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484450 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a surgical patient suffering from cardiogenic shock. the placement and access was done without company representative. the team later observed a groin site bleed/hematoma that was treated by manual pressure, blood products, and fluids. the team achieved hemostasis only with the placement of the gore dry seal. the pump was replaced after just over one hour of use and escalated to an impella 5.5 pump. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13235 | 1220648-2024-24570 | 20855889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20855889&pc=OZD | 04/16/2022 | 05/27/2025 |  | 12/05/2024 |  |  | 04/16/2022 | 12/05/2024 |  |  | 09/30/2023 |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | Asystole | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had a right coronary artery stent placed before the medical team decided to implant the impella cp to stabilize the patient for transport to another hospital. during transport, the patient had multiple episodes of pulseless electrical activity and a loss of pulsatility, which was treated with epinephrine and atropine. after arrival at the new hospital, right heart catheterization was placed and it was determined the patient was in right heart failure. the medical team then decided to implant an impella rp for increased support of the patient. prior to leaving the catheterization laboratory, the patient had an episode of ventricular tachycardia which was successfully resolved with defibrillation. while on impella cp and impella rp support, the patient experienced multiple episodes of ventricular tachycardia and ventricular fibrillation that required defibrillation to resolve. the patient was reported to have gone into asystole a couple of times but return of spontaneous circulation was achieved spontaneously and without intervention. the patient was successfully weaned from the impella rp to continue on impella cp and extracorporeal membrane oxygenation for mechanical circulatory support. later that day, the patient¿s family made the decision to withdraw care, and the patient expired while on impella cp support. there is not enough evidence to exclude the impella cp as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ note: this is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-18309 for the impella rp pump, with serial number (b)(6). this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13236 | 1220648-2024-24573 | 20855916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20855916&pc=OZD | 11/20/2024 | 01/13/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565594 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after nine hours of impella cp support, the impella cp became malpositioned and was exchanged for a new impella cp with smartassist. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue has been completed. no product was returned for evaluation. data logs were reviewed and rt logs showed "impella in aorta" alarms were triggered with a loss of pulsatility in motor current and pump flow. alarm was able to resolve with return of lv waveform becoming ventricular again. however, the following rt log shows additional "impella in aorta" alarm, with lv waveform becoming completely aortic and not resolving for remainder of case. clinical details noted that pump became malpositioned and was explanted and exchanged for a new cpsa. t-b lock was tightened and blue butterfly were secured. no patient movement was noted at time of positioning issue. the root cause of the positioning issues cannot be definitively determined as limited clinical details were provided.
13237 | 1220648-2024-24566 | 20855989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20855989&pc=OZD | 11/19/2024 | 01/15/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1744643 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a controller failure occurred during impella 5.5 support resulting in a pump stop. the automated impella controller was swapped and support was successfully reinitiated with the implanted pump. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the console (aic) pump stop has been completed. console logs from the reported day of event are mostly consistent with complaint and show that after 6 days of uninterrupted hemodynamic support, after the staff reduced p-level to p-1, the pump stopped and ¿impella stopped¿ alarm (#119, due to currents mismatch) was presented on console screen. multiple manual attempts to restart the pump were unsuccessful. during aic data analysis, couple other irregularities, unrelated to the complaint, were noted: ¿piston blocked¿ event and pud communication issues during initial deairing (from which console self-recovered), and multiple interruptions of data streaming throughout the case. data logs from the backup console were not available for analysis. the console was returned for evaluation and functional testing was performed with a known-good engineering pump, but the issue of pump stop was not reproduced. the ¿piston blocked¿ issue was most likely caused by purge motor shaft being momentarily immobilized by solidified purge fluid residue (issue self-resolved). rlm data streaming issues were caused by unexpected rlm reboots, most likely due to microsd card corruption. both of these issues were unrelated to the complaint. the root cause of pump stop could not be determined. | The device was returned for evaluation. upon completion of the investigation, a supplemental will be filed.
13238 | 1220648-2024-24572 | 20856104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20856104&pc=OZD | 11/20/2024 | 01/08/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523077 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp was placed post percutaneous coronary intervention for cardiogenic shock via the right femoral artery. however, the physician attempted to snap and peel away the peel-away sheath before it was completely out of the patient. this caused the impella cp to be pulled back into the aorta. the staff was unable to get the impella cp back across the valve. the impella cp insertion was aborted and an intra-aortic balloon pump was placed via the left femoral artery. the patient remained stable and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. no product was returned for evaluation. the logs from field show "impella position in aorta" alarms with a loss of pulsatility in both motor current and pump flow. clinical details noted that physician attempted to snap and remove peel-away introducer sheath before it was completely out patient and it caused pump to be pulled back into aorta. the root cause of the positioning issues was most likely a use related issue.
13239 | 1220648-2024-24563 | 20856186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20856186&pc=OZD | 11/19/2024 | 01/09/2025 |  | 12/05/2024 |  |  | 11/19/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569323 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhagic Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a hematoma at the access site. blood products were given to the patient and anticoagulation was held until neurological assessment was obtained due to extensive cardiopulmonary resuscitation time. furthermore, the patient's lactate dehydrogenase was 773. two units of packed red blood cells were provided to the patient. also, the patient had a computed tomography scan that found that the patient had a small subdural hematoma. the staff monitored and support was continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding, hemolysis, and stroke has been completed. the pump was returned for investigation. no physical abnormalities were seen on the returned product. biomaterial was found wrapped around the impella, and it could not be removed. the pump was tested in a bucket of water, and low pump flow was reported. the data log shows the trending placement signal and pump flow while running on a steady p-level. additionally, there were some small spikes reported in the motor current during steady p-level operation with reported low pump flow. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. the cause of the hemolysis issue was most likely biomaterial, based on given returned product investigation and data log review. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. d9 date device returned to manufacturer added.
13240 | 1220648-2024-24567 | 20857067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20857067&pc=OZD | 11/20/2024 | 01/08/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570579 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the next day after impella cp was placed, impella flows were low, and p-level could not go above p2 without suction alarms. repositioning was attempted at bedside however this was unsuccessful. the patient was taken to the cardiac catheterization lab to attempt repositioning, but impella came out of the ventricle and was unable to recross the aortic valve. the impella was exchanged. the patient is stable. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. no product was returned for evaluation. data logs were reviewed and rt log at end of case shows "impella in aorta" alarms with both lv and ao placement signals having an aortic waveform. additionally, towards the end of the rt log, pulsatility of both motor current and pump flow decreased. clinical details noted that pump "popped out" of ventricle and was unable to re-cross av. it was noted that blue butterfly was secured and tb was locked. the root cause of the positioning issues cannot be definitively determined as limited clinical details were provided. added h6 impact code 4628.
13241 | 1220648-2024-24574 | 20857551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20857551&pc=OZD | 11/20/2024 | 01/07/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542822 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient with cardiomyopathy had impella 5.5 implanted for over 30 days and over the past two days pfhgb increased and labs were hemolyzing. the 5.5 was attempted to be repositioned but unsuccessful. the impella was toward the mitral valve. tissue plasminogen activator was also unsuccessful. the impella was exchanged and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and positioning issues has been completed. pump was not returned for investigation. data logs were investigated. there were no positioning issues, suction alarms or motor current spikes observed. the clinical information provided is limited and data logs are inconclusive. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information was provided. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided.
13242 | 1220648-2024-24569 | 20857721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20857721&pc=PYX | 11/20/2024 | 01/03/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  | 05/31/2025 |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024345489 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues were encountered. the impella rp pump was inserted and zeroed but a kink in the cannula was noted and there was difficulty crossing the pulmonic valve without bowing in the cannula. the pump crossed but fell into the right ventricle shortly after the wire was removed and the pump started. multiple attempts were made to re-advance the pump into the left pulmonary artery. the attempts were unsuccessful as a kink in cannula was noted again. as a result, the pump was aborted and a new device was opened due to the presumed damage to the cannula. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy.
13243 | 1220648-2024-24568 | 20858031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20858031&pc=OZD | 11/20/2024 | 01/06/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573942 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was transferred from an outside hospital on impella cp support post high-risk percutaneous coronary intervention and escalated to an impella 5.5 device which was successfully completed. the patient was transported back to the unit, and while on support, the nurse noticed purge fluid leaking from purged side arm, which was determined to be from the connection. the purge cassette was changed and it was observed the luer connector was cracked and leaking. plumbers tape around the luer connector was used and tightened with new line. the impella cp line was used so that the next purge cassette change would be from the cp line connection to the cassette to avoid further manipulation at the fracture site. it was reported there was no change in purge pressure or flow, pump remained operational, the patient was unaffected and stable following the event. | [5 empty] | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. data logs were reviewed and when looking over the whole case there are no clear issues with the purge. there were 7 triggers of the purge pressure low alarms and 1 air in purge alarm, but all of these events align with purge cassette/bag changes. additionally, none of them triggered on the reported date of the issue. when looking at the reported date of the issue, the only event that potentially could have been interpreted as abnormal by the user was a drop in purge pressure from 500 mmhg to 400 mmhg. however, no alarms trigger, and the purge pressure remains stable and within specification from that point forward. this does align with the comment in the clinical narrative that there were no changes to purge pressures or flows. returned product was investigated and the sidearm luer was confirmed to have been wrapped in plumbing tape. all the tape on the pump was removed and cracks in the yellow luer were observed. the pump was then connected to a console and left to run for 3 hours and there were no leaks noted. when reviewing the data logs of the run in the lab no alarms triggered and purge pressures and flows were within specification. based on data logs and returned product, the root cause of the reported purge leak was determined to most likely be a damaged yellow luer due to improper technique applying an excessive side load to the luer. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock states the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
13244 | 1220648-2024-24575 | 20858422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20858422&pc=OZD | 11/20/2024 | 01/08/2025 |  | 12/05/2024 |  |  | 11/20/2024 | 12/05/2024 |  |  |  |  | 12/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562169 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support and during support, the patient was taken to the operating room for a washout. the clinical representative confirmed that the wash out was not impella related. the patient had a washout of the chest from the patient¿s coronary artery bypass graft. the physician was unable to reposition the pump due to tortuous anatomy after the washout. the impella 5.5 was explanted and an impella cp was implanted. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. no product or data logs were returned for evaluation. clinical details noted that pump was unable to be repositioned due to patient's tortuous anatomy. the root cause of the positioning issues was most likely patient condition related from patient's tortuous anatomy.
13245 | 1220648-2024-24540 | 20842100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20842100&pc=OZD | 10/18/2024 | 01/08/2025 |  | 12/04/2024 |  |  | 11/15/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: labs consistent with hemolysis. the patient expired on impella cp support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. b2 date of patient death is unknown. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13246 | 1220648-2024-24541 | 20842587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20842587&pc=OZD | 11/18/2024 | 01/20/2025 |  | 12/04/2024 |  |  | 11/18/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555331 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was major access site bleeding, noted by the hematoma. this was treated by giving the patient 2 units of blood. the patient eventually expired due to multiple organ shock. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13247 | 1220648-2024-24542 | 20842980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20842980&pc=OZD | 11/18/2024 | 01/03/2025 |  | 12/04/2024 |  |  | 11/18/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a thrombus was found on trans-esophageal echocardiogram while patient was on impella cp support. the thrombus was seen in the inflow cannula. the cardiologist decided they did not want to exchange or upgrade the pump. the pump was successfully weaned and removed a few days later. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13248 | 1220648-2024-24544 | 20843055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20843055&pc=OZD | 11/20/2024 | 01/09/2025 |  | 12/04/2024 |  |  | 11/20/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557140 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 insertion the 5.5 kinked and was unable to advance through the graft anastomosis. the 5.5 was exchanged for an impella cp. there was no harm to the patient. | [5 empty] | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues product damage could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13249 | 1220648-2024-24545 | 20843424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20843424&pc=PYX | 11/20/2024 | 12/27/2024 |  | 12/04/2024 |  |  | 11/20/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025563956 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was found to have acute bleed in middle right lobe. it was thought that the reason for the bleed was most likely related to wire placement during insertion of the rp flex and that the lung tissue was found to be very friable. the device was removed shortly afterwards. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿.
13250 | 1220648-2024-24543 | 20843623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20843623&pc=OZD | 05/08/2024 | 01/07/2025 |  | 12/04/2024 |  |  | 11/18/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening cardiogenic shock with a "unknown (undetermined)" relationship to the impella device used: the patient experienced severe in-stent restenosis of the left anterior descending artery stents and underwent complex percutaneous coronary intervention with an impella placement. the patient developed worsening cardiogenic shock. after a palliative care meeting with the family, it was decided to remove the impella device. the impella device was removed and the patient expired same day. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported cardiogenic shock has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the cardiogenic shock could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13251 | 1220648-2024-24547 | 20844165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20844165&pc=PYX | 11/19/2024 | 01/17/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025585015 |  | 0046-0011 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | Defective Component | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient on impella rp and while on support the 0.027 inch guidewire became unsterile during implant and that there was a pump stop. the guidewire issue was treated by using a backup guidewire from another rp set instead. the pump stop was resolved by explanting the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella; rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | D.9 device available for evaluation was left blank and should have been yes on the follow up manufacturer device report number 1220648-2024-24547. | The investigation has been completed. the returned guidewire was inspected and there were no abnormalities. the returned pump was severed along the catheter so only a partial pump was returned for analysis. the returned pump had biomaterial in both the inflow and ouflow cages. there was biomaterial by the impeller. the data logs confirmed 'impella stopped - motor current high alarms" and showed a motor current spike prior to pump stop. the root cause of the pump stop was determined to be biomaterial. the root cause of the delivery issue was most likely use-related. d.9 date device returned/information was added.
13252 | 1220648-2024-24546 | 20844218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20844218&pc=PYX | 11/19/2024 | 12/04/2024 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025461511 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient had an impella rp flex implant and following the implant, there was a potential inlet obstruction. there was a known clot near the inlet. there was a decrease in pump flows and purge flows with an increase in motor current. tissue plasminogen activator was added to the purge and the staff monitored. the patient expired after 12.67 hours of impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
13253 | 1220648-2024-24549 | 20844468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20844468&pc=OZD | 11/18/2024 | 01/03/2025 |  | 12/04/2024 |  |  | 11/18/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567462 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, there were false readings on waveforms and flows. the patient was bridged to a left ventricular assist device and survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation has been completed. the positioning issue was changed to an optical signal issue as it was stated there were false readings and the position was verified. no product was returned. according to the data logs, about 32 hours in there was a spike in motor current and placement signal followed by saturated flows. during this time, the purge flow began to decrease. all other 5s parameters were stable. based on the data log analysis, the root cause of the optical signal issue is most likely due to ingested biomaterial. d.4 serial number was previously entered incorrectly on the initial manufacturer device report number and was revised. h.6 medical device problem code was revised due to investigation.
13254 | 1220648-2024-24548 | 20844743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20844743&pc=OZD | 11/12/2024 | 01/09/2025 |  | 12/04/2024 |  |  | 11/12/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565512 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a positioning issue involving an implanted impella cp pump during patient support. it was noted via echocardiogram that the pump was deep and directed towards the mitral valve. the pigtail appeared to be behind the papillary muscle or cardiac structure. as a result, the decision was made to escalate to an impella 5.5. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue has been completed. the clinical details confirmed that the pump was out of position. the root cause of the positioning issue was unable to be determined as no product was returned and limited clinical details were provided.
13255 | 1220648-2024-24550 | 20845298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20845298&pc=OZD | 11/19/2024 | 01/28/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557313 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella 5.5 support developed acute mesenteric ischemia requiring emergent abdominal surgery twice. the physician stated that this event was attributed to micro-emboli shot off from the impella. before the ischemia event, the impella 5.5 level was dropped to p3 and anticoagulation was on hold for two hours prior. the patient remained stable and survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the pump was retuned for investigation and no abnormalities were observed upon investigation. data logs review showed the irregular spike in purge pressure and spike in the motor current indicating the biomaterial. to determine the true root cause of the thrombus, the clinical information need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. d.9 date device returned/information was added. e.1 revised facility name and address line 2 as the facility name was previously entered in the address line 2 field on initial manufacturer device report number 1220648-2024-24550.
13256 | 1220648-2024-24551 | 20845668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20845668&pc=OZD | 11/19/2024 | 01/09/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon removing the impella cp a large clot was found in the femoral artery. penumbra was used to remove the clot. the patient is stable. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
13257 | 1220648-2024-24553 | 20846321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20846321&pc=OZD | 11/19/2024 | 01/03/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and while on support, the patient had access site bleeding. pressure dressing was applied and one unit of blood products was given to the patient. support continued. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added.
13258 | 1220648-2024-24552 | 20846337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20846337&pc=OZD | 11/19/2024 | 01/07/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567217 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support was turn to one side and this caused placement issues which caused hemolysis. troubleshooting was done to correct the issues, but hemolysis continued therefore the doctor decided to explant the pump. the patient remained stable and survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. pump was returned for investigation. investigation results found out that there was a very big chunk of biomaterial found around the impeller blades and inside the cannula. clinical data mentioned that the pump was out of position with "impella in aorta" alarms due to positioning issue which occurred when patient was made to lay on his side. data logs were confirmed position issue and the "impella in aorta" alarms. as per the data logs the motor current spike was after the position alarms. therefore, as per the data log review and the clinical information provided, the root cause of the hemolysis issue was improper positioning of the device resulting in biomaterial ingestion.
13259 | 1220648-2024-24555 | 20846624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20846624&pc=OZD | 11/19/2024 | 01/22/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568030 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that delivery issues were encountered. a first attempt was done to insert the impella 5.5 over a .035 wire. the device was unable to pass through the subclavian innominate junction and fluoroscopy was utilized to further facilitate delivery. the device was manipulated with several attempts to advance and a wire exchange was performed with .018 and 5f angle tapered glide. the device was removed and catheter shaft showed concern for damage. the decision was made to attempt delivery of a new impella 5.5 device despite difficult vasculature. the patient remained stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation delivery issues has been completed. the root cause of the delivery issue was most likely patient condition since there was difficulty with delivery due to substantial calcium at the subclavian innominate junction.
13260 | 1220648-2024-24554 | 20847066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20847066&pc=OZD | 11/19/2024 | 01/15/2025 |  | 12/04/2024 |  |  | 11/19/2024 | 12/04/2024 |  |  |  |  | 12/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426098 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) priming does not proceed. the purge set was replaced but the priming screen did not proceed. the aic was replaced and the device operated normally. no patient harm has been reported. | [5 empty] | Logs from the complaint date showed multiple occurrences of the purge disc not detected. the issue of purge disc not detected was reproduced by connecting the purge cassette. it was confirmed that the purge flag was broken. the root cause for the purge disc not detected alarm was broken purge flag. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-24554. a5 race and ethnicity was missing. g6 type of report was missing. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13261 | 1220648-2024-24513 | 20831410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20831410&pc=OZD | 11/16/2024 | 01/07/2025 |  | 12/03/2024 |  |  | 11/16/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563705 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support due to anterior st elevated myocardial infarction. while on support, the pump began to have suction which lead into a purge pressure high alarm on p6 with purge pressure being 1057 and impella flows decreased to 1.8 liters. the staff decreased the p level to p5 and the patients blood pressure dropped to 87/59. at p5 the flows were 2.1/2. the motor current was flat and a purge system blocked alarm was present as well as a suction alarm. at this time it was apparent that the pump was not functioning properly and the doctor made the decision to exchange the pump for a new one. once the new pump was placed the doctor decided not to proceed with any more interventions at this time. the patient was unable to tolerate dialysis. with sedation on hold, patient still not responsive. the patient¿s family had meeting with physicians and decided to withdraw and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the high purge pressure has been completed. pump was returned for investigation. there was biomaterial present around the impeller and in the purge gap and the high purge pressure was reproduced. no blood ingress was observed. data logs were returned and investigated and a motor current was observed followed by a rise in purge pressure was observed over the span on 18 minutes leading to the high purge pressure and purge system blocked alarms. these are the characteristics of biomaterial ingestion which was confirmed in the product. therefore, the root cause of the high purge pressure issue was biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
13262 | 1220648-2024-24514 | 20831786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20831786&pc=OZD | 11/16/2024 | 01/03/2025 |  | 12/03/2024 |  |  | 11/16/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567230 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had a gastrointestinal (gi) bleed. the patient received one unit of red blood cells overnight along with two units of platelets. heparin remained off due to gi bleed. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the gastrointestinal bleed (gi bleed) has been completed. data logs are not relevant to the complication and product was not returned for investigation. clinical details provide limited information regarding what could have caused the onset of the gi bleed. due to lacking clinical information, the root cause of the vascular damage could not be determined. the instructions for use for the impella 5.5 for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13263 | 1220648-2024-24515 | 20831837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20831837&pc=OZD | 11/17/2024 | 01/07/2025 |  | 12/03/2024 |  |  | 11/17/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539750 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the patient developed acute onset stroke symptoms with left sided weakness, facial drooping, and slurred speech. the doctor is unsure if stroke symptoms are related to pump. there were also runs of ventricular tachycardia. the patient remains on support. | [5 empty] | The investigation for the stroke and ventricular tachycardia (vt) has been completed. the pump was returned for investigation and no abnormality were noted upon investigation. however, without sufficient clinical details, the root cause of the stroke and ventricular tachycardia (vt) could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13264 | 1220648-2024-24516 | 20832191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832191&pc=OZD | 11/17/2024 | 01/09/2025 |  | 12/03/2024 |  |  | 11/17/2024 | 12/03/2024 |  |  | 09/30/2025 |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444890 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient lost pulses on the impella side. the pump was explanted and an intra-aortic balloon pump was implanted. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. h.6 code 4641 was reported in error on initial manufacturer device report number 1220648-2024-24516 and the correct code was added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13265 | 1220648-2024-24517 | 20832314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832314&pc=OZD | 11/17/2024 | 01/02/2025 |  | 12/03/2024 |  |  | 11/17/2024 | 12/03/2024 |  |  | 07/31/2025 |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440509 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to anterior st elevated myocardial infarction (stemi). the impella cp was repositioned after rp flex placement. unable to achieve good position away from papillary muscles, so, the physician opted to remove device and then replace to optimize position. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13266 | 1220648-2024-24518 | 20832346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832346&pc=OZD | 11/17/2024 | 01/03/2025 |  | 12/03/2024 |  |  | 11/17/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486038 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient had cold extremity. pulse was difficult to locate. additional arterial sheath was placed to the same leg and attached to reaccess port. the patient is stable and the procedure outcome was successful. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13267 | 1220648-2024-24519 | 20832465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832465&pc=OZD | 11/17/2024 | 01/08/2025 |  | 12/03/2024 |  |  | 11/17/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565596 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. right heart catheterization was repeated and an impella rp was placed. heparin was stopped due to the medical doctor feeling that the patient was in disseminated intravascular coagulation. care was withdrawn the following day and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported death has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of death could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13268 | 1220648-2024-24520 | 20832557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832557&pc=OZD | 11/18/2024 | 01/03/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546782 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella 5.5. wire access was lost during insertion. the physician was not comfortable with wireless insertion. multiple attempts to recross were made while weaning off cardiopulmonary bypass. the patient was weaned off the pump. the decision was made to abort further insertion attempts. there was no reported patient harm and the patient remained stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely use issue since wire access during insertion was lost. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13269 | 1220648-2024-24521 | 20832717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832717&pc=OZD | 11/18/2024 | 09/18/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568402 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that impella 5.5 was unable to get above p4 due to suction. significant flow observed on trans-esophageal echocardiography with color. no patient hemodynamic compromise. the patient survived the event. | Additional information indicated that, during patient support, the pump experienced placement signal. |  |  |  |  | The investigation for the low pump flow has been completed. no product was returned for analysis. the data logs show suction alarms before p-level was lowered to p4. the data logs confirm suction alarms. there were no motor current spikes or positioning issues. placement signal was trending downwards slightly over the duration of support. the root cause of the low pump flow was most likely patient condition related. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the failure mode "placement signal issue" was changed to "optical signal issue" as complaint is related to cp's optical sensor. no product was returned for analysis. the data logs do not show any optical related alarms. the 5s parameters were within specification. the root cause of the optical signal issue was not determined as no product was returned for analysis and insufficient information related to failure was provided a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number.
13270 | 1220648-2024-24522 | 20832909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20832909&pc=OZD | 11/17/2024 | 12/23/2024 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1230739 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an unexpected automated impella controller shutdown (aic). the bedside clinical staff was recommended to do a controller exchange and return affected aic. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation automated imeplla controller (aic) - shutdown or restart issue has been completed. console logs from the reported day of event are consistent with complaint. console was tested, but issue was not reproduced. the root cause could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-24522. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13271 | 1220648-2024-24523 | 20833148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20833148&pc=OZD | 11/16/2024 | 01/08/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, there was a hematoma at the access site. the impella was removed due to the hematoma and the patient was placed on extra corporeal membrane oxygenation support. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h.6 code 1888 was reported in error on the initial manufacturer device report number 1220648-2024-24523. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13272 | 1220648-2024-24524 | 20833596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20833596&pc=PYX | 11/18/2024 | 01/08/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025547386 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella rp flex placed for the support of a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed but flows had drastically dropped to less than 1 liter per minute. it was reported that there were signs of hemolysis. the pfhg levels had risen. flows were slowly increased back to 1.8 liters per minute. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation for the low pump flow and the hemolysis has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started and run for 56 hours and 30 minutes, motor current (mc) spiked and increased followed low flow. this event remained to the rest of the case. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. the root cause of hemolysis was biomaterial ingestion. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13273 | 1220648-2024-24525 | 20833835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20833835&pc=OZD | 11/17/2024 | 01/08/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558215 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted to the tertiary center from the community due to acute myocardial infarction / cardiogenic shock. the pump was placed but there was a bleed observed at the access site. the team treated the site with two units of red blood cells, dressing adjustment, and stitch to the site. the pump remained on for two more days until care withdrawn and patient expired. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related. the dressing was removed, an additional stitch was placed, and the site was dressed correctly with angle matching. bleeding resolved as per the clinical description provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13274 | 1220648-2024-24526 | 20834047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20834047&pc=OZD | 11/18/2024 | 12/31/2024 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the next day after impella 5.5 was inserted the patient was taken to the catheterization lab for a pericardial drain for a pericardial effusion. additionally, there was concern for gastrointestinal bleed and systemic anticoagulant was held. the patient survived the event. | [5 empty] | The investigation of vascular damage - great vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13275 | 1220648-2024-24527 | 20834413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20834413&pc=OZD | 11/09/2024 | 01/16/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025478192 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was had an impella cp implant, and after the implant, the patient had chills and a fever. the patient became hypotensive and ecmo was implanted. the impella access site was swollen therefore they immediately surgically removed the impella, since during the positioning the femoral artery had torn. the patient continued on extracorporeal membrane oxygenation support and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. the pump was not returned for investigation, and a data log review was not required for this case. according to the clinical details, the femoral artery was torn during positioning, and the pump was surgically explanted. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided.
13276 | 1220648-2024-24528 | 20834692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20834692&pc=OZD | 10/03/2024 | 01/07/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  | 12/31/2024 |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317833 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team was notified via protect iv regarding a possible relationship between the patient's hemolysis and the impella device. the patient had elevated lactate dehydrogenase (ldh) and plasma free hemoglobin post procedure. the patient survived the event. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13277 | 1220648-2024-24529 | 20836008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20836008&pc=OZD | 11/18/2024 | 01/09/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550393 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered delivery issues with the impella cp. after the short sheath was successfully inserted, the pump was unable to pass. it was removed and after administering a shockwave with a 7.5 millimeter balloon, a second unsuccessful attempt to insert the pump was made. the short sheath was subsequently switched out with the long peel away sheath and the pump was still unable to pass through despite being placed over multiple stiff wires. the decision was ultimately made to abort the procedure. there was no reported harm and the patient remained stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was determined to be patient condition because of the calcified condition of the vessels and the pump was unable to be delivered despite multiple attempts. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13278 | 1220648-2024-24532 | 20836127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20836127&pc=OZD | 11/18/2024 | 01/03/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569784 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to non-st elevated myocardial infarction. while on support, heparin was suspended due to nosebleed and oozing at insertion site. the patient received three units of packed red blood cells due to continued bleeding at impella site. patient had to have vascular repair on explant. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13279 | 1220648-2024-24530 | 20836169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20836169&pc=OZD | 11/18/2024 | 01/06/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a hematoma formed during the procedure and manual pressure was applied to help control it. additionally, when the repositioning sheath was advanced, bleeding intensified and an angiography suggested a small perforation present. the decision was made to remove the pump to help the patient become hemodynamically stable and a covered stent was placed with prolonged manual pressure applied to close the arteriotomy. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported vascular damage has been completed. as per the clinical information provided, the patient had pad/ pvd. once the reposition sheath was removed, the bleeding intensified and a stent was placed which suggests that due to patients diseased vessels, the elasticity of the artery. therefore, the root cause of the access site bleeding-major issue was patient condition related to patients diseased vessels.
13280 | 1220648-2024-24531 | 20836217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20836217&pc=OZD | 11/19/2024 | 01/09/2025 |  | 12/03/2024 |  |  | 11/19/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella 5.5. the doctor was unable to advance because patient has a dilated aorta. the case was aborted and there was no reported harm to the patient. | [5 empty] | The investigation for the delivery issue has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the doctor was unable to advance the impella because the patient has a dilated aorta as per the clinical description provided. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24531 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
13281 | 1220648-2024-24533 | 20837847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20837847&pc=PYX | 11/20/2024 | 01/07/2025 |  | 12/03/2024 |  |  | 11/20/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025514500 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a rp flex was implanted to support a patient admitted in with a left ventricular assist device placement two days prior with ventricular tachycardia and right ventricular failure. the impella rp flex was placed but may have obtained damage, which caused high purge pressures. the physician speculated that possibly the microwire interfered with the motor. the high purge pressure prompted replacement of the impella rp flex pump. the patient was reported to be without harm from this event. | [5 empty] | The investigation of high purge pressure has been completed. as per the product investigation and data log review, the root cause of the high purge pressure issue was biomaterial. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
13282 | 1220648-2024-24536 | 20837917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20837917&pc=OZD | 11/18/2024 | 01/09/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  | 06/30/2025 |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373665 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp had access site bleeding, noted by the oozing. this was treated by withholding systemic heparin. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, bleeding from all sites including impella as per the clinical description provided.
13283 | 1220648-2024-24538 | 20838007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20838007&pc=OZD | 11/19/2024 | 01/08/2025 |  | 12/03/2024 |  |  | 11/19/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543808 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported delivery issues were encountered for an impella 5.5. the pump was primed and loaded onto .018 wire and was unable to pass. contrast imaging showed no calcification or stenosis of the patient. another pump insertion attempt was made with a cannula coated with propofol for lubrication and a buddy wire with .035 amplatz without success. then, another attempt was made by trying to pass the pump after the aortic cannula was removed without any success. ultimately, the decision was made to abort the impella 5.5 placement. the patient remained stable and no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella was aborted as heavy calcified area at subclavian innominate junction as per the clinical description was provided. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13284 | 1220648-2024-24539 | 20838022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20838022&pc=OZD | 11/20/2024 | 01/09/2025 |  | 12/03/2024 |  |  | 11/20/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558820 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs) and had low flows at start of support. the patient was planned for protected percutaneous coronary intervention to circumflex and left anterior descending arteries. the impella was implanted, started on auto and increased with low flows of 1.2l/min. a kink was noticed on the cannula close to the motor. consequently, the impella cp pump was explanted and replaced. it was reported the patient was not harm from the event. | [5 empty] | The investigation for the low pump flow has been completed. the pump was returned for investigation. the returned pump passed the electrical testing and was able to run according to specifications in a bucket of water. the only physical damage reported on the returned pump is the cannula kink near the motor housing. the data log showed low pump flow with an active suction alarm from the start of the case run without seen motor current spike or positioning elated abnormalities. as per the clinical notes, the doctor saw the kink in the cannula near the motor housing during low pump flow event. the cause of the low pump flow issue was a kinked cannula, based on returned product and data log review. d.4 catalog number was inadvertently left blank on the initial manufacturer device report 1220648-2024-24539 and was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24539 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
13285 | 1220648-2024-24534 | 20838900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20838900&pc=OZD | 11/18/2024 | 01/03/2025 |  | 12/03/2024 |  |  | 11/18/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523211 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there were positioning issues. the impella was pointing towards the mitral valve. multiple unsuccessful attempts were performed to reposition the pump under echocardiography. urine output is a concern as it is low and dark. the impella cp was explanted. the patient survived the explant. | [5 empty] | The investigation of positioning issues has been completed. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on manufacturer device report 1220648-2024-24534. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13286 | 1220648-2024-24535 | 20838908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20838908&pc=OZD | 11/15/2024 | 01/22/2025 |  | 12/03/2024 |  |  | 11/19/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536153 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the motor waveform suddenly became flat. the impella placement position was checked with an ultrasound but no abnormalities were found. the flow display was also yellow indicating that the flow rate was not optimal. the service life had exceeded its expected lifespan by just under a month and in consideration of deterioration of the motor over time, the impella 5.5 was replaced. | [5 empty] | The investigation for the low pump flow has been completed. data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for about 279 hours, motor current (mc) spiked followed by low flow that triggered auto suction response and suction alarms. also a mc spike occurred and mc dropped down and flatlined with no change in p-levels. flows reduced too after mc become flat. the pump was returned for evaluation. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of low flow was biomaterial ingestion. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13287 | 1220648-2024-24537 | 20838950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20838950&pc=OZD | 11/19/2024 | 01/06/2025 |  | 12/03/2024 |  |  | 11/19/2024 | 12/03/2024 |  |  |  |  | 12/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529850 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a nurse was having difficulty obtaining right lower extremity pulses for a patient on impella cp support. formal dopplers studies were performed. additionally, the patient¿s urine color transitioned from clear yellow to pink tinged. the patient was escalated to an impella 5.5 after 11 hours of impella cp support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes: may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely patient condition related, based on data log review. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. e4 should have been left blank on manufacturer device report 1220648-2024-24537.