13288 | 1220648-2024-24474 | 20818225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818225&pc=PYX | 11/16/2024 | 01/09/2025 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025443729 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Low Cardiac Output | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp support and during support, the patient had oozing at most invasive line sites. the patient had a direct impella insertion. one unit of packed red blood cells were provided to the patient. additionally, during turning of the patient, suction alarms occurred. the patient was having volume issues and two units of packed red blood cells were administered. furthermore, a purge system blocked alarm appeared. there were low purge flows. the purge system was de-aired and the purge system block alarm resolved but the flows were still low. the patient outcome was that they expired after the family decided to withdraw care. | [5 empty] | The investigation for the low pump flow, access site bleed, and high purge pressure has been completed. the pump was not returned for investigation. data logs were returned for investigation which showed suction alarms followed by low flows. data logs showed that there was an instant rise in purge pressure and placement signal. low pump flow was also observed without any p-level change. motor current spikes were also observed after that. since both purge, motor current, and pump flow were affected simultaneously, the cause of the high purge pressure and low pump flow was most likely biomaterial ingestion. the low pump flow and high purge pressure occurred simultaneously with an identical root cause. the cause of the access site bleeding was most likely patient condition related since there was bleeding at multiple lines. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13289 | 1220648-2024-24475 | 20818483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818483&pc=OZD | 11/15/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558155 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding at the impella access site. the bleeding was managed by tightening the perclose slightly, manual pressure, and femstop. the impella was weaned and explanted during the night due to the bleeding. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the bleeding was managed by tightening the per close slightly, manual pressure and femstop. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
13290 | 1220648-2024-24476 | 20818487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818487&pc=OZD | 11/15/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had two impella cp implants, the first one, prior to this complaint, had low pump flows so a second impella cp was implanted (this complaint). the second impella cp had low pump flows as well. the decision was made to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of low pump flow has been completed. the product and data were not returned for review. based on clinical description, the root cause of low flow was most likely due to patient condition. e4 should have been left blank on manufacturer device report 1220648-2024-24476. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
13291 | 1220648-2024-24477 | 20818571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818571&pc=OZD | 11/14/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/14/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444566 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, there was chest output which resulted in heparin being on hold. the patient was brought back to the operating room for bleeding. the physician performed mediastinal exploration and evacuation of hematoma. chest tube output slowed down but the patient still required blood products. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the access site bleed and the vascular damage has been completed. pump was not returned for investigation. clinical information provided mentions that the doctor had to operate the patient for evacuation of hematoma but no other supporting information was provided. therefore, the root cause of the vascular damage leading to the access site bleed issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
13292 | 1220648-2024-24478 | 20818591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818591&pc=OZD | 11/16/2024 | 12/30/2024 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568584 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, a crack was noted on the top of the purge filter. the patient had been ambulating. staff did not see any leaking from the crack, however a nurse said that they thought they saw fluid during ambulation. purge pressure and flows are within limits. there was no harm to the patient and patient remains on support. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. the root cause of the purge leak - pump is most likely due to a damaged sidearm but the exact location of the leak is unknown as the product was not returned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-24478. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
13293 | 1220648-2024-24479 | 20818601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818601&pc=OZD | 11/17/2024 | 01/17/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564180 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, there was a pump stop. the pump was explanted and the patient was stable. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. no product was returned for evaluation. insufficient data logs at time of reported pump stop were provided. clinical details noted patient was on ecpella support when the pump stop occurred. the root cause of the pump stop could not be definitively determined as insufficient product and data logs were provided. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number.
13294 | 1220648-2024-24480 | 20818604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818604&pc=OZD | 11/17/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemolysis | Tachycardia | Hematuria | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had a long high risk percutaneous coronary intervention procedure and in the middle of the case the patient developed hematuria. attributed to the impella as the pigtail was observed attached to the pap muscle after the case. the patient went to unit like that as team determined benefits of flow outweighed risk of hemolysis. it was further reported that the impella was all the way in the aorta. original plan discussed was to keep pump in dca on p1 and remove in the cardiac catheterization lab in the morning. the team pulled back impella by a few cm to ensure pump was actually in aorta, to prevent pigtail or teardrop from stenting open valve. the patient went into ventricular tachycardia and died shortly after. | [5 empty] | The investigation of hemolysis, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on manufacturer device report 1220648-2024-24480. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13295 | 20818679 | 20818679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818679&pc=OZD | 09/12/2024 | 10/30/2024 | 10/30/2024 | 12/02/2024 | 10/30/2024 |  |  | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  | 518049 |  | 0048-3092 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report |  | Non-Healthcare Professional | [6 empty] | 89 YR |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During an emergent case the impella cp was inserted in left femoral artery. the catheter kinked due to patient tortuosity. a second impella cp was opened to finish the case. manufacturer response for impella cp catheter kit, (brand not provided) (per site reporter). rep requested return of product and will issue a no charge replacement.
13296 | 1220648-2024-24481 | 20818714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818714&pc=OZD | 11/15/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565480 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the implant, there was persistent suction noted. the pump was repositioned multiple times to no avail. the pump was explanted and a new cp was placed. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
13297 | 1220648-2024-24483 | 20818746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818746&pc=OZD | 11/18/2024 | 01/08/2025 |  | 12/02/2024 |  |  | 11/18/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562886 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp constantly need to be repositioned despite no patient movement. the impella has been repositioned multiple times per hour. the was no harm to the patient. the patient remains on impella support. | [5 empty] | The investigation for the positioning issue has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient condition as the patient has very large dilated ventricle. d.6b date of explant was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24483 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13298 | 1220648-2024-24482 | 20818788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818788&pc=OZD | 11/18/2024 | 01/29/2025 |  | 12/02/2024 |  |  | 11/18/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525552 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported reported that the impella cp has leakage at the purge filter. it was recommended to exchange the impella cp. the medical doctor wanted other options so a purge filter bypass was discussed. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation for the purge leak has been completed. due to lack of product returned, the root cause of the purge leak - pump could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
13299 | 1220648-2024-24485 | 20818807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818807&pc=OZD | 11/15/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565508 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had a hematoma at the access site. manual pressure and pressure dressing were applied. the patient was transferred to a different hospital and the cp was explanted. it is unknown if the cp was explanted due to the hematoma. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13300 | 1220648-2024-24486 | 20818824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818824&pc=OZD | 11/19/2024 | 01/07/2025 |  | 12/02/2024 |  |  | 11/19/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570604 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the implant, the patient had a hematoma at the access site. manual pressure was held and a femstop was applied. support continued. additionally, the impella was measuring shallow after the implant. the doctor elected not to reposition the pump due to the patient having ectopy when they attempted to reposition. support continued. the patient survived the event. | [5 empty] | The investigation for the ectopy and positioning issue has been completed. clinical mentions that the pump was seating deep but no other information was provided. therefore, the root cause of the positioning and ectopy issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13301 | 1220648-2024-24484 | 20818835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818835&pc=OZD | 11/19/2024 | 01/07/2025 |  | 12/02/2024 |  |  | 11/19/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient developed a cold leg. decision was made to remove cp. thrombus was noted. pump was explanted due to limb ischemia. the patient survived the event. | [5 empty] | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24484. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13302 | 1220648-2024-24487 | 20818893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20818893&pc=OZD | 11/15/2024 | 01/07/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559898 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was a mural thrombus. staff discussed removing the impella due to the thrombus, however the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13303 | 1220648-2024-24488 | 20819273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20819273&pc=OZD | 11/15/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support ongoing for a patient admitted in with acute myocardial infarction and cardiogenic shock. the impella cp site was bleeding after two days and the team explanted the pump, escalated to a 5.5, and gave blood products. the hematoma was surgically repaired and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13304 | 1220648-2024-24489 | 20819392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20819392&pc=OZD | 11/15/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had dysphasia and central vision changes. computed tomography scan confirmed two localized areas of subarachnoid hemorrhage. the physician decided to remove the impella. the patient survived the event. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13305 | 1220648-2024-24490 | 20819652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20819652&pc=OZD | 11/15/2024 | 01/21/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491515 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the physician reported pump stop after motor current increase. several attempts were made to restart the first pump (see case / patient identifier 01358574). recommendation given to change the pump (sn: (b)(6). pump stop recurred. a decision was made to explant the pump. the patient¿s outcome is unknown. | [5 empty] | The investigation of pump stop has been completed. the pump was returned for investigation. the returned pump showed a large purge gap, which is a sign of bearing wear. no biomaterial was observed on the impeller. further investigation could not be performed due to the pump being cut. the data log indicates that the pump stopped. the root cause of the pump stop was bearing wear after 10 days of use, which is beyond the 8-day design life. motor current spike and temporary purge pressure rise were noted during case run, which could be sign of biomaterial interaction. the cause of the temporary pump stop issue was most likely biomaterial. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13306 | 1220648-2024-24491 | 20819841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20819841&pc=OZD | 11/15/2024 | 01/15/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090360 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller was not able to boot up. there was no patient involvement. | [5 empty] | The investigation for the console (aic) system self check error has been completed. data logs were returned for evaluation and during boot up, the aic began having pbm communication errors. this failed communication caused the console to not boot up normally as described in the complaint. the aic was returned for evaluation and the failure mode from the field was reproduced. the pbm board was found to have a corroded trace that carries communication for pbm1. the corrosion was caused by a leaking super capacitor on the pbm. the cause of the pbm communication failure was a contaminated communication pcb trace. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13307 | 1220648-2024-24492 | 20820019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20820019&pc=OZD | 11/15/2024 | 01/20/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547552 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue. the procedure was aborted and the pump was removed due to the patient anatomy. the patient's outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the insertion was aborted due to patient anatomy. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13308 | 1220648-2024-24493 | 20820424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20820424&pc=OZD | 07/10/2024 | 06/27/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491740 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Unspecified Vascular Problem | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured rv failure with a "unknown (undetermined)" relationship to the impella 5.5 device used: ¿after the patient was cannulated with impella 5.5 , he was noted to have worsening rv failure with cvp aroun 28 and hypotension with no response to max dose of iv diuretics for which he was started on crrt , there was an attempt to decrease the impella support to help rv function but patient became hypotensive. given severe rv failure and unresponsiveness to interventions supportive medicine was consulted and a family discussion was held. patient was then transitioned to comfort care and subsequently passed away.¿ the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of vascular damage - great vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13309 | 1220648-2024-24494 | 20820652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20820652&pc=OZD | 10/30/2024 | 12/26/2024 |  | 12/02/2024 |  |  | 11/12/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1228175 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | During product evaluation, it was reported to the complaints team that it was determined that the automated impella controller fan is defective. the investigation is on going and there was no patient involvement. | [5 empty] | The following have been reported incorrectly on manufacturer device report 1220648-2024-24494. b3 date of event. h5 labeled for single use. h8 usage of device. | The investigation into automated imeplla controller (aic) - damage before therapy has been completed. the fan's connections were examined and found to be severely pinched. the root cause of the issues that arose during preventative maintenance were due to a defective fan. | The impella device was returned, and an investigation has begun but has not yet been completed. upon completion of our analysis, a supplemental report will be filed.
13310 | 1220648-2024-24495 | 20820918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20820918&pc=OZD | 11/15/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had heparin stopped due to brain bleed. neuro was consulted. heparin was resumed the next day, but after a few days, there was concerned for a new finding of ischemic stroke. heparin was stopped. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13311 | 1220648-2024-24496 | 20821785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20821785&pc=OZD | 11/16/2024 | 01/08/2025 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534817 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support for a acute myocardial infarction and cardiogenic shock patient. the support lasted for 3 days and then the pump was explanted. after the explant the vessel achieved hemostasis by the use of a femostop but, the limb became ischemic and had thrombectomy performed. the patient remained stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
13312 | 1220648-2024-24498 | 20824483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20824483&pc=OZD | 11/15/2024 | 01/07/2025 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568062 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the nurse noted a pupillary change and a stroke code was called. sedation was stopped and cardiology was notified. computed tomography (ct) showed acute basilar and left posterior cerebral artery (pca) thrombus. it was also noted that left lower extremity (lle) had decreased flow, although still present, and was cold to the touch. after returning from ct, a bedside echocardiogram showed the impella in good position. given concerns for evolving limb ischemia, while awaiting coordination with neurology for acute basilar and left pca thrombectomy, the physician was consulted for impella removal and lle thromboembolectomy. the next day the patient was taken to neurology and clot was removed from the basal and pca. the patient was then taken to operating room for removal of left femoral impella, thromboembolectomy of left common/external iliac artery, common femoral, profounda, superficial femoral artery (sfa) and popliteal artery, common femoral endarterectomy and repair of the left sfa. in addition, there was 4 compartment fasciotomy of the left lower extremity. it was noted that the patient had an extensive clot throughout the iliac and femoral arteries with dense calcification plaque throughout. the patient was successfully weaned and the patient survived the explant. | [5 empty] | The investigation for the thrombus, limb ischemia, and stroke has been completed. the impella device was not returned for evaluation. without additional information, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13313 | 1220648-2024-24499 | 20824805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20824805&pc=PYX | 11/16/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025534026 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support and during support, there was a sudden drop in flow and spike in pa pressure. the physician decided to explant the pump. upon removing the rp flex clot was noted in the cannula. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of thrombus and low pump flow has been completed. the data logs confirmed the failure mode and clinical narrative. logs showed after pump started and run for 8 minutes, mc spiked followed low flow that triggered auto suction response & suction alarms. this event remained to the rest of the case. product analysis: product was returned in damaged condition (catheter was cut off with no pump plug). visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. the root cause of low flow was biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined g4 pma/510(k) number was reported incorrectly on manufacturer device report 1220648-2024-24499.
13314 | 1220648-2024-24500 | 20825086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825086&pc=OZD | 11/17/2024 | 01/15/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148920 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during a routine purge cassette change, the automated impella controller (aic) was continuously priming tubing. the aic was replaced and the issue resolved. no patient harm was reported. | [5 empty] | The investigation for the priming issue has been completed. console logs from the reported day of event are consistent with complaint and show three successful purge cassette changes, and one (last) unsuccessful purge cassette change, during which console presented with ¿defective purge cassette¿ alarm (#403) and prompted operator to ¿replace purge cassette. press the purge menu soft key then select change cassette & bag¿. console logs entries at the time a new purge cassette was inserted indicate crc error while reading purge cassette data. console was returned for evaluation and was able to detect and communicate with a known-good purge cassette. by manipulating the content of purge cassette data causing its corruption, we were able to reproduce console log entries identical to the complaint data. inspection of the console revealed contamination of the purge insert, and particularly purge disk retainer and flag area, which resulted in failure of disk detect flag and retainer during testing - both unrelated to this complaint. the root cause of the issue could not be determined. the console functioned properly. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13315 | 1220648-2024-24497 | 20825092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825092&pc=OZD | 11/15/2024 | 01/17/2025 |  | 12/02/2024 |  |  | 11/16/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517019 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The eu complainant had a 5.5 pump placed to support a patient admitted in for high-risk surgery and suffering from post-cardiotomy cardiogenic shock. the pump supported for just two days and then was explanted/weaned quickly due to a purge leak. the leak was not treated by bypass. the patient survived the event. | [5 empty] | H.6 code 1491 was added due to investigation results. h.10 additional instructions for use were added due to new code being added. the investigation of purge leak ¿ pump and high purge pressure has been completed. data logs were reviewed and the high purge pressure event was confirmed on the first day of support. immediately after the pump was started, the purge pressure began to rise for 11 minutes until the high purge pressure alarm was triggered. there were no abnormal motor current spikes during the start up that would suggest biomaterial ingestion. however, there was variability in the motor current during periods of constant p-level while the purge pressure was rising. this suggests an impediment to impeller rotation. the purge pressure resolved 45 minutes later after a purge cassette change. the product was not returned for investigation and clinical details do not report any other troubleshooting methods. the root cause of the high purge pressure was unable to be determined without the product. during the high purge pressure event, there were three triggers of the purge system open alarm. the first correlated with a purge cassette change. however, the second two did not. for both instances, the alarm resolved within 30 seconds, and there are no further purge issues for the remainder of the case. clinical details reported a dripping of the purge solution, but the source could not be identified and a filter bypass was not reported to have been done. since product was not returned for investigation and clinical details do not provide any troubleshooting methods that resolved the issue, the root cause of the purge leak could not be determined. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ a.3 revised gender to reflect unknown as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24497. b.3 revise date of event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-24497. e.1 revised address line 1 and city as they were previously entered incorrectly on initial manufacturer device report number 1220648-2024-24497. g.3 30-day should of been marked as yes on the initial manufacturer device report number 1220648-2024-24497. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿.
13316 | 1220648-2024-24501 | 20825186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825186&pc=OZD | 11/17/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had purge high and purge blocked alarms shortly before the pump stopped. several unsuccessful attempts were made to restart the cp. the impella was explanted. | [5 empty] | Data logs were reviewed and high purge pressure and a pump stop were confirmed on the reported date. the returned product was investigated and there was biomaterial in the inflow cage, wrapped around the proximal end of the cannula, and in the purge gap. based on data logs and returned product the root cause of the pump stop was determined to be biomaterial. since the biomaterial caused the high purge pressure, which then resulted in a pump stop, the investigation into the pump stop will be representative of both reported failure modes. d9 date device returned to manufacturer added. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13317 | 1220648-2024-24502 | 20825200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825200&pc=OZD | 11/15/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/15/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566550 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the impella was unable to pull much volume due to the pump being behind the papillary. due to patient anatomy, the impella was unable to be removed from behind the papillary. the decision was made to drop impella to p-4 and maintain. suction alarms resolved with fluid boluses. patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on clinical details the root cause of the positioning issues was determined to most likely be patient condition. e4 should have been left blank on manufacturer device report 1220648-2024-24502. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13318 | 1220648-2024-24503 | 20825315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825315&pc=OZD | 11/17/2024 | 12/27/2024 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474285 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient's left ventricle was perforated after impella 5.5 placement. the impella 5.5 was removed and perforation was surgically repaired. the impella 5.5 then was reimplanted utilizing a pigtail catheter and wires without any difficulty. the patient was stable and the procedural outcome was successful with this device. | [5 empty] | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13319 | 1220648-2024-24507 | 20825335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825335&pc=OZD | 11/15/2024 | 01/16/2025 |  | 12/02/2024 |  |  | 11/18/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025527889 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella 5.5 support. while on support, hemolysis appeared worse and flows were 'false'. there was a dissociation of systolic signals and the decision was made to remove the impella 5.5 and was replaced with a left ventricular assist device (levitronix). the patient periarrested on the table during induction. the surgeon reported there was clot formation at the impella inlet. the impella 5.5 was removed due to hemolysis and incorrect flow data displayed. the patient survived the explant. | [5 empty] | The investigation for the hemolysis, thrombus, and low pump flow has been completed. the pump was not returned for investigation. according to the provided clinical details, thrombus-like material was observed outside the cannula, with no position-related issues. hemolysis worsened during the case, and the pump flow was also inaccurate. the doctor decided to explant the device. according to the data log, a motor current spike was reported, followed by pump flow saturation on day reported hemolysis. no abnormalities were reported in the placement signal or positioning issues. the low pump flow failure mode was changed to saturated pump flow failure mode after a data log review. additionally, thrombus failure was removed since it was addressed under the causes of hemolysis and saturated pump flow. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review and given clinical details. the cause of the hemolysis issue was most likely biomaterial, based on data log review and given clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure."
13320 | 1220648-2024-24506 | 20825347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825347&pc=OZD | 11/17/2024 | 12/30/2024 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502487 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in and transferred to the tertiary center with a cardiac history. the team did the cath and placed distal perfusion and then placed an impella cp for support. during the brief support, urine and labs were hemolyzed due to deep positioning. the hemolysis persisted despite all standard troubleshooting measures being taken. the pump was explanted and escalated to the impella 5.5 support. the patient was stable and the procedure outcome was successful. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. e4 should have been left blank on manufacturer device report 1220648-2024-24506. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13321 | 1220648-2024-24511 | 20825372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825372&pc=OZD | 11/20/2024 | 01/03/2025 |  | 12/02/2024 |  |  | 11/20/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569696 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, significant bleeding and hematoma to right groin was noted. dressing was removed and manual pressure for approximately 55 minutes and no further bleeding was noted and hematoma was softer and not increasing in size. dressing applied with angle matching. the doctor decided to explant at bedside due to risk of rebleeding and because patient had been weaned down to p2 and remained stable. manual compression held to achieve hemostasis. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient movement since the bleeding started occurring after nurse turned the patient to get her off the bedpan and immediately afterwards noted significant bleeding and hematoma to right groin. e4 should have been left blank on manufacturer device report 1220648-2024-24511. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13322 | 1220648-2024-24509 | 20825375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825375&pc=OZD | 11/19/2024 | 01/13/2025 |  | 12/02/2024 |  |  | 11/19/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524023 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue. the pump was placed three separate times inside the left ventricle. when the wire was removed, the pump would pull back above the aortic valve and not re-advance. after the third rewire and placement attempt, the pump was exchanged for another cp device. the patient survived the event. | [5 empty] | The investigation of delivery issues has been completed. the pump was returned and upon investigation it was found that the catheter was damaged and had a hole. the root cause of delivery issue was determined to be use issue because after the third rewire and placement attempt pump was exchanged for another impella cp device and the insertion was successful. d.9 date device returned/information was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-24509 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13323 | 1220648-2024-24508 | 20825384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825384&pc=OZD | 11/18/2024 | 01/07/2025 |  | 12/02/2024 |  |  | 11/18/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. after the peel away sheath was removed, a large hematoma was noted at the impella access site. manual pressure was held to the site followed by femstop being placed. vascular repair involved the right femoral artery and right iliac thrombectomy. additionally, the patient developed hemolysis. one unit of red blood cells was given to the patient. the impella cp was removed due to the hematoma. the patient survived the explant. | [5 empty] | The investigation for the access site bleed has been completed. the pump was not returned for investigation. a data log review was not required for this case. according to the provided clinical details, a hematoma at the access site required a surgical approach for removal. vascular repair was performed. clinical details indicating that the access site complication required a surgical approach. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13324 | 1220648-2024-24512 | 20825385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20825385&pc=OZD | 11/20/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/20/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) cart is unstable. the screw used to secure the stand can no longer be tightened; the screw turns in the thread. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) cart issue has been completed. without the device nor additional clinical details, the root cause of the reported issue could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13325 | 1220648-2024-24505 | 20826567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20826567&pc=OZD | 11/17/2024 | 01/27/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569063 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had a large hematoma at the access site. when the pump was being repositioned, force was used which caused the hematoma. blood products were given to the patient and the pump was explanted. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was returned for evaluation and no problems were found with the device. there was no sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13326 | 1220648-2024-24510 | 20827702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20827702&pc=OZD | 11/19/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/19/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541794 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was access site bleeding, noted by the oozing from the impella site as well as from the venous sheath. this was treated by giving the patient one unit of blood. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, oozing from impella site as well as from venous sheath as per the clinical description provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13327 | 1220648-2024-24504 | 20827803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20827803&pc=OZD | 11/17/2024 | 01/06/2025 |  | 12/02/2024 |  |  | 11/17/2024 | 12/02/2024 |  |  |  |  | 12/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025563727 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant placed an impella cp to support a patient admitted in for acute myocardial infarction / cardiogenic shock. the pump was supporting but due to an access site hematoma, the site was treated by adjustment of the sutures and manual pressure. this did not fully remedy the issue and so the sheath was exchanged to a 16 french cook sheath. the hemostasis was achieved and a new impella cp was placed in the opposite limb. the patient survived the explant. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24504. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13328 | 1220648-2024-24445 | 20813584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813584&pc=OZD | 11/09/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 11/09/2021 | 11/30/2024 |  |  | 10/31/2021 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1454815 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the patient was observed to be bleeding into the abdomen. the impella involvement in the reported injury, how the complication was managed, and if hemostasis was achieved are unknown. the procedural outcome was the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The udi , manufacturing date and expiry date of the device is unknown. should the device information be received, a supplemental mdr will be filed. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13329 | 1220648-2024-24446 | 20813585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813585&pc=OZD | 08/03/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 08/03/2021 | 11/30/2024 |  |  |  |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1441125 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, bleeding from the chest with central cannulation of venoarterial extracorporeal membrane oxygenation (va ecmo) was observed. it was reported that twelve units of red blood cells were transfused to the patient to resolve the issue. the patient had a direct impella access site. additionally, it was reported that during washout, a thrombus was identified and removed. it was stated that at removal of the thrombus, heart function slightly improved however, the patient decompensated soon after. the va ecmo was weaned and explanted, however the patient was in severe right heart failure and expired in the or, while on impella support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The udi, manufacturing date and expiry date of the device is unknown. should the device information be received, a supplemental mdr will be filed. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13330 | 1220648-2024-24447 | 20813671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813671&pc=OZD | 10/23/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 10/23/2021 | 11/30/2024 |  |  | 05/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022088789 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was to be implanted with an impella 5.0 for mechanical circulatory support. it was reported that the patient presented with covid-like symptoms on (b)(6) 2021. upon arrival, the patient coded three times in twenty-four hours. an electrocardiogram performed noted an anterior non-st elevation myocardial infarction. the patient was brought to the cardiac catheterization laboratory but was surgically turned down. the decision was then made to transfer the patient to another hospital but coded as soon as the transport team arrived and was denied. a return of spontaneous circulation was achieved. on (b)(6) 2021, ten days after admission, the 5.0 impella was decided to be placed to stabilize the patient for transfer. the impella 5.0 insertion was aborted due to the patient's vascular anatomy. the physician then decided to place an impella cp into the axillary artery, and the delivery was successful. the patient remained unstable, did not improve and expired three days later. the cp device was noted to have worked as expected. given the inability to deliver the 5.0 device, the patient underwent additional intervention to place the impella cp and this could have contributed to the patient's outcome although the patient expired three days later. therefore, conservatively reporting as vessels characteristics more likely have impacted the event. immediate needed support was unable to be delivered due to patient's vascular anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13331 | 1220648-2024-24448 | 20813673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813673&pc=OZD | 03/10/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 03/10/2022 | 11/30/2024 |  |  |  |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136276 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Connection Problem | [6 empty] | Renal Failure | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a thrombus and renal failure were noted. in addition, there were automated impella controller (aic) connectivity issues. the patient's ventricles were unable to be seen due to a suspected intra cardiac clot. the patient was immediately taken for a ct for further testing. the patient was switched from lasix to bumex. continuous renal replacement therapy (crrt) was started. ultimately, care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13332 | 1220648-2024-24449 | 20813674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813674&pc=OZD | 03/14/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 03/14/2022 | 11/30/2024 |  |  |  |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489795 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. while on support the patient developed hemolysis. it was reported that the patient received 3 units of packed red blood cells with heparin induced thrombocytopenia panel being sent to the lab. heparin was stopped. continuous renal replacement therapy was also in place. it is unclear if the heparin induced thrombocytopenia panel came back positive or negative. in addition, it was reported that the purge side arm was leaking at the leur lock connection and there was low purge pressure. additional information was provided that noted the patient expired after 139.48 hours of support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D4: expiration date and unique identifier (udi) # are unavailable. h4: device manufacture date are unavailable. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13333 | 1220648-2024-24450 | 20813677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813677&pc=OZD | 08/14/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 08/14/2021 | 11/30/2024 |  |  | 07/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1498346 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | Us complainant reported the patient with unknown ethnicity was on impella ld for hemodynamic support. while on support low purge flow, high purge pressure alarms, and bleeding was noted. tpa was administered in the impella purge due to purge alarms. heparin was removed from purge due to bleeding. it was reported that the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13334 | 1220648-2024-24451 | 20813679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813679&pc=OZD | 07/13/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 07/13/2022 | 11/30/2024 |  |  | 03/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022070993 |  | 005082-RO |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. it was reported while on impella ld support, the patient was taken to the or for exploratory surgery. the patient had significant oozing from root and at valve and had cardiac tamponade and the left ventricle was being compressed. a large clot was removed, and 1500ml of volume was removed from the chest. the patient was transfused with 2 cell saver, 1 unit of packed red blood cells, and 1 unit of platelets. the patient had been heparinized for impella use, heparin was discontinued. the patient remained on impella ld support following the intervention until the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13335 | 1220648-2024-24452 | 20813686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813686&pc=OZD | 06/05/2023 | 11/30/2024 |  | 11/30/2024 |  |  | 06/05/2023 | 11/30/2024 |  |  | 06/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121055 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Prime | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an unsuccessful implant attempt of an impella 2.5 device for a patient with cardiogenic shock from an acute myocardial infarction who would eventually expire. the patient underwent coronary artery bypass graft surgery and experienced an aorta perforation. the perforation was successfully treated. an intra-aortic balloon pump (iabp) was placed after heavy blood loss, and the chest was closed after the bleeding resolved. the patient was transported to the unit and began to spiral into cardiogenic shock. the patient was emergently taken to the cardiac catheterization lab for iabp to impella 2.5 device escalation. the impella 2.5 device was opened, and there was an inability to prime the device despite troubleshooting and phone support. during this time the patient began to rapidly deteriorate and the intent to implant the impella 2.5 device was aborted. an impella cp device was opened, primed, and inserted without an issue. the patient continued deterioration after the impella cp insertion with cardiopulmonary resuscitation for nearly one hour and ultimately required bipella support with an impella rp prior to stabilizing hemodynamics. the impella rp was inserted and started with outlet at the level of the swan in right pulmonary artery. however, return of spontaneous circulation (rosc) was not obtained. transesophageal echocardiogram revealed outflow was not fully across pulmonic valve. the impella rp was repositioned and rosc immediately obtained. the patient was transported to the intensive care unit in critical condition but would later expire the following day. although more likely the patient expired as a result of the cardiac event, the issue with the impella 2.5 device cannot be excluded as a contributing factor in the overall outcome for this event. the impella 2.5 was unable to be fully used and was exchanged for another impella device. at said moment, this resulted in loss of mechanical circulatory support to the patient at a critical time. | [5 empty] | The device and data logs were not returned for analysis. the root cause of the priming issue was not determined as a result of the product not being returned, data logs not received for analysis and insufficient clinical information provided. this report is being filed as part of a retrospective review of historical records.
13336 | 1220648-2024-24453 | 20813688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813688&pc=OZD | 01/27/2022 | 11/03/2025 |  | 11/30/2024 |  |  | 01/27/2022 | 11/30/2024 |  |  | 11/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the pump was unable to go above p2 without suction and flows were lower than expected at 1.8 l/min. an attempt was made to escalate the patient to an impella 5.0, however the 5.0 would not cross, and the insertion was aborted. the impella cp remained implanted, and was repositioned after the 5.0 attempt to resolve the flow issues. the patient remained on impella cp support until there was no longer a need for mechanical support, and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13337 | 1220648-2024-24454 | 20813689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813689&pc=OZD | 01/25/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 01/25/2022 | 11/30/2024 |  |  | 08/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1502422 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. it was reported while on impella ld support, the patient began to decompensate despite impella support. multiple pressors were administered and the patient was started on continuous renal replacement therapy. it was noted that there were suction alarms noted at p9, so the p level was decreased to p7 and because of a drop of hemoglobin value from 8.3 to 6.3, the patient was transfused with 2 units of packed red blood cells and platelets. bleeding was observed later that day and the medical team decided to reopen the patient's chest after the patient was hemodynamically unstable and had an increased need for pressors. the patient was brought to the catheterization laboratory to be escalated to an impella 5.5, however delivery was unsuccessful. an impella cp was used instead with success, and the impella ld was successfully removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13338 | 1220648-2024-24455 | 20813695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813695&pc=OZD | 03/22/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 03/22/2022 | 11/30/2024 |  |  | 04/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119927 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Positioning Failure | Restricted Flow rate | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. it was reported while on impella ld support, the pump had positioning issues which caused the pump flow to be <1.0 l/min. the physician reported issues with obtaining optimal placement of the device, and could not correct the low pump flow with repositioning. the physician decided to explant the impella ld, and replace the device with an impella 5.5. the patient underwent surgical intervention to place the impella 5.5. it was noted that the patient's care was withdrawn 11 days later and the patient expired. the positioning issues that required surgical intervention to resolve could not be excluded as a contributing factor with the limited information at hand, even though the patient expired 11 days later. | [5 empty] | The investigation for the reported low flow and positioning issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical account, the root cause of the low flow was improper positioning due to patient anatomy. this report is being filed as part of a retrospective review of historical records.
13339 | 1220648-2024-24456 | 20813696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813696&pc=OZD | 08/16/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 08/16/2021 | 11/30/2024 |  |  | 05/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022094479 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported while on impella 5.0 support, the purge filter began leaking due to patient mobilization. the issue was resolved by performing a purge reservoir and filter bypass. it was noted that no cleaning agents were used. the purge filter began leaking again, and it was decided to wean the patient and explant the impella device. there was no patient harm reported, and the patient was hemodynamically stable after explant. | [5 empty] | The investigation into the reported purge leak - pump has been completed. the impella device was not returned for evaluation. the cause of the purge leak was sidearm filter damage as a result of patient mobilization. this report is being filed as part of a retrospective review of historical records.
13340 | 1220648-2024-24457 | 20813697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813697&pc=OZD | 01/12/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 01/12/2022 | 11/30/2024 |  |  | 09/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022135069 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient experienced a cerebrovascular accident (cva) due to a cerebral infarct lesion, which was reportedly a consequence of prior medical management. however, as the patient was heparinized for impella use it cannot be said that there is not a relationship between impella use and the cva. it was noted that the patient's care was withdrawn and the patient expired. the impella cannot be excluded as a possible contributing factor, however the patient's condition may be more likely have impacted the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13341 | 1220648-2024-24458 | 20813698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813698&pc=OZD | 01/21/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 01/21/2022 | 11/30/2024 |  |  | 06/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022136952 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was escalated to an impella ld from an impella cp for mechanical circulatory support. it was reported after the impella ld was weaned and explanted, the medical team found a clot in the graft connected to the aorta. there were no medical intervention required. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. biomaterial was observed on the pump at the time of pump removal. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13342 | 1220648-2024-24459 | 20813702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813702&pc=OZD | 12/24/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 12/24/2021 | 11/30/2024 |  |  | 05/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022092423 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support it was reported that the purge flow rate was gradually decreasing, and long-term support would be needed. the impella 5.0 was explanted. in addition, ventricular tachycardia was occurring frequently, and direct current defibrillation was being applied each time. the surgeon informed the family that it was difficult to save the patient's life. additional information was provided that noted that care was withdrawn and a do not resuscitate order was put in place. the patient passed away after 216.42 of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13343 | 1220648-2024-24460 | 20813703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813703&pc=OZD | 03/25/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 03/25/2022 | 11/30/2024 |  |  | 02/29/2024 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023184269 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device along with extracorporeal membrane oxygenation for mechanical circulatory support. while on support the patient was restricted on anticoagulation due to bleeding from abdominal surgery/hemicolectomy (ischemia). additional information was provided that noted care was withdrawn due to the patient's critical condition and the patient expired. there were no reported issues with the impella; however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13344 | 1220648-2024-24461 | 20813704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813704&pc=OZD | 02/08/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 02/08/2022 | 11/30/2024 |  |  |  |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1466157 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported the patient experienced sepsis and went into septic shock during impella support. the sepsis was treated with endotoxin adsorption therapy, and the administration of antibacterial drugs. the patient required reintubation after entering septic shock. the following day, it was reported the patient had bleeding from an unknown origin. the patient was transfused with blood products, the number of units the patient received is unknown. the patient was in prolonged shock and his general condition was poor, and it was noted that the cause of the bleeding was unknown. a thoracentesis performed, and 2l of blood was removed. bleeding was noted to have resolved the following day; however, the patient's condition was still severe. on (b)(6) 2022, the patient experienced a cerebral hemorrhage and was given a large dose of catecholamines. the patient was then made a do not resuscitate order, care was withdrawn and the patient and expired. there is not enough evidence to exclude impella as an associated factor in the patient's demise. | [5 empty] | D4: udi, expiration date is unknown. h4: manufacturing date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13345 | 1220648-2024-24462 | 20813710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813710&pc=OZD | 11/02/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 11/02/2021 | 11/30/2024 |  |  | 04/30/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1475709 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device to replace an existing impella cp for increased mechanical circulatory support. the complainant reported that the pump had purge pressure high alarms. troubleshooting the alarm with rapid deceleration technique was utilized with minimal success. despite aggressive inotropes and vasopressors, the patient continued to deteriorate, and the family opted to withdraw care. the patient expired quickly after support was removed. the high purge pressure is a product malfunction which cannot exclude the impella use from the outcome of demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13346 | 1220648-2024-24463 | 20813715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813715&pc=OZD | 01/27/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 01/27/2022 | 11/30/2024 |  |  | 03/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022067970 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the impella 5.0 would not cross into aorta, therefore the implant was aborted. the physician had the 5.0 passed the graft but could not cross the innominate artery. the device was removed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13347 | 1220648-2024-24464 | 20813821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813821&pc=OZD | 06/22/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 06/22/2022 | 11/30/2024 |  |  | 10/31/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. the impella 5.0 replaced an impella cp, as the impella cp had been implanted for longer than recommended. it was reported after the impella 5.0 was inserted; liquid was found to be leaking from the purge housing. the purge cassette was replaced immediately and resolved the issue. additionally, it was noted that the patient had an infection that could not be controlled, and bleed from all over his body. the patient's anemia progressed, and he developed bradycardia. the patient was made a do not resuscitate, care was withdrawn, and the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Note: the device was not returned for evaluation, however the purge cassette was returned on june 24, 2022. the complaint purge cassette was visually inspected and cut open, a leak test was performed. the problem was reproduced, a purge leak was observed from the piston seal. the cause of the purge leak was a piston failure. this report is being filed as part of a retrospective review of historical records.
13348 | 1220648-2024-24465 | 20813836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813836&pc=OZD | 01/11/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 01/18/2022 | 11/30/2024 |  |  | 07/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492391 |  | 005066 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.0 for mechanical circulatory support and experienced low purge flow and cerebral infarction with a demise outcome. the patient underwent percutaneous coronary intervention with support impella cp and extracorporeal membrane oxygenation (ecmo) support. the patient was then escalated to impella 5.0 device remaining on ecmo. the impella 5.0 was inserted from the right subclavicle with complications. with subclavian insertion, the vasometer was sufficient. however, the patient had swelling overall, it was difficult to insert but eventually completed smoothly. the pump was turned on. immediately following, the impella 5.0 had a low flow (4.0l/min), and therefore, weaning was carried out early. act had no problem at the time of insertion and management. sometime following the low purge flow within the impella 5.0 support period, the patient had a cerebral infarction, consequently treatment was discontinued, and the patient and expired two days later per the report. the treating physician commented the cause of death is due to the do not resuscitate policy for the cerebral infarction, and the relationship between the impella 5.0 and the cerebral infarction is possibly related. this report is being filed as part of a retrospective review of historical records. | [5 empty] | The pump was received for analysis. visual inspection revealed no abnormalities were observed. the pump was run in a basin of water and the pump flow was within specification at 5.0l/min. in conclusion, the cause of the low pump flow was most likely patient condition since the returned product operated to specification. this report is being filed as part of a retrospective review of historical records.
13349 | 1220648-2024-24466 | 20813839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813839&pc=OZD | 08/24/2022 | 05/12/2025 |  | 11/30/2024 |  |  | 08/24/2022 | 11/30/2024 |  |  | 05/31/2024 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023224248 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced intermittent bleeding from the pump insertion site as well as the arterial cannula site. it was reported that the pump was explanted, and the procedural outcome was the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | B2: date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13350 | 1220648-2024-24467 | 20813851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813851&pc=OZD | 07/07/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 07/07/2021 | 11/30/2024 |  |  | 07/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1499590 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. it was reported that the impella 5.0 was unable to be fully utilized, the patient's axillary was too small for the delivery of the 5.0 device. the physician decided to remove the device and replace it with an impella cp. the exchange resulted in loss of mechanical circulatory support to the patient at a critical time. the patient expired the next day; care was withdrawn. although more likely the patient expired as a result of the cardiac event, the issue with the impella 5.0 device cannot be excluded as a contributing factor in the overall outcome for this event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13351 | 1220648-2024-24468 | 20813854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813854&pc=OZD | 07/08/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 07/08/2021 | 11/30/2024 |  |  | 08/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503737 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that the patient experienced sepsis, it was noted that the plan was to let the patient rest and treat the sepsis. the type of treatment is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. the relationship between the impella 5.0 and the patient's diagnosis of sepsis is unknown, however there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13352 | 1220648-2024-24469 | 20813927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813927&pc=OZD | 03/28/2022 | 10/31/2025 |  | 11/30/2024 |  |  | 03/28/2022 | 11/30/2024 |  |  | 03/31/2024 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023184271 |  | 05060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient was noted to be septic, and thrombocytopenia was diagnosed due to low platelets. it was noted the next day that the heparin induced thrombocytopenia was not confirmed through testing. it was reported that 3 days later the patient was in multiple organ failure and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of infection/sepsis and thrombocytopenia has been completed. the log was reviewed and no high purge pressure alarms were found. the returned pump was visually inspected and no biomaterial or catheter kinks were observed. the returned pump cassette was connected to the pump and high purge pressure was not reproduced. the purge operated to specification. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. the cause of the thrombocytopenia was not determined due to insufficient clinical details. d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24469.
13353 | 1220648-2024-24470 | 20814078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20814078&pc=OZD | 07/15/2021 | 11/30/2024 |  | 11/30/2024 |  |  | 07/21/2021 | 11/30/2024 |  |  | 12/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022055488 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 2.5 device for mechanical circulatory support. the patient had been hospitalized for heart failure for several days, and his heart failure worsened, and lactate increased, and it was decided to implant the impella 2.5. it was reported that that patient was anuric and required hemodialysis prior to impella insertion. the patient was placed on continuous renal replacement therapy after the impella 2.5 was implanted. it was noted that the patient's dialysate drainage indicated the patient was experiencing hemolysis. the impella 2.5 was repositioned, and volume was given to the patient. it was noted that the hemolysis was non-serious, however the hemolysis did not resolve while the patient was on impella support. the patient's condition gradually worsened and extracorporeal membrane oxygenation was added for support. circulation could not be maintained even with the use of large amounts of nicotinamide adenine dinucleotide, and metabolic acidosis progressed. the patient experienced cardiopulmonary arrest. it was determined that further treatment was difficult, so care was withdrawn and the patient expired. additional information reported that the patient expired during impella support due to metabolic acidosis and circulatory insufficiency. there is not enough information to exclude the impella device as an associated factor in the patient's demise, however, the patient's presenting condition may be more likely have impacted the event. | [5 empty] | According to clinical, per j-pvad, hemolysis occurred during support and was reported to be related to impella. hemolysis was confirmed by properties of dialysate drainage. no pfhg measurements taken. position was verified by echocardiogram. to troubleshoot, the position of the category was adjusted and volume loading to the left ventricle. the patient did not have right heart failure or left ventricle filling issues. the patient was concurrently supported by extracorporeal membrane oxygenation (ecmo). patient died during impella support. no causal relationship between impella and death. cause of death was metabolic acidosis, circulatory insufficiency. impella was working until the patient died. the device and data logs were not returned for analysis. it was concluded the cause of the hemolysis was most likely improper positioning and no causal relationship between impella and death. this report is being filed as part of a retrospective review of historical records.
13354 | 1220648-2024-24471 | 20814092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20814092&pc=OZD | 06/23/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 06/23/2022 | 11/30/2024 |  |  | 12/31/2022 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022051874 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited a placement signal [ps] reading of 20/10¿s and the a-line did not correlate. support advised the medica staff that it seemed to be a sensor issue, and to treat patient from a-line. in addition, the patient experienced ventricular tachycardia and was defibrillated twice. cardio-pulmonary resuscitation (cpr) was administered along with boluses of epinephrine and levosimendan. fluids were also infused. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13355 | 1220648-2024-24472 | 20814121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20814121&pc=OZD | 02/26/2022 | 10/06/2025 |  | 11/30/2024 |  |  | 03/02/2022 | 11/30/2024 |  |  | 06/30/2023 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022097225 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient experienced renal failure and sepsis, treatment for these issues is unknown. additional information noted that the patient died due to a serious infectious disease, this was attributed to the patient being immunocompromised after surgery and undergoing several thoracotomy operations. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24472. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-24472. g1 reporting contact email revised on manufacturer device report 1220648-2024-24472 in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-24472.
13356 | 1220648-2024-24400 | 20814125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20814125&pc=OZD | 05/20/2022 | 11/30/2024 |  | 11/30/2024 |  |  | 05/20/2022 | 11/30/2024 |  |  | 02/29/2024 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197296 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombocytopenia | [10 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had been escalated from an impella cp to the impella 5.5 as hemodynamic support need had increased. after being implanted with the impella 5.5 it was reported that the patient was bleeding from multiple sites and was transfused with blood products as a result of the bleeding. on (b)(6) 2022, the patient was diagnosed with heparin induced thrombocytopenia. it was reported the patient's platelet count was 19,000 cells per microliter. the patent was transfused with platelets and blood products. additionally, the patient had three episodes of ventricular fibrillation and required defibrillation to resolve each episode. a family meeting was called to discuss withdrawing care. it is unknown if care was withdrawn, however it is noted that the patient expired while on support. there is not enough evidence to exclude impella as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13357 | 1220648-2024-24473 | 20814145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20814145&pc=PYX | 06/05/2023 | 05/28/2025 |  | 11/30/2024 |  |  | 06/05/2023 | 11/30/2024 |  |  | 12/31/2024 |  | 11/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024295658 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported positioning issue with the impella rp device immediately after implant to a patient with cardiogenic shock from an acute myocardial infarction who would eventually expire. the patient underwent coronary artery bypass graft surgery and experienced an aorta perforation. the perforation was successfully treated. an intra-aortic balloon pump (iabp) was placed after heavy blood loss, and the chest was closed after the bleeding resolved. the patient was transported to the unit and began to spiral into cardiogenic shock. the patient was emergently taken to the cardiac catheterization lab for iabp to impella 2.5 device escalation. the impella 2.5 device was opened, and there was an inability to prime the device despite troubleshooting and phone support. during this time the patient began to rapidly deteriorate and the intent to implant the impella 2.5 device was aborted. an impella cp device was opened, primed, and inserted without an issue. the patient continued deterioration after the impella cp insertion with cardiopulmonary resuscitation for nearly one hour and ultimately required bipella support with an impella rp prior to stabilizing hemodynamics. the impella rp was inserted and started with outlet at the level of the swan in right pulmonary artery. however, return of spontaneous circulation (rosc) was not obtained. transesophageal echocardiogram revealed outflow was not fully across pulmonic valve. the impella rp was repositioned and rosc immediately obtained. the patient was transported to the intensive care unit in critical condition but would later expire the following day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. this is one of two medical device reports associated with this event; the associated impella 2.5 device is reported under manufacturing report number: 1220648-2024-24452. note: the date of death should reflect (b)(6) 2023 for both reports. | No product was returned for analysis. clinical details mention imaging revealed pump positioning was shallow - outflow not fully across pulmonic valve. clinical details do not mention how pump positioning became shallow after initial placement. pump was repositioned and issue resolved. the root cause of the positioning issue was not determined as limited clinical information was provided. this report is being filed as part of a retrospective review of historical records.
13358 | 1220648-2024-24336 | 20810224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810224&pc=OZD | 10/17/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 10/17/2022 | 11/29/2024 |  |  | 03/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023193509 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, the patient was bleeding. an echocardiogram verified the impella was in a correct position. the patient received blood products. in addition, there were perfusion issues after the implant, a subclavial intervention was performed due to perfusion issues and was resolved. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13359 | 1220648-2024-24337 | 20810395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810395&pc=OZD | 12/10/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 12/10/2021 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121608 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia arrest. reportedly, cardio-pulmonary resuscitation was performed for forty-five minutes. the patient¿s wife made the decision to discontinue the lifesaving efforts, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13360 | 1220648-2024-24338 | 20810407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810407&pc=OZD | 12/14/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 12/14/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, that due to the amount of bleeding and active clotting time, the cardiac surgeon decided to go heparin-free for about ten hours, and then gradually use heparin as a purge solution while monitoring. the patient remained on support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13361 | 1220648-2024-24340 | 20810425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810425&pc=OZD | 01/23/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 01/23/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1483212 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, it was noted that microbubbles were seen in the left ventricle. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13362 | 1220648-2024-24339 | 20810426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810426&pc=OZD | 11/16/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 11/16/2022 | 11/29/2024 |  |  | 05/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023217911 |  | 005060 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported that the patient required resuscitation due to ventricular tachycardia. the impella flow was increased to p9, and preparations were made for a percutaneous coronary intervention (pci) of the main vessel. following resuscitation, the impella device was observed to be functioning without issues, and its position was verified. care was eventually withdrawn, and the patient expired while still on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13363 | 1220648-2024-24341 | 20810427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810427&pc=OZD | 03/23/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 03/23/2022 | 11/29/2024 |  |  | 08/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131603 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient experienced bleeding from the right groin. the patient was suspected of having neurological issues overnight and underwent a computed tomography (ct) scan, which revealed two strokes. additionally, the patient was observed to be bleeding from the ecmo arterial site (rfa), requiring 1¿2 units of blood per shift. care was eventually withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13364 | 1220648-2024-24342 | 20810430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810430&pc=OZD | 03/24/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 03/24/2022 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022114102 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient remained in critical condition, with ongoing organ failures. bleeding occurred from the nose, venous lines, and via the endotracheal tube, but these bleeding issues were resolved. the patient passed away shortly after the impella device was deactivated, despite limited therapy. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13365 | 1220648-2024-24343 | 20810433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810433&pc=OZD | 04/20/2022 | 10/13/2025 |  | 11/29/2024 |  |  | 04/20/2022 | 11/29/2024 |  |  | 09/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504585 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. an echocardiogram was performed, revealing a thrombus in the left ventricle. it was further reported that heparin was removed from both the purge and systemic lines due to the patient testing positive for heparin-induced thrombocytopenia (hit). during the course of treatment, the patient experienced a high purge pressure alarm overnight, with attempts made to stop and restart the pump the following morning. ultimately, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that the patient (pt) was being placed on impella 5.0 for hemodynamic support. it was reported that the patient had an echo done with a noted a thrombus in the left ventricle. discussed what to look for if the thrombus gets ingested into the pump. a thrombus while on support. it was reported that heparin was removed from the purge and systemic due to the patient being positive for heparin induced thrombocytopenia. heparin induced thrombocytopenia while on support is a repheparin induced thrombocytopenia while on support. it was reported that the patient experienced high purge pressure alarm overnight with pump stop/start attempted this morning. several days later, it was reported that the purge cassette was visibly leaking with high purge flows. the leaking cassette was changed and purge stabilized. purge cassette leaking is a user inconvenience while on support and non-reportable due to the device not being removed until a later date. ultimately care was withdrawn and the patient expired. the investigation results shows that the event could not be determined as insufficient clinical details were provided. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13366 | 1220648-2024-24344 | 20810434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810434&pc=OZD | 07/24/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 07/24/2022 | 11/29/2024 |  |  | 07/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1502237 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support the patient was determined to be experiencing heparin induced thrombocytopenia, which was addressed by adding bicarbonate to the purge and systemically running bivalirudin. additionally, a computer tomography scan, on (b)(6) 2022, revealed bleeding from the orogastric tube. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13367 | 1220648-2024-24345 | 20810438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810438&pc=OZD | 12/17/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 12/17/2022 | 11/29/2024 |  |  | 03/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023188152 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported a that there was a high purge pressure issue resulting in the purge being changed to tissue plasminogen activator lyse therapy for a patient in critical condition. the impella ran after the lyse therapy, however the impella therapy was stopped due to the patient's worsening condition. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13368 | 1220648-2024-24346 | 20810476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810476&pc=OZD | 05/31/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/31/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1483644 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella ld device for mechanical circulatory support as a bridge to therapy. while on support, the patient experienced bleeding issues during the coronary artery bypass graft requiring six units of packed red blood cells, 6 units of fresh frozen plasma, and two units of platelets. it was reported that due to high purge pressure issues that tissue plasminogen activator was started resulting in a cleared purge pathway. the patient remained on impella support and was reported to have survived a successful weaning and explant. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13369 | 1220648-2024-24347 | 20810502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810502&pc=OZD | 06/16/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/16/2022 | 11/29/2024 |  |  | 09/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013268 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. while on support, the patient experienced bleeding at an unknown location resulting in heparin being withheld. reportedly on (b)(6) 2022 the patient experienced multiple episodes of rapid atrial fibrillation with rapid ventricular response requiring cardioversion. it was reported on (b)(6) 2022 that the patients¿ urinary output decreased, and the patient was started on continuous renal replacement therapy to manage the complication. additional information noted the patients breathing became agonal and the patient was do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13370 | 1220648-2024-24348 | 20810527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810527&pc=OZD | 10/02/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 10/02/2022 | 11/29/2024 |  |  | 04/30/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217711 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced absent pulses in their right arm and the skin was cool to the touch. the patient had been in critical condition, and family decided to withdraw care and declined any further treatment or interventions. the patient's care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13371 | 1220648-2024-24349 | 20810537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810537&pc=OZD | 12/05/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 12/05/2022 | 11/29/2024 |  |  | 04/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022073441 |  | 005060 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the venous cannula was exchanged, but rapid clotting was again observed, with the ventricles also clotting with activated clotting time (act) of greater than 200. no further therapeutic options were available. the patient expired in the catheterization lab while still on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13372 | 1220648-2024-24350 | 20810603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810603&pc=OZD | 12/18/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 12/18/2022 | 11/29/2024 |  |  | 05/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022085037 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, the patient experienced an uncontrollable hemorrhage, which ultimately resulted in the patient's death. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13373 | 1220648-2024-24351 | 20810664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810664&pc=OZD | 10/06/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/06/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022070994 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. at the end of support, once the chest was closed, the patient had significant bleeding. the patient had to be reopened two times, and the physician found a bleeding site and employed shunt to the right artery. patient received six platelets, six units' fresh frozen plasma, four cryo, four units red blood cells, 2000 anti-inhibitor coagulant complex (fieba) full dose factor seven administered in procedure due to bleeding. no peripheral heparin or dti running. it was later noted that the patient was still bleeding, so dextrose 5% in water/bicarbonate running in purge. patient taken back to cardiovascular operating room, evacuated 4.5 cm wall of clot from the mediastinum. hemodynamics greatly improved after evacuation. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13374 | 1220648-2024-24352 | 20810731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810731&pc=OZD | 06/17/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/17/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella ld device for mechanical circulatory support post cardiac surgery. there was access complications and the physician had difficulty positioning the device but was placed successfully. no patient harm was reported. | [5 empty] | D4: the device information such as udi, manuf. date, expiry date, sn, and lot # are not available at this time. should any new information be received, a supplemental mdr will be filed. the investigation into the reported issues is complete. the device was not returned for evaluation, but the returned data logs show the purge flow steadily decrease from 21 ml/hr to around 0 ml/hr accompanied by a high purge pressure alarm. the pp remains high for 2 minutes, then abruptly decreases to around 300 mmhg followed by a spike in pf to 30 ml/hr. the pf and pp then return to normal. the pump was returned and no kinks or abnormalities were found, the high purge pressure could not be reproduced. thus, the root cause of the sudden brief high purge pressure was most likely due to a temporary purge line kink. this report is being filed as part of a retrospective review of historical records.
13375 | 1220648-2024-24353 | 20810797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810797&pc=OZD | 09/07/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 09/07/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022070995 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella ld device for mechanical circulatory support. while on support, the patient was observed to be bleeding post coronary artery bypass graft (cabg). it was stated that the patient was transfused 6 units of blood products in an effort to manage the complication. however, the patient's bleeding was persistent. it was stated that the medical staff withheld heparin until the bleeding was under control. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13376 | 1220648-2024-24354 | 20810809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810809&pc=OZD | 03/04/2022 | 05/22/2025 |  | 11/29/2024 |  |  | 03/04/2022 | 11/29/2024 |  |  | 12/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023162646 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. it was reported that during implantation there was resistance felt when removing the guide wire, so insertion was redone. the doctor tried to insert it again, but the situation suddenly changed, and the insertion was canceled. no further information is available. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of this issue and its relation to the impella device was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records. | The root cause of the delivery issues was unable to be determined as the pump and guidewire were not returned for evaluation. this report is being filed as part of a retrospective review of historical records.
13377 | 1220648-2024-24355 | 20810874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810874&pc=OZD | 03/28/2023 | 11/29/2024 |  | 11/29/2024 |  |  | 05/11/2023 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022110407 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced a ventricular septal perforation (vsp), which caused the patient to expire. there was no issues or malfunctions reported, and the impella was operating as expect. the relationship between the impella and the vsp is unknown. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13378 | 1220648-2024-24356 | 20810883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810883&pc=OZD | 05/09/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 05/09/2022 | 11/29/2024 |  |  | 11/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022148748 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that a patient was escalated from femoral impella cp and peripheral ecmo to central rvad and impella ld. (5.0) it was reported that vascular was called for cutdown of left groin due to absent pulse in left foot. additionally, physician found dissected femoral artery. artery was repaired. it was also noted that the patient was evaluated for possible add of d5/heparin 6.25 to control bleeding. the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13379 | 1220648-2024-24357 | 20810915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810915&pc=OZD | 07/26/2021 | 12/13/2024 |  | 11/29/2024 |  |  | 08/28/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1483213 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Failure to Sense |  |  |  |  | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the pump was explanted and replaced due to low flows/high purge pressure. no patient harm was reported. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0 support, the pump purge pressure was high, and the purge flow was decreasing. tissue plasminogen activator lysis protocol was performed. the purge pressure did not decrease; however, it did not continue to increase. on (b)(6) 2021 the lysis protocol was performed again, and the purge cassette was changed. there was no change to the purge pressure or purge flow and the lysis protocol was repeated a third time without any effect on (b)(6) 2021. additionally, on (b)(6) 2021 the patient began experiencing hemolysis. the medical team decided to proactively replace the impella. on (b)(6) 2021 the impella 5.0 was replaced. there was no additional patient harm reported. |  |  |  |  | The investigation for the reported placement signal and low pump flow issues has been completed. the product was returned, and the issue was confirmed. the field logs were reviewed, and placement signal not reliable and high purge pressure alarms were confirmed. evaluation of the returned pump did not reveal any abnormalities. the pump was connected to a console and purged normally with purge flow of 6.0ml/hr and purge pressure of 600mmhg. also, the differential pressure sensor appeared to be functioning normally. no alarms were reproduced. per the results of the clinical review and investigation, the root cause could not be determined through this investigation as the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. corrected information from initial mfr 1220648-2024-24357. b1: corrected to include adverse event. b2: corrected to include required intervention to prevent permanent impairment/damage (devices). b5: corrected date of event. corrected complaint to include hemolysis and purge pressure issues. h1: corrected to serious injury. h6: corrected codes for hemolysis, explantation, unexpected medical intervention, and increase in pressure.
13380 | 1220648-2024-24358 | 20810962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810962&pc=OZD | 07/02/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 07/03/2022 | 11/29/2024 |  |  | 03/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023195857 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported that a leak was observed during the implant from the device's side arm yellow luer. the purge cassette was replaced but the leak was found to be on the pump. the pump was replaced as a result. there was no patient harm reported. | [5 empty] | The root cause of the purge leak was unable to be determined as no product was returned for evaluation. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13381 | 1220648-2024-24359 | 20810980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810980&pc=OZD | 06/13/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/20/2022 | 11/29/2024 |  |  | 03/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023195870 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The hospital reported that one of the two impella 5.0 packages that were shipped via fedex was received damaged on arrival, stating that "the box looks battered and has been opened". a replacement device was requested. there was no patient involvement. | [5 empty] | An image received from the field showed considerable damage to the underside of the impella box exterior with bending and crumpling. in the same area, it appears that the seal label was compromised. therefore, the cause of the package issue was a damaged label. it is likely that the label was damaged during shipment. this report is being filed as part of a retrospective review of historical records.
13382 | 1220648-2024-24360 | 20810985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20810985&pc=OZD | 04/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/09/2021 | 11/29/2024 |  |  | 04/30/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1428395 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported during a bedside insertion of the impella 5.0, the physician was unable to advance the pump through the anastomosis. the physician confirmed the vessel size was appropriate, and used a 10x30 hemoshield gold despite being informed that best practice was to use a hemoshield platinum. all attempts at delivery were unsuccessful, and the decision was made to abort the impella 5.0 insertion. the patient was successfully placed on an impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13383 | 1220648-2024-24362 | 20811007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811007&pc=OZD | 06/26/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/26/2022 | 11/29/2024 |  |  | 08/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022146843 |  | 005062-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the impella 5.0 was unable to pass through vasculature due to a turn in innominate artery with several failed attempts. the decision was made to abort the 5.0 implant and implant the impella 5.5. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was patient condition related, as the impella was unable to pass through vasculature due to turn in innominate artery as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
13384 | 1220648-2024-24361 | 20811009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811009&pc=OZD | 05/29/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/29/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022064663 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. while on support, oozing was noted. heparin was put on hold until bleeding around swan site was evaluated. care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13385 | 1220648-2024-24363 | 20811015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811015&pc=OZD | 10/19/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/19/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021014990 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. the complainant reported that there was long-term support for transplantation required. impella was removed and an attempt was made to insert a central extra-corporeal membrane oxygenation, but mediastinitis was observed and percutaneous cardiopulmonary support was inserted. the patient was administered antibiotics. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the cause of the patient injury was patient condition and the relation to impella was determined to be unknown/unrelated. there is no evidence to provide a causal relationship with mediastinitis and the use of impella and thus it's relation could not be determined. this report is being filed as part of a retrospective review of historical records.
13386 | 1220648-2024-24364 | 20811021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811021&pc=OZD | 11/11/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 11/11/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492674 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. during insertion, the impella 5.0 motor housing would not take the curve of the anatomy. the procedure was aborted and an impella cp was placed instead. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13387 | 1220648-2024-24365 | 20811030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811030&pc=OZD | 09/14/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 09/14/2021 | 11/29/2024 |  |  | 04/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022076603 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | Failure to Sense | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. while on support, purge pressure low alarms were observed. the patient was in a very bad condition and care was withdrawn. it was noted that the patient expired, and that the outcome was not related to the impella. however, there were no further details received regarding the patient's condition, type of death and cause of death. given the limited information surrounding the details the patient's condition, and the cause of death, the impella cannot be excluded as a possible contributing factor. | [5 empty] | The investigation for the reported high purge pressure and placement signal issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
13388 | 1220648-2024-24366 | 20811032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811032&pc=OZD | 10/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 03/02/2022 | 11/29/2024 |  |  | 12/31/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1454565 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0 support, the patient experienced hemolysis, and it was noted it was possibly related to the impella. there was reported suction alarms that were frequent, and the impella was repositioned, and the right ventricular assist device flow rate was adjusted. the patient remained on support for 16 days, and it was noted that the patient's left heart function improved, however the right heart failure persisted. it was additionally noted that the patient expired on support due to a sudden cardiac event. it was noted there is a possibly that the impella was related to the patient's outcome, however, the patient's condition may be more likely have impacted the event. no additional information provided. | [5 empty] | The cause of the hemolysis was not determined since product was not returned and insufficient clinical details were provided. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13389 | 1220648-2024-24367 | 20811042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811042&pc=OZD | 10/25/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/25/2021 | 11/29/2024 |  |  | 04/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022082493 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | Intraoperative Cardiac Valve Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient had mitral valve injury after using 5.0 pump and mitral valve repair (mvr) was performed. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the site reported that the patient had worsened mitral regurgitation following removal of impella 5.0. mitral valve repair was performed, and the removed valve leaflet was reportedly bent. additional information indicated that the impella position was poor, and the pump was facing the back wall of the heart. therefore, the root cause of the reported heart valve damage was most likely improper positioning. this report is being filed as part of a retrospective review of historical records.
13390 | 1220648-2024-24368 | 20811044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811044&pc=OZD | 05/27/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/27/2021 | 11/29/2024 |  |  | 01/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022049370 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the purge line was leaking at the reservoir. the complication was managed, and the case continued. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13391 | 1220648-2024-24369 | 20811052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811052&pc=OZD | 05/17/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/17/2021 | 11/29/2024 |  |  | 05/31/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1434505 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. at the end of support, the device was explanted successfully, and the physician back-bled the graft, which cleared it of clots and then surgically closed the axillary site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13392 | 1220648-2024-24370 | 20811055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811055&pc=OZD | 12/08/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 12/08/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021014995 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the complainant reported that during impella placement, the patient experienced a cerebral hemorrhage. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13393 | 1220648-2024-24371 | 20811062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811062&pc=OZD | 02/17/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 02/17/2022 | 11/29/2024 |  |  | 06/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022103894 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in st elevation myocardial infarction was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported while prior and during insertion, the patient was coding intermittently. the physician was unable to get the pump to cross into the left ventricle. the patient began coding again, and the insertion attempt was paused. the medical team was unable to achieve a return to spontaneous circulation and the patient expired. although the impella catheters were not inserted, it appears a guidewire was placed based on the information at hand and therefore there is an identifiable connection between the impella device and the event. however, the patient's condition presenting may be more likely have impacted the event. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was difficult patient anatomy at the innominate. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13394 | 1220648-2024-24372 | 20811066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811066&pc=OZD | 07/07/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 07/07/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492674 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported that heparin was stopped due to bleeding out of the og (orogastric) tube. after 141.47 hours of impella 5.0 support, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13395 | 1220648-2024-24373 | 20811071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811071&pc=OZD | 05/07/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/07/2021 | 11/29/2024 |  |  | 10/31/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1455361 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. the attempt to implant the impella 5.0 was unsuccessful, and the cardiac surgeon partner came in and switched to impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13396 | 1220648-2024-24374 | 20811074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811074&pc=OZD | 06/02/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/02/2021 | 11/29/2024 |  |  | 07/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1502978 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. during impella placement, the pump flows would not increase above 0.3 on pump speed level p-8. the decision was made to remove the initial pump due to the suspicion of thrombus. a new impella 5.0 was then implanted. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13397 | 1220648-2024-24375 | 20811080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811080&pc=OZD | 06/04/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/04/2021 | 11/29/2024 |  |  | 01/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1467005 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient experienced significant blood loss from multiple sites. packed red blood cells were administered due to low hemoglobin levels. bleeding was observed at the impella pocket, and a hematoma was noted at the insertion site. vascular surgery was consulted, and the site was closed with a femstop applied to manage the complication. the patient was also septic and febrile. additional information noted the patient expired while still on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13398 | 1220648-2024-24376 | 20811081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811081&pc=OZD | 10/15/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/15/2021 | 11/29/2024 |  |  | 04/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078638 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a 5.0 supporting a patient when there were suction alarms due to low pump flow. the pump remained on for support and was troubleshot. in addition, the patient had episode of ventricular tachycardia and coded, was shocked 7 times. untimely care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13399 | 1220648-2024-24377 | 20811085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811085&pc=OZD | 08/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/09/2021 | 11/29/2024 |  |  | 01/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022039119 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the device was noted to be in a suboptimal position, with the inlet directed toward the patient¿s septum. the medical team was unable to torque the inlet towards the apex as intended. the patient¿s condition was critical, requiring a massive transfusion to maintain mechanical circulatory support. twelve units of fresh frozen plasma and red blood cells, along with ten liters of fluids, were administered. additionally, approximately six liters of blood were identified in the patient¿s abdomen, and a liver rupture was observed. the medical staff successfully managed the patient¿s bleeding. the impella device was subsequently removed, care was withdrawn, and the patient expired at the time of explant. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13400 | 1220648-2024-24378 | 20811088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811088&pc=OZD | 09/01/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 09/01/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063213 |  | 005062 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the patient experienced bleeding from all access sites. heparin was removed from the purge line to manage the bleeding complication. the left groin was catheterized, and blood products along with a femostop were required. a subarachnoid hemorrhage was identified, and the patient also experienced episodes of ventricular tachycardia and asystole. the impella device was explanted, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13401 | 1220648-2024-24379 | 20811090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811090&pc=OZD | 06/28/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/28/2022 | 11/29/2024 |  |  | 02/29/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023178991 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with unknown ethnicity in the eu has a 5.0 support ongoing for 25 days and is still supporting. the team had alarms for purge pressure and loss of signal during the support. additionally, the purge cassette was leaking, therefor it was changed with nurse, purge system condition decreased. the patient also had bleeding and due to the bleedings issues they will no longer increase the heparin. family withdrew care and patient expired, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the reported problems and its relation to the impella device was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
13402 | 1220648-2024-24380 | 20811110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811110&pc=OZD | 09/15/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 09/15/2021 | 11/29/2024 |  |  | 06/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022092119 |  | 05060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. the pump on day three had placement signal loss and suction. the pump remained on for support despite this. ultimately therapy stopped due to diagnosis of intercranial brain injury. family withdrew care, and patient expired but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13403 | 1220648-2024-24381 | 20811111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811111&pc=OZD | 01/14/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 01/14/2022 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022109794 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during washing, the impella device was dislodged from the ventricle. multiple unsuccessful attempts were made at the bedside to reposition the device. the patient was subsequently brought to the catheterization lab, where the impella was advanced across the ventricle, but it was found to be mispositioned. a decision was made to explant the impella and replace it with a new device. later, it was reported that the patient experienced a hemorrhagic stroke with a midline shift and was found to be neurologically unresponsive. additionally, the patient tested positive for heparin-induced thrombocytopenia (hit), and the purge solution was switched from heparin to bicarbonate. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13404 | 1220648-2024-24382 | 20811121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811121&pc=OZD | 04/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/09/2021 | 11/29/2024 |  |  | 10/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021018111 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was going to be implanted with an impella 5.0 for mechanical circulatory support. it was reported that the physician encountered positioning issues. the impella 5.0 was advanced to the left ventricle (lv) and was activated without issue. however, it was observed that the 5.0 appeared to be behind a papillary muscle and an attempt was made to reposition the pump. upon repositioning, the impella 5.0 fell out of the left ventricle (lv) and into the ascending aorta. all attempts to re-cross the aortic valve back into the lv failed and the impella was subsequently removed. a second attempt at implantation was repeated using all insertion steps. all attempts to reinsert the impella 5.0 failed as the doctor could not get the motor housing to recross the graft anastomoses again. it was stated that the impella 5.0 did not seem damaged in any way, but the team did not feel comfortable continuing to force the pump through the anastomoses. the decision was made to insert an impella cp instead and it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13405 | 1220648-2024-24383 | 20811126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811126&pc=OZD | 04/22/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/22/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1419831 |  | 005082-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella ld device for mechanical circulatory support. during insertion, the impella ld was advanced into r carotid artery without issue, however, was unable to cross atrioventricular with multiple unsuccessful attempts. the decision was made by the physician to abort, so the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13406 | 1220648-2024-24384 | 20811137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811137&pc=OZD | 08/11/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/11/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044170 |  | 005066-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a purge system leak occurring during priming. the purge cassette was changed and the issue resolved. there was no adverse impact to the patient. just over 1 month later, the pump stopped and could not be restarted. the pump was run beyond the intended use life. no patient harm was reported. | [5 empty] | The udi, manuf. date and expiry date of the device is unknown. should any information be received, a supplemental mdr will be filed. the investigation for the reported issue has been completed. the impella device was not received from the customer. data logs showed that the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. no abnormal performance metrics were observed prior to the pump stop. since no product was returned for investigation, the cause of the purge leak and the pump stop could not be determined. it was noted the impella was used beyond indication per design trace matrix. this report is being filed as part of a retrospective review of historical records.
13407 | 1220648-2024-24385 | 20811147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811147&pc=OZD | 07/21/2021 | 10/06/2025 |  | 11/29/2024 |  |  | 07/21/2021 | 11/29/2024 |  |  | 02/28/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | CANADA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022054768 |  | 005064-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  | White | Death | Yes | No | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, the angle of insertion made it difficult to advance the impella 5.0. a thrombus developed compromising blood flow to the patient¿s upper extremity and left axillary access was abandoned. it was reported that femoral access was gained without issue and the axillary thrombus was resolved. the device was removed, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13408 | 1220648-2024-24386 | 20811151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811151&pc=OZD | 11/08/2024 | 06/27/2025 |  | 11/29/2024 |  |  | 11/08/2024 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528948 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, pump was running on p1 due to suction alarms. reviewed ic and advised it appears to have ingested clot or biomaterial at one point on p7 with 0.0 flows as well. patient remained on support for two more days then impella was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined". based on data log analysis low pump flow was added: no product was returned for analysis. the logs confirm 0.0l/min flow while the pump was at p1and also at p7. the logs do not show any motor current spikes before low flows outside p-level changes. there were suction alarms while p-level was shifted between levels and clinical mentions that was why pump turned to p1. clinical mention action was taken to confirm pump position via echo but it is unknown if pump was in wrong position before low flows. there were no position related alarms before low pump flow occurrence. no patient related issues or volume administration was noted. the pump flows recovered and clinical mention possible clot ingestion. the root cause of the low pump flow was unable to be determined as there was no identifiable log pattern or abnormality before low flow occurrence, no product was returned for analysis, and limited clinical information was provided. b2 selection was erroneously entered on the initial manufacturer device report number 1220648-2024-24386.
13409 | 1220648-2024-24387 | 20811256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811256&pc=OZD | 07/17/2021 | 05/23/2025 |  | 11/29/2024 |  |  | 07/17/2021 | 11/29/2024 |  |  | 08/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503993 |  | 004680-AU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient had active bleeding at the extracorporeal membrane oxygenation and central line sites. the plan was to discontinue ecmo heparin and leave purge heparin as a single source for anticoagulation. the next day bleeding from ecmo sites overnight was noted and target activated partial thromboplastin time (aptt) at 40s. three days later, therapeutic aptt. the pacing wire to be transferred to other side and for evacuation of hematoma just above the clavicle, left side. the patient remained on support wans was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13410 | 1220648-2024-24388 | 20811503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811503&pc=OZD | 04/26/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/26/2021 | 11/29/2024 |  |  | 07/31/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male | 62 KG |  | Asian | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Life Threatening | Required Intervention |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.0 device for pre-operative mechanical circulatory support (mcs) and experience bleeding and ultimately expired. the patient was a transfer from another hospital five days prior to the event for proposed mitral valve replacement surgery. the patient's heart failure worsened from the acute mitral regurgitation of a papillary muscle tear and an intra-aortic balloon pump was inserted. after arrival to designated receiving hospital, the patient was patient was implanted with the impella 5.0 from the left subclavian artery. p-level was at p-8 with a flow rate of 4.5l; a transesophageal echocardiogram was completed to check device placement. however, it was noticed the patient was bleeding "considerably" during insertion of the indwelling sheath. consequently, the patient required treatment with fluid and blood transfusion (note: number of units transfused were not provided). the volume and the position of the impella were corrected and the procedure was completed. the following day follow-up noted there were no impella alarms. the patient had good urine outflow. the patient was barely breathing before insertion; however last night the almost went into cardiac arrest from hypoxia and was introduced to percutaneous cardiopulmonary support (pcps). the patient planned for cardiopulmonary rest for a few days prior to the surgery. days following now approximately day eight of impella mcs, it was noted the patient's condition suddenly deteriorated from an alveolar hemorrhage and family notified the patient was unlikely to survive if surgery was completed. at this time the patient was made a do not attempt resuscitation, and active treatment was discontinued. eventually, the patient expired. the cause of death was noted to be multi-organ failure that had progressed, and the physician stated there was no causal relationship to the impella. however, given the "considerably" access site bleeding that required giving the patient fluids and blood transfusions, and also with and impella position correction, it cannot be ruled out these may have added to an already complex situation or compounded the situation. it is possibly unlikely but unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13411 | 1220648-2024-24389 | 20811633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811633&pc=OZD | 11/11/2024 | 01/08/2025 |  | 11/29/2024 |  |  | 11/14/2024 | 11/29/2024 |  |  | 07/31/2025 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024369975 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support expired from an infection. the source of infection was unable to be identified as the patient was being supported by other devices as well as impella. the relationship of the infection to the impella is unknown. the cause of death was infection, and the causal relationship to impella is unknown. | [5 empty] | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13412 | 1220648-2024-24390 | 20811646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811646&pc=OZD | 07/30/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 07/30/2021 | 11/29/2024 |  |  | 04/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075932 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0, the pump had alarms for high purge pressure and system is blocked. a deceleration was performed with optimal results. ultimately, the patient's care was withdrawn and the procedural outcome was the patient expired. the patient was noted to have been very sick prior to impella insertion. while the high purge pressure could have potentially compromised reliability, the issue cannot be excluded as a possible contributing factor to the patient's outcome. however, the patient's condition was more likely the key factor in the outcome.. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13413 | 1220648-2024-24391 | 20811983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20811983&pc=OZD | 08/24/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/24/2021 | 11/29/2024 |  |  | 06/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022099194 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during impella removal, biomaterial was visible at the outflow of the used pump and was retained for investigation. additionally, it was noted that the patient had occlusion of the m1 segment. care was subsequently withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported stroke/cva and thrombus has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the reported observations could not be determined. this report is being filed as part of a retrospective review of historical records.
13414 | 1220648-2024-24398 | 20812112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812112&pc=OZD | 04/12/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 04/12/2022 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022059110 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support and experienced high purge pressure. approximately four days into support, there was a sharp increase in purge pressure / decrease in purge flow rate. the purge cassette was suspected to be defective, and the purge cassette was replaced. however, the event was unresolved after replacing the purge cassette. a decision was made to replace the device due to due to high risk of pump stoppage. the device was explanted and observed with no blood clots or other adhesion to it. the patient expired approximately 14 days later after explant and replacement of the initially implanted device. of note, there was no report/information that indicated the replacement device did not function as intended. the initially placed pump needed to be exchanged for purge issues, and it cannot be said given the limited amount of information that the pump exchange did not contribute to the overall outcome. | [5 empty] | The device was not returned for analysis. investigation concluded, the root cause of the purge pressure increase was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
13415 | 1220648-2024-24401 | 20812125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812125&pc=OZD | 01/13/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 01/13/2022 | 11/29/2024 |  |  | 10/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155247 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was escalated from an impella 2.5 to an impella cp as the patient required more support. during the placement of the impella cp the patient went into ventricular tachycardia requiring cardio-pulmonary resuscitation and intubation. the impella cp was able to be successfully placed, and the patient began to stabilize. while on impella cp support, the patient was bleeding from the lungs resulting in heparin being removed. the patient then developed a gastrointestinal bleed two days later, heparin was in the purge and received two units of packed red blood cells. additional information noted care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." this report is being filed as part of a retrospective review of historical records.
13416 | 1220648-2024-24395 | 20812136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812136&pc=OZD | 04/29/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/29/2021 | 11/29/2024 |  |  | 08/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1496801 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, numerous suction alarms were received. medical staff discussed problems with the patient experiencing heparin induced thrombocytopenia (hit). the patient was in poor condition. subsequently, the device was explanted, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13417 | 1220648-2024-24394 | 20812154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812154&pc=OZD | 06/25/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/25/2022 | 11/29/2024 |  |  | 11/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022148710 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia Stroke | [12 empty] | Disability | Required Intervention | Life Threatening | Death |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.0 for mechanical circulatory support. the patient was an emergent transport by acute myocardial infarction. although it was admitted that the left ventricle could be stretched, it was difficult to respond urgently, therefore extracorporeal membrane oxygenation (ecmo) and intra-aortic balloon pump (iabp) were inserted and transferred to the hospital. it was noted emergency left ventricular closure surgery would be performed after surgery to replace iabp with impella 5.0. the impella 5.0 was successfully placed and approximately five days after start of support, an alarm triggered for increase in purge pressure. the cassette was replaced without change. therefore, heparin was increased 5000 units to 25000 units. approximately three days later it was noted the doctor determined that the patient could withdraw from ecmo, but the function of the right ventricle could not keep up and gradually deteriorated. further noted, the patient also had multiple cerebral infarctions, and aggressive treatment was abandoned. the patient expired the next day, and additional information noted the impella was not related to the cerebral infarctions. no further details were provided. there is not enough evidence to exclude impella as an associated factor in the patient's demise. patient had multiple cerebral infarcts despite any best efforts, which may be unlikely related to the impella device used given the noted challenges the patient had. there are not enough details surrounding this adverse experience. additionally, there was also high purge pressure which is a product malfunction and cannot exclude this from the outcome of demised. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13418 | 1220648-2024-24396 | 20812159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812159&pc=OZD | 11/14/2024 | 12/30/2024 |  | 11/29/2024 |  |  | 11/15/2024 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had hemolysis. staff reported lactate dehydrogenase was trending up and haptoglobin trended down. the patient also developed acute kidney injury (aki) as well. the cp was explanted and replaced with an impella 5.5. the patient survived the event. | [5 empty] | The investigation for the renal failure and hemolysis has been completed. no product was returned. the date that hemolysis was noted, the clinical details stated the pump was not in an optimal position as the patient had a small left ventricle. the data logs showed that there were suction alarms two days before. there are no positioning alarms, motor current spikes or a trend in the placement signal. due to lack of sufficient information, the root cause of the hemolysis could not be determined. the root cause of the renal failure could not be determined. h.6 health effect - impact code 4641 was entered on the initial manufacturer device report number and should not have been. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13419 | 1220648-2024-24403 | 20812162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812162&pc=OZD | 06/20/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/20/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. during setup, there were low purge pressure alarms with no fluid exiting the catheter during setup. a new console and new purge cassette were attempted, to no avail. a new pump was used with no harm to the patient. | [5 empty] | D4 serial number, catalog number, lot number, and unique identifier (udi) # unknown h4 device manufacture date unknown. the investigation for the reported purge leak issue has been completed. the product was returned, and the issue was confirmed. a purge leak was found on the sidearm. a crack in the plastic was observed between the yellow luer and the reservoir. per the results of the clinical review and investigation, the root cause was determined to be sidearm yellow luer damage.
13420 | 1220648-2024-24405 | 20812171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812171&pc=OZD | 12/22/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 12/23/2022 | 11/29/2024 |  |  | 09/30/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253930 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with unknown demographics in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the white valve on the red side arm was leaking and the pin at the white valve dislocated and could not be reinserted with no further information provided. | [5 empty] | D6b explant date unknown. the investigation for the reported leak issue has been completed. the product was returned, and the issue was confirmed. the returned product was visually inspected and a damaged white flush valve was confirmed. per the results of the clinical review and investigation, the root cause was determined to be a damaged white flush valve. this report is being filed as part of a retrospective review of historical records.
13421 | 1220648-2024-24408 | 20812174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812174&pc=OZD | 11/30/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 11/30/2022 | 11/29/2024 |  |  | 08/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266419 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the patient received shockwave in order to open the lumen large enough for the impella. the physician attempted to implant the impella without success. shockwave was used a second time to open up the lumen to insert the impella. attempt number two to insert impella failed as well. at this time the physician moved on from impella cp and placed the impella rp for emergency support to try and intervene on the rca and the patient coded before intervention could begin. 35-40 mins of cpr without return of spontaneous circulation (ros. resuscitation called off, and the patient expired on the table. the physician noted that there was nothing wrong with the pump, everything worked fine. this patient just had very bad vessels and heavy calcification at every access site. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was determined to be patient condition as the patient had small/diseased vessels. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. this report is being filed as part of a retrospective review of historical records.
13422 | 1220648-2024-24409 | 20812179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812179&pc=PYX | 06/11/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/11/2021 | 11/29/2024 |  |  | 02/28/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022060842 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, there was a rp alarm high motor current and the impella stopped. the pump was restarted and the error occurred again. the physician decided against changing out the pumps due to the alarm resolving and the aic change. ultimately, care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported pump stop issue has been completed. only the cassette was received from the customer and not the pump; therefore, an evaluation of the device was not possible. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
13423 | 1220648-2024-24402 | 20812187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812187&pc=OZD | 04/30/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/30/2021 | 11/29/2024 |  |  | 12/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032839 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was unable to increase p level beyond p4 due to persistent suction alarms. the pump was explanted and replaced with an impella rp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13424 | 1220648-2024-24404 | 20812191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812191&pc=OZD | 10/04/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 10/04/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician inserted the cp sheath without complication and then put the cp in and when he tried to go around the aortic arch, it would not pass through the calcification. he then decided to use an impella 2.5 and was able to pass it around the arch. | [5 empty] | D4 serial number, catalog number, lot number, expiration date, and unique identifier (udi) # unknown. h4 device manufacture date unknown. the investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause which was patient condition related and due to calcification. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
13425 | 1220648-2024-24406 | 20812194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812194&pc=OZD | 12/05/2021 | 10/01/2025 |  | 11/29/2024 |  |  | 12/07/2021 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114346 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to implantation of the impella, the patient is already close to brain death, and it is difficult to determine whether the cause is traumatic or cardiac. while on support, the patient endured a cerebrovascular accident (cva) after device removal. the physician could not make a determination of a relationship to the device. the patient expired due to hypoxic encephalopathy that was unrelated to impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b2 required intervention was erroneously selected on the initial manufacturer device report number 1220648-2024-24406. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-24406 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - impact code 2199 was erroneously entered on the initial manufacturer device report number 1220648-2024-24406.
13426 | 1220648-2024-24393 | 20812212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812212&pc=OZD | 11/27/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 11/27/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1464083 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient's blood-tinged urine upon arrival. the patient was stable with urine output being darkened some concern for hemolysis, and the position to be confirmed tonight with echo. the patient started dialysis because of hemolysis. it was reported that the device was normally explanted and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13427 | 1220648-2024-24392 | 20812218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812218&pc=OZD | 12/08/2021 | 10/08/2025 |  | 11/29/2024 |  |  | 12/09/2021 | 11/29/2024 |  |  | 10/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022143082 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Rupture | [12 empty] | Life Threatening | Required Intervention | Death |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with postcardiotomy cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support and experienced a ventricular "rupture" resulting in the demise of the patient. the patient was transferred from a hospital under the control of extracorporeal membrane oxygenation (ecmo) intra-aortic balloon pump (iabp) one day prior to the impella implant. a surgery was scheduled but was canceled due to a thrombus in aorta. since the thrombus improved, "vsp" closure was performed. impella 5.0 was added to ecmo due to the difficulty of withdrawal from the heart-lung machine. the physician commented that the impella 5.0 device was used after vsp and upon arrival, an artificial blood vessel has been constructed under the right clavicle. after implant of the impella 5.0, the patient was transferred to the unit for continued support. echocardiogram was completed, it was noted the patient had right ventricular insufficiency, and the physician adjusted the ecmo flow rate while observing the movement of the right ventricle, aiming for ecmo withdrawal at the beginning of the said week. day four of support there was a leak in the purge cassette, and the cassette was replaced. day five of support, it was reported the patient expired due to a left ventricular "rupture," hemodynamic failure. the physician commented there was no causal relationship to the impella. no further details were provided. based on the information at hand, as captured above, the patient had a left ventricular (lv) rupture while the impella was implanted. the ventricular rupture is probably due to the pre-morbid state. it is surmised that the patient likely had vaso intervention provided there is mention of a left ventricular rupture. it was not specified that the lv rupture was spontaneous, and it was also not specified the lv rupture was related to the impella (which would be more of a perforation). the impella 5.0 is unlikely associated with the lv rupture; however, there are not enough details to make this determination even with the physician's comment of no relationship. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13428 | 1220648-2024-24397 | 20812222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812222&pc=OZD | 06/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/09/2021 | 11/29/2024 |  |  | 12/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022034447 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, the patient was actively bleeding. the patient experienced an esophageal bleed and vomited roughly three hundred cubic centimeters of blood. it was also reported that the patient was provided with one unit of red blood cells. heparin was halted and replaced with sodium bicarbonate to manage the complication. subsequently, the device was explanted, care was withdrawn, and the procedural outcome was the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13429 | 1220648-2024-24399 | 20812285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812285&pc=OZD | 06/12/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/12/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1454569 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient's left extremity was cold to the touch indicating limb ischemia, and it was difficult to find a pulse. it was reported that due to the patient being heparin resistant, argatroban was utilized in the purge. subsequently, the device was explanted, and the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D4: udi and expiration date are currently unknown. h4: the manufacturing date is currently unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13430 | 1220648-2024-24407 | 20812290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812290&pc=OZD | 06/14/2023 | 11/29/2024 |  | 11/29/2024 |  |  | 06/14/2023 | 11/29/2024 |  |  | 12/31/2024 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300756 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after switching from an intra-aortic balloon pump (iabp) to the impella cp, the patient's condition gradually improved, but then suddenly went into cardiac arrest and passed away due to multiple organ failure (mof). arrhythmia may have been the problem. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
13431 | 1220648-2024-24410 | 20812292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812292&pc=OZD | 11/15/2024 | 01/03/2025 |  | 11/29/2024 |  |  | 11/15/2024 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567793 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Defective Device | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was having an impella cp implant and during the implant, the pump was inserted over .018 wire through 14 fr short peel away sheath in right femoral artery. pump and wire prolapsed into the left iliac at the distal abdominal aorta bifurcation. the pump bent in half/kinked at the cannula. while removing .018 wire and pump through sheath, damage occurred at proximal portion of sheath just below the valve/hub of the sheath. a large amount of bleeding occurred, but sheath was able to be re wired successfully into the femoral artery. peel away sheath was removed and a new long peel away sheath was placed. the patient received a blood transfusion for the bleeding. the doctor was concerned about the integrity of the pump. the pump implant was aborted and a new pump was placed. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The product was returned for investigation. a data log review was not required for this case. several kinks were noted on the returned guidewire. additionally, a catheter kink was reported at the 20 cm mark on the pump. the returned introducer tip was damaged, and a split in the scoreline was noted, with a kink at the hub. as per the given clinical, pump was inserted over a .018 wire through a 14 fr peel-away sheath in the right femoral artery. the pump and wire prolapsed into the left iliac at the aorta bifurcation and the pump kinked. during removal, damage occurred to the sheath near the valve/hub, causing significant bleeding, which required a blood transfusion. the guidewire damage and introducer damage failure modes were ruled out as the cause, as the product damage was addressed under the delivery issue with access site bleeding, which resulted from the patient's condition. the cause of the access site bleeding issue was most likely introducer sheath split along the scorelines due to diseased/small vessels. the cause of the delivery issue was patient condition related to diseased/small vessels. the patient condition could cause the pump damage and guidewire damage.
13432 | 1220648-2024-24411 | 20812298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812298&pc=OZD | 01/30/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 01/30/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021018102 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during support, an echocardiogram was performed which confirmed a new large left ventricle [lv] thrombus. disseminated intravascular coagulation [dic] was suspected as well. patient decompensated and expired. the device cannot be ruled out as a contributing factor to the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13433 | 1220648-2024-24412 | 20812301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812301&pc=OZD | 10/01/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/01/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022062381 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. while on support, the flow and suction were unable to be monitored due to a failure of the devices position sensor. troubleshooting was performed but this did not fix the issue. the device was replaced due to a ventricular assist device (vad) transplantation. | [5 empty] | The investigation for the reported placement signal issue has been completed. the product was returned, and the issue was confirmed. the data logs from the case were not sent. no damage or abnormalities were found on the returned pump. when the pump was run, all performance metrics were within specification and no relevant placement signal alarms occurred. additionally, all metrics were visible on the aic display. per the results of the clinical review and investigation, the root cause was not determined since the failure mode could not be reproduced and the data logs were not provided. this report is being filed as part of a retrospective review of historical records.
13434 | 1220648-2024-24413 | 20812305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812305&pc=OZD | 10/11/2021 | 10/06/2025 |  | 11/29/2024 |  |  | 10/11/2021 | 11/29/2024 |  |  | 01/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1465640 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome, with an open chest for a surgical ablation, mitral and aortic valve replacement, and aortic root aneurysm repair, was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0 support, the patient experienced bleeding from the procedures that required the patient to be transfused with 7 units of blood, 2 units of platelets, and 2 units of factor vii. the patient's coagulation studies were reported as normal, however the patient continued to bleed from their chest tubes. the patient's chest was reopened to investigate the source of bleeding, but the source was unable to be located. the physician was not concerned with the impella as a contributing factor for the bleeding. the patient was transferred back to the icu on impella support, where shortly after arriving, the impella flows continued to reduce with frequent suction alarms. the patient expired while on support. given the limited information surrounding the details the patient's condition and cause of death the impella cannot be excluded as a possible contributing factor, however, the patient's condition may be more likely have impacted the event. | [5 empty] | H4 date of manufacturer is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
13435 | 1220648-2024-24414 | 20812306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812306&pc=OZD | 06/10/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/10/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1465640 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the 20-year-old male patient was transferred to the operating room for hematoma from the axillary insertion site. an arterial bleed was observed near the pectoralis minor. it was reported that the site was stitched, and heparin was stopped for the first twelve hours of the patient¿s care. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13436 | 1220648-2024-24415 | 20812319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812319&pc=OZD | 11/19/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 12/09/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1482109 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the complainant reported "worsening hemodynamics" with "buildup in outflow" cage of the device. the pump was ultimately removed, and the patient was placed on extracorporeal membrane oxygenation (ecmo), however, expired days later. no additional information was given. | [5 empty] | The investigation for the reported thrombus issue has been completed. the product was returned, and the issue was confirmed. data logs were not returned. visual inspection of the returned pump revealed the presence of biomaterial on the impeller. no other damage or abnormalities were found. per the clinical information, clot ingestion was suspected, and biomaterial was observed on the pump at the time of pump removal. per the results of the clinical review and investigation, the root cause was determined to be clot ingestion was suspected and biomaterial was observed on the pump at the time of pump removal. this report is being filed as part of a retrospective review of historical records.
13437 | 1220648-2024-24416 | 20812322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812322&pc=OZD | 05/31/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/31/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, it was noted that a blood clot was adhered near the suction site. the anesthesiologist wanted the pump to be turned off until the ultrasound could be viewed. it was determined that the impella pump had ingested a clot. an attempt to restart the impella pump was unsuccessful. the decision was made to remove the device. it was stated that a thrombus was also identified to be in the aortic valve and the process of removing the clot was initiated. | [5 empty] | D4 serial number, catalog number, lot number, expiration date, and unique identifier (udi) # unknown. h4 device manufacture date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13438 | 1220648-2024-24417 | 20812333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812333&pc=OZD | 07/31/2021 | 10/08/2025 |  | 11/29/2024 |  |  | 07/31/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1449886 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Renal Failure | [11 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella implant, the 35-year-old male patient¿s temperatures began to rise and blood cultures were required. it was reported that the patient suffered from staff infection of the blood and was placed on continuous renal replacement therapy (crrt). weaning of the impella was successful and the procedural outcome was the patient survived. | The complainant reported that during impella implant, the patient¿s temperatures began to rise and blood cultures were required. it was reported that the patient suffered from staff infection of the blood and was placed on continuous renal replacement therapy (crrt). weaning was successful and the procedural outcome was the patient survived. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the placement signal issue was unable to be determined without the product and data returned.
13439 | 1220648-2024-24418 | 20812487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812487&pc=OZD | 08/05/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/05/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1501268 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, the patient experienced access site and gastrointestinal bleeding. ten units of packed red blood cells were transfused to the patient to manage the complication. subsequently, the device was explanted, and the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13440 | 1220648-2024-24419 | 20812687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812687&pc=OZD | 08/05/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/05/2021 | 11/29/2024 |  |  | 10/31/2021 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1448540 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the patient experienced heparin induced thrombocytopenia during treatment. systemic argatroban was used. subsequently, the device was explanted, and the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13441 | 1220648-2024-24420 | 20812922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812922&pc=OZD | 02/14/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 02/14/2022 | 11/29/2024 |  |  | 10/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021018105 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, the patient experienced bleeding from the chest. it was reported that blood products were provided, and the purge fluid remained dextrose with five percent water until bleeding was resolved. the device was explanted, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13442 | 1220648-2024-24421 | 20812966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20812966&pc=OZD | 09/21/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 09/21/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1479063 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of an unknown age and gender in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. impella implantation attempted to add for unloading. it was not possible to pass the atrioventricular (no opening of it) attempts by surgeons and cardiologists; therefore, the implant was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13443 | 1220648-2024-24422 | 20813322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813322&pc=OZD | 11/05/2021 | 10/10/2025 |  | 11/29/2024 |  |  | 11/05/2021 | 11/29/2024 |  |  | 07/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022041415 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient had been experiencing ventricular tachycardia (vt) before the impella was removed, so a cardiac resynchronization therapy with a defibrillator (crt-d) was inserted. no other information was provided however the procedural outcome was the patient survived. | It was reported that a patient was on impella 5.0 for hemodynamic support. it was noted that the crt-d was inserted because there was vt before leaving impella. no other information was provided however the procedural outcome was the patient survived. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
13444 | 1220648-2024-24423 | 20813440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813440&pc=OZD | 02/03/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 02/03/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 1479063 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during support, femoral arteria was noted with multiple lesions due to calcifications. when repositioning sheath was pulled back, diffuse bleeding was noted. patient¿s vessel was cut and graft as bridge was brought in. the patient ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13445 | 1220648-2024-24424 | 20813486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813486&pc=OZD | 08/05/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/16/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1483647 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.0 for mechanical circulatory support and experienced bleeding. the patient was a transfer in and was replaced with the impella 5.0. however, the patient¿s general condition was poor, and the patient expired without improvement in cardiac function. during escalation to the impella 5.0, it was difficult for the patient to stop bleeding and blood transfusion was required (number of units were not provided). heparin was administered at 5 units/ml, but it was temporarily turned off and followed up (details are unknown for the off-time period). thereafter, the purge flow rate gradually decreased, the amount of heparin was gradually increased, and heparin was 20 units/ml. however, the situation did not improve, and the purge flow was about 1 ml/h. the patient eventually expired. the physician commented the patient's condition itself was poor, and the blood vessels were tattered. further, hemostasis was difficult, and heparin control was followed; bleeding and death are not related to the impella device. there is not enough evidence to exclude impella as an associated factor in the outcome even with the physician's comment for unrelatedness to the impella device. there was a purge issue and also a bleeding issue from the anastomotic site requiring a blood transfusion. the patient was already in a poor state, and the issue with the device and adverse event cannot be excluded as a contributing factor in the patient¿s demise. | [5 empty] | According to clinical, there was bleeding at the access sit occurred during the use of an impella 5.0. the physicians reported that the patient had poor vessel properties making it difficult to stop the bleeding. as a result, heparin was removed from the purge solution and red blood cells were given. the purge pressure increased, and purge flow decreased shortly after. heparin was increased again; however, the low purge flow did not resolve. the patient expired on the 11th day of support; the physician commented that the cause of death was unrelated to the use of impella device. no field logs were returned for analysis. however, the product was received and device analyzed. visual inspection noted the repo unit was intact, and there were no notable kinks were found in the catheter. functional testing was completed by running the pump, and it was noted the purge pressure and flow were within specification although slightly elevated. some biomaterial could be seen in the outflow cage. in conclusion the following root causes were identified. 1. the root cause of the access site bleeding was determined to be patient condition and anticoagulation management as the patient has poor blood vessel conditions and act concerns. 2. the root cause of the high purge pressure and low purge flow was most likely biomaterial buildup due to the lack of heparin in the purge system. 3. the root cause of the patient death was determined to be unrelated to the use of impella as the physician specified that there was no causal relationship between the patient death and the use of impella. this report is being filed as part of a retrospective review of historical records.
13446 | 1220648-2024-24425 | 20813491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813491&pc=OZD | 05/21/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/21/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1434496 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. while on support, the patient experienced generalized medical bleeding around the chest tubes and received 3 units of packed red blood cells. additionally, the device exhibited suction issues indicating low flow. it was also noted that the physician believed when the chest was closed the pressure change pulled up on the graft resulting in the impella being pulled back. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D4: the udi and expiration date is unknown. h4: the manufacturing date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13447 | 1220648-2024-24426 | 20813497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813497&pc=OZD | 04/09/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 04/09/2022 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492674 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. multiple attempts were made to place the impella into position without any success. the surgeon tried to manually reposition, however the impella pigtail was caught on heart structures. the impella was unable to be removed or repositioned. the patient started to decompensate as blood pressure dropped into the 40s systolic. a decision was made to do an emergent central venoarterial extracorporeal membrane oxygenation (va-ecmo). multiple pressors were initiated and 2 units of packed red blood cells were given. the device was removed and the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13448 | 1220648-2024-24427 | 20813501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813501&pc=OZD | 05/12/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/12/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1419831 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A complainant reported that a patient was admitted for a high-risk three vessel coronary artery bypass graft. it was reported that the patient had several sustained runs of ventricle tachycardia during procedure that required defibrillation. after completion of the procedure patient did not tolerate coming off bypass and decision was made to place an impella ld. it was later reported that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The udi, manufacturing date and expiry date of the impella device is not available at this time. should the device information be received, a supplemental mdr will be filed the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13449 | 1220648-2024-24428 | 20813503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813503&pc=OZD | 06/01/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/01/2021 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1459486 |  | 005080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced bleeding from the chest and heparin administration was stopped to combat the issue. it was stated that the patient required blood products as hemolysis was observed. in addition, the patient expired on support due to suffering from septic shock. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The udi , manufacturing date and expiry date of the device is unknown . the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13450 | 1220648-2024-24429 | 20813505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813505&pc=OZD | 06/07/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 06/07/2021 | 11/29/2024 |  |  | 03/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022085143 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella ld device for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced bleeding complications due to an extensive procedure. it was reported that the patient's aorta was repatched, and the chest was open post operation. it was stated that seven units of red blood cells were provided and heparin was removed from the purge in an effort to manage the complication. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13451 | 1220648-2024-24430 | 20813508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813508&pc=OZD | 07/23/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 07/25/2022 | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC |  |  | US |  | 1496784 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Male |  |  | Asian | Death | Yes | Yes | Failure to Pump | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.0 device for mechanical circulatory support and was supported for approximately two months. the patient was made a do not attempt resuscitation because no further recovery of cardiac function was expected. however, the pumped stopped before withdrawal of treatment and the patient expired with three hours of the pump stopping. additional information noted the physician provided the cause of death heart failure and further commented, the patient had irreversible heart failure and circulation was pump dependent. as a result, it became difficult to maintain the hemodynamics of the patient due to the pump stop, and as a result, the patient died. | [5 empty] | According to clinical, the pump was implanted on and after 59 days there was a pump stop. the patient was do not resuscitate with irreversible heart failure and expired a few hours following the pump stop. the product and data logs were not returned for analysis. the pump stop occurred after the 28th day impella 5.0 design life per design trace matrix. in conclusion, the cause of the pump stop was not determined as the product was not returned. the cause of the unrelated death was patient condition related.
13452 | 1220648-2024-24431 | 20813526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813526&pc=OZD | 04/01/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/01/2021 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1485428 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella 5.0 device for mechanical circulatory support. it was noted that the patient was unstable from returning from or. they will be giving packed red blood cells shortly for hemoglobin of 8.3. no purge heparin has been started yet due to bleeding at the trach site. in addition, the nurse stated that the patient¿s urine started turning red. an echo was ordered and showed that the impella was in good position. 1 unit red blood cells were given. ultimately the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13453 | 1220648-2024-24432 | 20813530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813530&pc=OZD | 03/07/2021 | 10/08/2025 |  | 11/29/2024 |  |  | 03/07/2021 | 11/29/2024 |  |  | 08/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503993 |  | 004680-AU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient presented with bleeding from coagulopathy. ultimately care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Unique identifier (udi) is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13454 | 1220648-2024-24433 | 20813533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813533&pc=OZD | 08/26/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 08/26/2021 | 11/29/2024 |  |  | 08/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495914 |  | 005060 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the patient experienced an episode of ventricular tachycardia and was cardioverted. patient was suffering from neurological disorders. in addition, it was noted that sepsis was the focus of treatment and raising fewer was depressed by coolgard. as the patient continued to suffer from neurological disorders during the case, therapy provided by the impella device was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13455 | 1220648-2024-24434 | 20813535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813535&pc=OZD | 10/26/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 10/26/2021 | 11/29/2024 |  |  | 06/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022109803 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  |  | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. impella 5.0 was advanced into axillary artery; however, unable to advance through innominate. repeated attempts were made involving various manipulations of the patient¿s arm (abduction), the graft, and the catheter. eventually, the device was removed. percutaneous transluminal angioplasty was performed on innominate prior to continued attempts to re-advance the impella 5.0. direct/open insertion as well as femoral insertion was discussed: however, the patient decompensated rapidly in the operating room with increasing difficulty to achieve adequate perfusion. the decision was made to abort impella 5.0 procedure with the patient being returned to the unit on an impella cp for hemodynamic support where family made the decision to withdraw care and patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13456 | 1220648-2024-24435 | 20813537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813537&pc=OZD | 11/10/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 11/10/2021 | 11/29/2024 |  |  | 03/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1473509 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombocytopenia, the resolution for this issue is unknown. additionally, the patient suffered a stroke. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported thrombocytopenia and stroke has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of these issues and its relation to the impella device was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
13457 | 1220648-2024-24436 | 20813543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813543&pc=OZD | 11/05/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 11/05/2021 | 11/29/2024 |  |  | 01/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1465640 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombocytopenia, and bicarbonate was added to the purge. additionally, it was reported that the physician was unable to pass the pump into the axillary artery from the graft and the axillary artery was dissected. the vessel was repaired. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13458 | 1220648-2024-24437 | 20813544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813544&pc=OZD | 02/07/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 02/07/2022 | 11/29/2024 |  |  | 06/30/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022095636 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. complainant reported a change in heparin concentration to 12.5 ie/ml due to bleeding issues. patient had small bleeding issue on right groin and was able to be resolved quickly with vacuum-assisted closure (vac) therapy. patient also experienced heparin induced thrombocytopenia [hit]. physician switched to new heparin free purge solution and changed purge cassette & solution to sodium bicarbonate. impella was successfully weaned and explanted. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13459 | 1220648-2024-24438 | 20813547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813547&pc=OZD | 02/28/2022 | 10/06/2025 |  | 11/29/2024 |  |  | 03/16/2022 | 11/29/2024 |  |  | 07/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022105542 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma. the patient was administered a blood transfusion. survival outcome at explant noted the patient expired due to multiple organ failure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-24438. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-24438.
13460 | 1220648-2024-24439 | 20813551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813551&pc=OZD | 04/09/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 04/09/2021 | 11/29/2024 |  |  | 08/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1502526 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to implantation, the patient had coded with an hour of cpr and placed on extracorporeal membrane oxygenation with an attempted 5.0 insertion. a decision was made to pull the impella 5.0 and place an impella cp. according to transesophageal echocardiogram, the left ventricle emptied very quickly with the impella at p1. in addition, the patient had been bleeding possibly due to the cpr (chest) and not able to provide enough volume. the patient was also implanted with an intra-aortic balloon pump. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output; use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states); ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13461 | 1220648-2024-24440 | 20813561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813561&pc=OZD | 03/24/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 03/24/2022 | 11/29/2024 |  |  | 04/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1476519 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding, was not able to be stabilized and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13462 | 1220648-2024-24441 | 20813563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813563&pc=OZD | 05/28/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 05/28/2022 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1487976 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. it was reported the patient was febrile and hypotensive. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13463 | 1220648-2024-24442 | 20813565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813565&pc=OZD | 06/10/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 06/10/2022 | 11/29/2024 |  |  | 06/30/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495906 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. it was reported that site bleeding was observed during this process and hemostasis was achieved through surgery. it was reported that the patient became febrile thought the night. it is unknown how the complication was managed. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13464 | 1220648-2024-24443 | 20813575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813575&pc=OZD | 10/20/2022 | 11/29/2024 |  | 11/29/2024 |  |  | 10/20/2022 | 11/29/2024 |  |  | 12/31/2023 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023164532 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, a computed tomography scan (ct) was performed and because the patient was bleeding from an unknown location. additionally, the patient experienced ventricular tachycardia throughout support. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13465 | 1220648-2024-24444 | 20813576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20813576&pc=OZD | 05/21/2021 | 11/29/2024 |  | 11/29/2024 |  |  | 05/21/2021 | 11/29/2024 |  |  | 07/31/2022 |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1492909 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient experienced abdominal compartment syndrome, confirmed by a bladder pressure of 22. exploration of abdomen was performed, and 4 liters of blood was removed from around abdomen. sodium bicarbonate was added to the purge and heparin was withheld. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13466 | MW5162997 | 20820367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20820367&pc=OZD | 11/05/2024 | 11/22/2024 |  | 11/29/2024 |  |  |  | 11/29/2024 |  |  |  |  | 11/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. | [6 empty] | 1000100 |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 27 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [5 empty] | ECMO EXTRACORPOREAL MEMBRANE OXYGENATION | [22 empty] | Initial submission | [9 empty] | Impella device alarming inappropriately. suspect issues with fiberoptic cable. device running and supporting patient. alarms included placement signal lost, retrograde flow alarms and placement in aorta. troubleshooting done echocardiogram, abiomed rep contacted. red waveform (placement signal) with sporadic reads-both high and low.
13467 | 1220648-2024-24270 | 20790037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20790037&pc=OZD | 09/03/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 09/03/2022 | 11/27/2024 |  |  | 03/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022062389 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. after meeting resistance, insertion with the impella 5.0 was aborted and the physician proceeded with implanting an impella cp device. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the delivery issues was patient condition related. the patient's vessel size was not compatible for successful 5.0 delivery. this report is being filed as part of a retrospective review of historical records.
13468 | 1220648-2024-24249 | 20790118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20790118&pc=OZD | 08/04/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 08/04/2021 | 11/27/2024 |  |  | 01/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1480624 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced shortness of breath, tachycardia. and a quick drop of blood pressure. the patient lost pulse, and despite all efforts the patient coded and expired. the relationship between the patient's outcome and the impella is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13469 | 1220648-2024-24247 | 20790505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20790505&pc=OZD | 08/02/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 08/02/2021 | 11/27/2024 |  |  | 01/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022049140 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The clinician reports the implant was inserted (b)(6) 2023 in fdi 16. on (b)(6) 2024, non-osseointegration was verified. the device was forwarded to the manufacturer. there were no reported patient injuries or complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13470 | 1220648-2024-24236 | 20790634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20790634&pc=PYX | 11/13/2024 | 01/03/2025 |  | 11/27/2024 |  |  | 11/13/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025561686 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | Hemolysis | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support. while on support, the doctor was unable to pass balloon tip catheter through 23fr sheath. sheath replacement resolved issue. additionally, impella rp flex had poor flows. clot was ingested. hemolysis was also noted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | Pump was returned for investigation cut from the catheter. there was biomaterial found around the impellers. data logs were investigated and there was a motor current rise and rise in the placement signal followed by low pump flows. therefore as per the data log review and product investigation, the root cause of the low pump flow issue was biomaterial. delivery issue was changed to introducer damage since as per the clinical information the introducer was kinked and not the pump. without the product, the failure cannot be investigated. therefore, the root cause of the introducer damage issue was not determined since product was not returned for investigation. pump was returned and biomaterial was observed around the impeller blades. data logs show rise in placement signal and motor current on the same p-level. tissue plasminogen activator (tpa) was added to purge and p-level was reduced to break clot. therefore, the root cause of the hemolysis issue was biomaterial. thrombus failure mode was deleted since that was a chain event for low pump flow and is covered as the root cause of the failure. d9 date device returned to manufacturer added. b2 outcomes attributed to adverse event was missing on manufacturer device report 1220648-2024-24236.
13471 | 1220648-2024-24256 | 20791034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791034&pc=OZD | 06/11/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 06/11/2021 | 11/27/2024 |  |  | 06/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492908 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, plans were made to continue weaning the impella as the physicians felt the impella was causing hemolysis. it was also noted that systemic heparin was put on hold due to potential bleed in axillary area associated with a jp drain. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13472 | 1220648-2024-24238 | 20791210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791210&pc=OZD | 07/30/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 07/30/2021 | 11/27/2024 |  |  | 01/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022043285 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. following impella insertion, the patient¿s liver function values progressively increased. intestinal bleeding was suspected, prompting a gastroenterological consultation. the patient¿s condition continued to deteriorate, and the decision was made to remove the impella device and withdraw care. the patient ultimately expired at the time of explant. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13473 | 1220648-2024-24254 | 20791507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791507&pc=OZD | 02/08/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 02/08/2022 | 11/27/2024 |  |  | 08/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131587 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. the complainant reported a tear in the sterile sleeve of the repositioning sheath that occurred during use of the device. there was no reported harm or injury to the patient. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported product damage (sterile sleeve damage) has been completed. the impella device has not been returned for evaluation. however, the root cause of the failure mode was a damaged sterile sleeve. this report is being filed as part of a retrospective review of historical records.
13474 | 1220648-2024-24273 | 20791550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791550&pc=OZD | 02/15/2023 | 11/26/2024 |  | 11/27/2024 |  |  | 02/16/2023 | 11/27/2024 |  |  | 01/31/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023171559 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a sensor failure occurred. calibration did not resolve the placement signal lost. additionally, the patient experienced hemolysis. although, the impella positioning was good per echo, the physician wanted to exchange pump for an impella 5.5. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
13475 | 1220648-2024-24266 | 20791626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791626&pc=OZD | 07/02/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 07/04/2022 | 11/27/2024 |  |  | 11/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022154238 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure and a purge leak. the purge cassette was replaced. it was reported that the patient survived normal explant. | [5 empty] | The investigation into the high purge pressure issue has been completed. the impella 5.0 was returned for investigation. the visual inspection of the returned device did not observe any external damage or abnormalities on the returned purge cassette. the purge cassette was cut open to perform leak reproduction testing, and leaking was observed from the piston seal. therefore, the cause of the purge leak was a piston failure. the cause of the purge leak was a piston failure. this report is being filed as part of a retrospective review of historical records.
13476 | 1220648-2024-24269 | 20791648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791648&pc=OZD | 08/01/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 08/01/2022 | 11/27/2024 |  |  | 11/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  |  |  | 2021014149 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. during implantation, the pump failed to start. another pump was successfully implanted with no reported harm or injury to the patient. | [5 empty] | The investigation for the reported pump did not start issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the pump not starting was a damaged guidewire. this report is being filed as part of a retrospective review of historical records.
13477 | 1220648-2024-24272 | 20791663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791663&pc=OZD | 01/10/2023 | 11/26/2024 |  | 11/27/2024 |  |  | 01/12/2023 | 11/27/2024 |  |  | 06/30/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023232224 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. implantation was aborted due to the patient¿s vascular anatomy. the pump was exchanged to an impella 5.5. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was received for evaluation. per the review of the clinical information and investigation, the root cause was determined to be related to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
13478 | 1220648-2024-24260 | 20791811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20791811&pc=OZD | 03/02/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 03/02/2022 | 11/27/2024 |  |  | 12/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022075891 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Death | Yes | Yes | Break | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced bleeding at the access site along with hemolysis. the clinical team decided to leave the introducer in place to achieve hemostasis. no further interventions were reported. the patient exhibited hemolytic urine and dialysate drainage, along with an elevated ldh level; however, details regarding treatment were unknown. additional information noted care was withdrawn and the patient expired due to myocardial infarction. | [5 empty] | The investigation for the reported issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the hemolysis was most likely a result of patient condition as the patient had issues with left ventricular filling. the root cause of the access site bleeding was unable to be determined as the damaged introducer was not returned for analysis. this report is being filed as part of a retrospective review of historical records.
13479 | 1220648-2024-24241 | 20792176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792176&pc=OZD | 06/27/2024 | 12/30/2024 |  | 11/27/2024 |  |  | 11/13/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486154 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured pericardial effusion with a possible relationship to the impella device used: during impella supported percutaneous coronary intervention procedure, patient was receiving cardiopulmonary resuscitation and pressor support and remained hypotensive. 2 units packed red blood cells transfused. bedside trans-thoracic echocardiography performed and revealed pericardial effusion. 2l of blood aspirated from the pericardium, and pericardium continued to fill with blood. repeat coronary angiography did not reveal any evidence of coronary perforation. it was felt that the patient probably developed free wall rupture. it was reported that the patient/event has not recovered/not resolved. the patient expired after 1.50 hours of impella support. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
13480 | 1220648-2024-24248 | 20792178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792178&pc=OZD | 05/30/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 05/30/2022 | 11/27/2024 |  |  | 10/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022145143 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient had frequent ventricular fibrillation episodes required sedation, ecmo still in place at 1,9l/min, bp 72/70 (71), impella p8 at 4,4l/min, ps 27/12. the patient expired after 391.50 hours of impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13481 | 1220648-2024-24251 | 20792368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792368&pc=OZD | 11/25/2021 | 10/06/2025 |  | 11/27/2024 |  |  | 02/22/2022 | 11/27/2024 |  |  | 07/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024926 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. during impella placement, the patient experienced bleeding at the insertion site. no further information was provided. the patient remained on impella support and was successfully weaned. | It was further reported the patient experienced cerebrovascular accident (cva) during support. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-24251. b5 additional failure mode was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24251. g1 reporting contact email revised on manufacturer device report 1220648-2024-24251 in accordance with updated procedures. h6 health effect - clinical code 1770 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24251. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-24251.
13482 | 1220648-2024-24252 | 20792466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792466&pc=OZD | 11/14/2024 | 01/07/2025 |  | 11/27/2024 |  |  | 11/14/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568586 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an oncology patient admitted in with mitral regurgitation, mild aortic stenosis / aortic insufficiency, severe tricuspid regurgitation, and other cardiac and systemic risk factors. the team was placing a 5.5 to support while awaiting decisions. the 5.5 failed to deliver and so the team placed a cp. the cp supports the patient now. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the cp was successfully placed after aborting the insertion of impella 5.5 due to resistance at aorta.
13483 | 1220648-2024-24261 | 20792574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792574&pc=OZD | 04/07/2023 | 11/26/2024 |  | 11/27/2024 |  |  | 04/07/2023 | 11/27/2024 |  |  | 06/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022148668 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Patient-Device Incompatibility | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. the impella 2.5 was prepped and placed onto a wire, but it continuously tracked into the descending aorta and would not navigate into the ascending aorta. after further attempts with viperslide to facilitate delivery, the device still failed to track into the ascending aorta, and the decision was made to abort the insertion attempts. the patient's mitral valve had growth impeding into ventricular septal wall causing sepsis and cold body temperature. the patient was removed from cardiopulmonary bypass (cpb), but expired in the operating room. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was due to patient condition related with patient having difficulty for access traversing innominate to ascending aorta. this report is being filed as part of a retrospective review of historical records.
13484 | 1220648-2024-24268 | 20792611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792611&pc=OZD | 03/08/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 03/08/2022 | 11/27/2024 |  |  | 08/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1499601 |  | 005082-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the pump had low flows which prompted explant and a planned escalation of support to the 5.5. there was no reported harm or injury to the patient. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, no abnormalities were found and the pump ran within specification. therefore, the root cause of the low pump flow was mostly likely patient condition related, as the physician stated that it could have been, due to poor atrial and ventricular function. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
13485 | 1220648-2024-24262 | 20792775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792775&pc=OZD | 05/10/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 05/10/2022 | 11/27/2024 |  |  | 07/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 1492909 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the patient still had some bleeding from the chest tubes that staff was trying to control. packed red blood cells and fresh frozen plasma were administered to the patient. it was reported that after 76.27 hours of impella 5.0 support that care was withdrawn. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13486 | 1220648-2024-24232 | 20792810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792810&pc=OZD | 09/08/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 09/08/2021 | 11/27/2024 |  |  | 06/01/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1502911 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | No | Crack | Fluid/Blood Leak | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported there was a leak in the purge set, therefore, the purge cassette was replaced. additionally, the patient developed hemolysis. treatment details were not provided. patient survived the procedure. | [5 empty] | The investigation for the reported purge leak - pump and hemolysis issues has been completed. the impella device has not been returned for evaluation. therefore, the cause of the purge leak was not determined since product was not returned. the cause of the hemolysis was not determined since the relation was unknown and the product was not returned. this report is being filed as part of a retrospective review of historical records.
13487 | 1220648-2024-24239 | 20792855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792855&pc=OZD | 10/29/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 10/29/2021 | 11/27/2024 |  |  | 07/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503732 |  | 005066 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. during implant attempt, catheter kinked. the pump was removed without harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. kinks were found in the catheter along with strain on cannula close to the motor housing. this confirmed the difficult delivery and use of force. the root cause of the delivery issues was determined to be patient condition as the physician mentioned that the patient's vessels narrowed. this report is being filed as part of a retrospective review of historical records.
13488 | 1220648-2024-24259 | 20792899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20792899&pc=OZD | 12/20/2021 | 11/26/2024 |  | 11/27/2024 |  |  | 12/20/2021 | 11/27/2024 |  |  | 03/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060529 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Failure to Sense | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was a defect on the 2.5 pump. the flush valve from the red side arm was defective and could not be flushed. blood was dripping from the flush valve itself, and no ps was displayed. the alarm was turned off, with the advice to check the pump position more often with echo. blood continued to drip from the flush valve, recommendation was given to carefully disconnect the red side arm to prevent further bleeding. ultimately, care was withdrawn and the patient expired on support. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
13489 | 1220648-2024-24255 | 20793031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20793031&pc=OZD | 03/23/2022 | 11/26/2024 |  | 11/27/2024 |  |  | 03/23/2022 | 11/27/2024 |  |  | 06/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022098337 |  | 005066 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced hemolysis, the resolution for this issue is unknown. however, patient status was noted to be remission and the patient survived normal explant. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
13490 | 1220648-2024-24265 | 20793314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20793314&pc=OZD | 06/25/2022 | 03/13/2025 |  | 11/27/2024 |  |  | 06/25/2022 | 11/27/2024 |  |  | 12/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023162534 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Death | Yes | No | Positioning Failure | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella placement, the device was removed after fulminant pump dislocation. care was withdrawn, and the procedural outcome was the patient expired at explant. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
13491 | 1220648-2024-24275 | 20795754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795754&pc=OZD | 11/14/2024 | 01/08/2025 |  | 11/27/2024 |  |  | 11/14/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477764 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemolysis | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, there was restricted flow down the impella leg so a reperfusion catheter was placed to provide distal leg flow. patient remained stable and patient was prepared for transport when he suddenly went into pulseless electrical activity with pressure in the 40s. patient was recharged showing early stent thrombosis. after cardiopulmonary resuscitation and defibrillation, the patient had additional percutaneous coronary intervention to affected arteries which returned the patient to hemodynamic stability. additionally, the patient developed hemolysis. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis, arrhythmia, and limb ischemia has been completed. the pump was not returned for investigation. data logs showed no motor current spikes, no suction alarms or positioning issues. no other information critical to hemolysis was provided. therefore, the root cause of the hemolysis issue was not determined since the product was not returned and data logs were inconclusive. the root cause of the limb ischemia and arrhythmia could not be determined without additional clinical details. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number: 1220648-2024-24275 in accordance with updated procedures.
13492 | 1220648-2024-24276 | 20795760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795760&pc=OZD | 07/21/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 07/21/2022 | 11/27/2024 |  |  | 12/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023164590 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. initial implant via the right axillary failed due to vascular anatomy. the pump could not be advanced further on shortly after graft anastomosis. the decision was made to change to the right femoral artery. surgical cut-down and graft anastomosis uneventful and a new impella 5.0 was implanted without problems. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned, and the issue was confirmed. a visual inspection of the pump revealed multiple kinks in the catheter near the motor housing. after reviewing the clinical information, it was determined that the root cause of the delivery issue was patient condition related, specifically the patient's vessel anatomy. this report is being filed as part of a retrospective review of historical records.
13493 | 1220648-2024-24277 | 20795780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795780&pc=OZD | 08/12/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 08/16/2022 | 11/27/2024 |  |  | 11/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023204415 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation of the impella 5.0 pump, that it was unable to be connected to the impella connector due to a kinked pin inside the connector. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. the returned product was visually inspected and bent gray connector pins were observed. the white patient cable was not returned for review. per the results of the clinical review and investigation, the root cause was determined to be damaged connector pins. this report is being filed as part of a retrospective review of historical records.
13494 | 1220648-2024-24278 | 20795789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795789&pc=OZD | 08/31/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 08/31/2022 | 11/27/2024 |  |  | 01/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022052317 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, there was a damaged sensor upon attempted insertion via tortuous vasculature. the pump was removed and a new impella was placed successfully. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issues was patient condition related, specifically the patient's tortuous artery. this report is being filed as part of a retrospective review of historical records.
13495 | 1220648-2024-24279 | 20795799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795799&pc=OZD | 09/16/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 09/17/2022 | 11/27/2024 |  |  | 11/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022148717 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the initial impella 5.0 pump insertion was aborted and an impella cp was subsequently implanted for patient support. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause, which was most likely patient condition related since an impella cp was able to be delivered successfully. this report is being filed as part of a retrospective review of historical records.
13496 | 1220648-2024-24280 | 20795802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795802&pc=OZD | 09/18/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 09/18/2022 | 11/27/2024 |  |  | 04/30/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023207694 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, after 18 days the pump was explanted and replaced due to purge flow issues that were not effectively remedied by the placement of tissue plasminogen activator (tpa) in the purge. the team placed a new pump without harm or injury from the interruption in the support. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the product was returned, and the issue was confirmed. the long term log confirms high purge pressure during the case. the increase in purge pressure seen in the logs aligns with heparin being removed from the purge. a visual inspection of the pump revealed a kink in the catheter between the 41 and 42cm marks. a kink in the pigtail was also observed. the catheter was cut open at the kink and it was observed that the purge line was not kinked. the pump was attached to a purge cassette and a console and was run. the high purge pressure was not reproduced. biomaterial was observed near the outflow after running the pump. per the results of the clinical review and investigation, the root cause was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
13497 | 1220648-2024-24281 | 20795815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795815&pc=OZD | 09/30/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 09/30/2022 | 11/27/2024 |  |  | 08/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022146955 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation. the initial pump insertion was aborted due to the tortuosity of the vessel, and an impella cp was implanted in its place. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical account, root cause due to patient condition. this report is being filed as part of a retrospective review of historical records.
13498 | 1220648-2024-24282 | 20795820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795820&pc=OZD | 10/05/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 10/05/2022 | 11/27/2024 |  |  | 03/31/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023193535 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during prep, the inability to introduce the 0,018" guidewire through the pigtail was noted. many guidewire entrapments reported by the surgical team. the pump was exchanged. the patient was critical but survived change to the new pump. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned, and the issue was confirmed. the pump was returned and deformations in the pigtail suggested guidewire resistance. no guidewire was returned. per the results of the clinical review and investigation, the root cause was determined to be use related. this report is being filed as part of a retrospective review of historical records.
13499 | 1220648-2024-24283 | 20795829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795829&pc=OZD | 12/14/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 12/14/2022 | 11/27/2024 |  |  | 03/31/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023209498 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient has been weaned from the impella. a critical purge flow and pressure led to pump stop at p3. the patient remained stable after the pump stop and the impella was explanted. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned, and the issue was confirmed. the data logs confirm an increase in purge pressure on the 8th day of support. no low purge pressure event was present before the purge pressure increase. motor current rose beyond the pump stop threshold on the 14th day of support and the pump stopped. the product was returned and confirmed that blood had ingressed into the purge line. no long term log was returned so the anticoagulation in the purge at the time of the increase in purge pressure is unknown. per the results of the clinical review and investigation, the root cause was determined to be blood ingress. the cause of the blood ingress was unknown. this report is being filed as part of a retrospective review of historical records.
13500 | 1220648-2024-24286 | 20795865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795865&pc=OZD | 06/28/2023 | 11/27/2024 |  | 11/27/2024 |  |  | 06/28/2023 | 11/27/2024 |  |  | 08/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131607 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the medical team had difficulty implanting the device in a patient who required impella support. multiple attempts were made, however, the medical staff was unsuccessful and aborted the procedure. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. the cause of the delivery issue was patient condition related due to pump being unable to advance past the transcatheter aortic valve.
13501 | 1220648-2024-24287 | 20795929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20795929&pc=OZD | 11/14/2024 | 09/22/2025 |  | 11/27/2024 |  |  | 11/14/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555926 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella 5.5 support and the patient had slight oozing/hematoma at site. after several days with the hematoma heparin was stopped due to continued oozing at insertion site. the staff observed and mashed out hematoma and bleeding was resolved. the patient remains on support therefore the explant date and patient outcome are unknown. | The complainant reported that the patient on impella 5.5 support had placement signal unreliable alarm. |  |  |  |  | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding major was most likely patient movement related as it occurred after being transferred. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. snr drops to 0, light is noisy, contrast is oscillating, gain is noisy, and lamp is at 100 since case start. a similar issue is on the pump prior to this one on the same console. mc is mostly stable throughout the case. ps is noisy with psnr alarms consistent throughout the case. the cause of the optical issue was not determined since product was not returned. fm was changed from placement signal issue to optical signal issue as the 5.5 does not have a dp sensor. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. added h6: for optical signal issue per a retrospective review of optical signal issues in accordance with fda recommendation.
13502 | 1220648-2024-24288 | 20796121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796121&pc=OZD | 09/06/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 09/06/2022 | 11/27/2024 |  |  | 10/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157627 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. during the coronary intervention, the right coronary artery was dissected from the origin the to mid section. during support, the impella 2.5 was not performing optimally with only 0.5lpm flows and irregular waveforms while the patient was receiving chest compressions. it was recommended to the medical team to replace the pump with either a new 2.5 impella or to escalate to an impella cp. due to the patient's heavily calcified femoral and iliac arteries, the physician opted to replace the existing 2.5 pump with a new 2.5 impella. the second impella 2.5 had optimal flows for 10 minutes. the patient was reportedly completely unloading and non-pulsatile at 1.5 lpm and below, which the physician attributed to right ventricle heart failure. the decision was made to send the patient to the icu as they were hemodynamically stable at the time, however during transport the patient's pressures dropped and never recovered. the patient expired within 15 minutes of arriving at the icu. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the event. refer to initial medwatch with date of report 24-nov-2024 for pump 2 impella 2.5 serial number (b)(6). this report is being filed as part of a retrospective review of historical records.
13503 | 1220648-2024-24289 | 20796123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796123&pc=OZD | 10/01/2018 | 11/27/2024 |  | 11/27/2024 |  |  | 05/20/2022 | 11/27/2024 |  |  | 05/31/2020 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1342102 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. after support and the impella was explanted, a thrombus was found in the surgical graft, for which a thrombectomy was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13504 | 1220648-2024-24290 | 20796157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796157&pc=OZD | 10/01/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 10/01/2021 | 11/27/2024 |  |  | 07/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022105537 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. while on support, impella had purge rate issues with high purge pressure. impella was successfully weaned and explanted. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the product was returned and no visible biomaterial blockages could be seen. a purge cassette from the lab was hooked up to the pump and the purge system was run. high purge pressure and low purge flow were reproduced. the pump was then run in a bucket of tergazyme and the purge pressure fell to within normal limits. the root cause of the high purge pressure was determined to be a biomaterial buildup. this is most likely a result of the lack of heparin used in the purge system. this report is being filed as part of a retrospective review of historical records.
13505 | 1220648-2024-24291 | 20796259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796259&pc=OZD | 10/13/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 10/13/2022 | 11/27/2024 |  |  | 10/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021008464 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.0 device for mechanical circulatory support. while on support there was increasing purge pressure and decreasing purge flow. the purge cassette was changed with no effect. the pump was explanted and a heartmate 3 inserted. a clot was found in the outflow of the impella. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the high purge pressure was most likely biomaterial since a clot was noted on explant. this report is being filed as part of a retrospective review of historical records.
13506 | 1220648-2024-24292 | 20796312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796312&pc=OZD | 09/06/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 09/06/2022 | 11/27/2024 |  |  | 11/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222901 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Low Blood Pressure/ Hypotension | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. during the coronary intervention, the right coronary artery was dissected from the origin the to mid section. it was reported that during support, the impella 2.5 was not performing optimally with only 0.5lpm flows and irregular waveforms while the patient was receiving chest compressions. it was recommended to the medical team to replace the pump with either a new 2.5 impella or to escalate to an impella cp. due to the patient's heavily calcified femoral and iliac arteries, the physician opted to replace the existing 2.5 pump with a new 2.5 impella, which is for this report. this second impella 2.5 had optimal flows for 10 minutes. the patient was reportedly completely unloading and non-pulsatile at 1.5 lpm and below, which the physician attributed to right ventricle heart failure. the decision was made to send the patient to the icu as they were hemodynamically stable at the time, however during transport the patient's pressures dropped and never recovered. the patient expired within 15 minutes of arriving at the icu. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2024-24288 for pump 1 with serial number (b)(6). this report is being filed as part of a retrospective review of historical records.
13507 | 1220648-2024-24274 | 20796343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796343&pc=OZD | 06/06/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 06/06/2022 | 11/27/2024 |  |  | 09/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022135061 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was a low purge flow exhibiting after restarting the device post-surgery. the pump was removed with no further impella support required. there was no harm to the patient. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the product was returned, and the issue was confirmed. the pump did ingest biomaterial which was determined to be the cause of the high purge pressure / low purge flow. per data log review, the 5 ml/hr purge flow and 600 mmhg purge pressure at the beginning of the run were within specification but purge flow steadily decreased, and purge pressure steadily increased throughout the run until an alarm triggered at the end. this report is being filed as part of a retrospective review of historical records.
13508 | 1220648-2024-24284 | 20796442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796442&pc=OZD | 01/27/2023 | 11/27/2024 |  | 11/27/2024 |  |  | 01/27/2023 | 11/27/2024 |  |  | 06/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022103913 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the pump was unable to advance into the aorta as a result of a calcified lesion in the subclavian artery. after multiple attempts the decision was made to abort placement of the impella 5.0 with an attempt to insert a smaller impella cp device. the cp would not pass the calcified lesion. after multiple attempts to pass the cp, the pump was removed, and the case was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause, which was patient condition related. this report is being filed as part of a retrospective review of historical records.
13509 | 1220648-2024-24285 | 20796448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796448&pc=OZD | 06/13/2023 | 11/27/2024 |  | 11/27/2024 |  |  | 06/13/2023 | 11/27/2024 |  |  | 11/30/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2024329060 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. the pump was not able to be delivered due to the vessel size being under-appreciated. the ct was done without contrast and the team found vessel too small. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. the cause of the delivery issue was patient condition related due to vascular anatomy and ct scan was without contrast and the vessel size was overappreciated.
13510 | 1220648-2024-24293 | 20796562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796562&pc=PYX | 11/14/2024 | 12/31/2024 |  | 11/27/2024 |  |  | 11/14/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529732 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when removing the impella rp the new peripherally inserted central catheter (picc) line that was placed became dislodged and coiled in the subclavian. the patient ended up getting a new picc line. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the removal issues has been completed. data logs are not relevant to the issue and product was not returned for investigation. clinical details report that while removing the rp flex as planned, it interacted with a picc line, moving it out of the desired position. based on clinical details the root cause of the removal issues was determined to most likely be interaction with another device. added d10 concomitant products.
13511 | 1220648-2024-24294 | 20796868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20796868&pc=OZD | 11/15/2024 | 01/16/2025 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541285 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, purge pressure high and purge flow low. purge lysis was in progress. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. there were no data logs or product returned. based on the clinical details provided, the root cause of the high purge pressure is most likely due to biomaterial as tissue plasminogen activator resolved the issue. d.4 revised serial number and unique identifier (udi) # as they were entered incorrectly previously on initial manufacturer device report number 1220648-2024-24294. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-24294 was submitted in accordance with updated procedures.
13512 | 1220648-2024-24298 | 20797126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797126&pc=OZD | 04/05/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/05/2021 | 11/27/2024 |  |  | 05/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1422495 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the patient was on support for twelve days and in spite of his native heart showing signs of recovery the patient expired due to sepsis, followed by disseminated intravascular coagulation. | [5 empty] | Explant date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13513 | 1220648-2024-24296 | 20797137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797137&pc=OZD | 08/05/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 08/05/2021 | 11/27/2024 |  |  | 04/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022077189 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fungal Infection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient developed a fever. a fungal infection was identified, and the medical team considered explanting the impella pump. the patient's fungal infection worsened, and the elevated temperature persisted. the device was ultimately explanted, and the patient expired while still on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13514 | 1220648-2024-24297 | 20797144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797144&pc=OZD | 11/15/2024 | 01/15/2025 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518320 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, bleeding issues were noted at the access site. the decision was made to explant the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13515 | 1220648-2024-24295 | 20797171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797171&pc=OZD | 04/04/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/04/2021 | 11/27/2024 |  |  | 06/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1486754 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported the patient had been on an impella cp and was escalated to the impella 5.0, the patient was noted to be on multiple pressors with high right heart pressures, however, right heart support was declined. the impella 5.0 was successfully inserted and the impella cp removed. while on support, high purge pressure was an ongoing issue resulting in the pharmacist running bivalirudin in the saline. when the clinician was notified of the switch by the pharmacist the nurse was advised to switch to straight dextrose 5% in water (d5w) and the clinician and nurse then performed the start/stop technique twice due to low purge flows. tissue plasminogen activator protocol was started, the purge flows were reported to have improved and were being monitored. the patient was reported to be septic, survival outcome at explant noted that the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13516 | 1220648-2024-24300 | 20797194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797194&pc=OZD | 04/06/2021 | 05/20/2025 |  | 11/27/2024 |  |  | 04/06/2021 | 11/27/2024 |  |  | 05/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1434505 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported that the impella 5.0 was successfully implanted and the patient transferred to the intensive care unit, however, a computer tomography confirmed the patient had a hemorrhagic stroke with herniation. additional information was provided that noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13517 | 1220648-2024-24299 | 20797201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797201&pc=OZD | 04/28/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/28/2021 | 11/27/2024 |  |  | 08/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1444654 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Sepsis | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a hematoma was noted at the insertion site, which was managed by withholding systemic heparin. also blood products were given as a precautionary measure to prevent further bleeding. the patient was also administered antibiotics to address a septic condition. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13518 | 1220648-2024-24302 | 20797376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797376&pc=OZD | 04/28/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/28/2021 | 11/27/2024 |  |  | 06/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1494055 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient¿s condition was critical, and the patient was maxed out on pressors with systemic vascular resistance of 410. the patient was oozing blood everywhere and one unit of packed red blood cells was administered. additional information noted the patient care was withdrawn by the family, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13519 | 1220648-2024-24301 | 20797426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797426&pc=OZD | 11/15/2024 | 12/26/2024 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551198 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, there was difficulty maintaining volume and patient was found to have a retroperitoneal bleed on impella side confirmed with computed tomography (ct) scan. the patient also noted to have bleed on ct scan on left side from arterial line access. the patient received x5 blood products overnight. the patient survived the event. | [5 empty] | The investigation for the vascular damage has been completed. no product was returned for analysis. there was no clinical mention of excessive force or relevant patient condition issue. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13520 | 1220648-2024-24303 | 20797456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797456&pc=OZD | 04/26/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/26/2021 | 11/27/2024 |  |  | 10/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 1451378 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported the impella was placed after an emergent coronary artery bypass graft was performed during forty-five minutes of cardio-pulmonary resuscitation, reportedly the team was unable to wean the patient from the heart lung device. while on support the patient developed hemolysis. additionally, it was reported that the patient experienced disseminated intravascular coagulation and massive bleeding from all insertion sites. reportedly care was withdrawn due to multiple organ failure. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13521 | 1220648-2024-24304 | 20797497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797497&pc=OZD | 04/22/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/22/2021 | 11/27/2024 |  |  | 05/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1425138 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Sepsis | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support the patient developed bleeding issues, the team was advised by clinical to decrease the purge. additionally, the patient¿s creatine kinase (ck) levels were increasing due to compartment syndrome of the left leg resulting in the sheath being removed from the left groin despite the patient¿s fragile state. the patient developed severe systemic inflammatory response syndrome with sepsis several days later and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13522 | 1220648-2024-24305 | 20797568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797568&pc=OZD | 11/15/2024 | 12/23/2024 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563793 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the device was loaded and powered on and was unable to get any substantial flow. after troubleshooting and continuing cardiopulmonary resuscitation, the decision was made to take that device out and place another cp in case there was something wrong with the first device. a second device (case: (b)(6)) was loaded and powered on, and again, there was no substantial flow from the impella. the patient expired on the second device, the separate case (patient identifier) (b)(6). | [5 empty] | Manufacturer device report number 1220648-2024-24305 was found to be a duplication of manufacturer device report number 1220648-2024-24476. any new information will be provided in manufacturer device report number 1220648-2024-24476. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
13523 | 1220648-2024-24306 | 20797644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797644&pc=OZD | 11/15/2024 | 01/07/2025 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583541 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an impella cp implant due to having acute myocardial infarction / cardiogenic shock. after the implant, the impella leg had low pulses. the leg became more ischemic and the doctor elected to remove the impella. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13524 | 1220648-2024-24307 | 20797790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797790&pc=OZD | 12/29/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 12/29/2022 | 11/27/2024 |  |  | 12/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022046792 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, after 30 minutes, the pump migrated out of optimal positioning within the left ventricle. the clinical team was unable to reposition the pump into the left ventricle, and as a result, the device was explanted, and a cp support device was placed and the patient expired shortly thereafter. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause which was calcified aorta patient condition related this report is being filed as part of a retrospective review of historical records.
13525 | 1220648-2024-24308 | 20797801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797801&pc=OZD | 04/20/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/20/2021 | 11/27/2024 |  |  | 12/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044612 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device to replace an impella cp device for mechanical circulatory support. the patient had ongoing bleeding from the impella cp access site after removal, while on impella 5.0 support. as a result of the bleeding, heparin was withheld, and the patient was taken to the operating room for surgical repair. additionally, the impella 5.0 device exhibited high purge pressure. the staff added tissue plasminogen activator (tpa) to the purge, and it was reported that patient's peak flow (pf) and pulse pressure (pp) improved. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13526 | 1220648-2024-24309 | 20797884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797884&pc=OZD | 11/15/2024 | 01/08/2025 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562471 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support placed for a patient admitted in with acute myocardial infarction and cardiogenic shock diagnosis. the pump was placed despite the patient being on multiple anticoagulants, including tenecteplase, plavix, heparin, and aspirin. the pump had a retroperitoneal bleed that prompted two units of blood and fresh frozen plasma to be infused. the pump was weaned successfully and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding is determined to be patient condition because the bleeding was reported from retroperitoneal space and not at the impella access site. hence the root cause is patient condition and unrelated to the use of impella. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13527 | 1220648-2024-24310 | 20798082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20798082&pc=OZD | 11/15/2024 | 01/15/2025 |  | 11/27/2024 |  |  | 11/15/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536601 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the patient showed dark discolored urine in the first 24 hours. there was no laboratory evidence of hemolysis in the first 48 hours. the following day, the patient showed clear signs of hemolysis, as a result the impella was removed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24310. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13528 | 1220648-2024-24311 | 20798386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20798386&pc=OZD | 10/31/2024 | 01/03/2025 |  | 11/27/2024 |  |  | 10/31/2024 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571666 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support placed for a patient admitted in for a high-risk percutaneous coronary intervention. it was reported that the pump stopped and there was retrograde flow. the pump was removed and replaced. there was no patient harm. | [5 empty] | The investigation for the pump stop has been completed. the pump was returned for investigation. according to the clinical report, the pump did not start after insertion, and the aic was showing no motor current or flow. a purge system block alarm was also observed during this event. the returned pump passed all electrical testing, and a significant amount of biomaterial was observed on the impeller. the pump was unable to start when placed in a bucket of water, so it was left soaked in the bucket of waste. some biomaterial was removed before reseating the impella. the pump was able to start, and some of the biomaterial was flushed out while running. the pump operated as expected, and the impeller was observed to be free of biomaterial. the cause of the pump did not start issue was most likely biomaterial, based on returned product investigation. added d9 device return date | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped. ¿if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
13529 | 1220648-2024-24312 | 20798764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20798764&pc=OZD | 03/16/2022 | 10/10/2025 |  | 11/27/2024 |  |  | 03/16/2022 | 11/27/2024 |  |  |  |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495914 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, after the sequential checks, and the de air procedure performed, there were persisting low pressure alarm. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
13530 | 1220648-2024-24318 | 20798793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20798793&pc=OZD | 02/14/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 02/14/2022 | 11/27/2024 |  |  | 10/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2021020027 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella ld device for mechanical circulatory support. while the physician was washing out the chest, the patient went into ventricular fibrillation (vf), requiring approximately 20 shocks with internal paddles throughout the case. after the implantation of the protek duo, the patient developed significant clots in the right ventricle (rv). the impella device was removed at the bedside. additional information was provided that noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13531 | 1220648-2024-24319 | 20799025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799025&pc=OZD | 12/15/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 12/15/2021 | 11/27/2024 |  |  | 11/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024738 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. the patient tested positive for heparin-induced thrombocytopenia (hitt), prompting the removal of heparin from the purge and the infusion of bicarbonate instead. argatroban was administered systemically but was discontinued due to concerns regarding a potential gastrointestinal bleed. additional information was provided that noted care was ultimately withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13532 | 1220648-2024-24320 | 20799078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799078&pc=OZD | 03/23/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 03/23/2022 | 11/27/2024 |  |  | 12/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023168325 |  | 005060 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.0 and extracorporeal membrane oxygenation for mechanical circulatory support. the patient had received the bypasses and could not be weaned off. the patient needed alot of volume to keep both devices running. it was reported that due to patient condition, the impella could not run in stable condition at p-3 0,5 l/min. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported placement signal issue has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13533 | 1220648-2024-24321 | 20799101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799101&pc=OZD | 09/05/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 09/05/2021 | 11/27/2024 |  |  | 08/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1447360 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported an infection that occurred from the surgical graft site and puss needed to be evacuated. two days later the impella was removed and intra-aortic balloon pump was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported infection/sepsis issues has been completed. the impella device has not been returned for evaluation. review of the provided data logs showed that the pump, sn (b)(6), was run for approximately 19 days. it was reported that the infection was not alleged to have been caused by the impella device. the root cause of the infection was determined to be unrelated to the impella device. this report is being filed as part of a retrospective review of historical records.
13534 | 1220648-2024-24313 | 20799335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799335&pc=OZD | 03/08/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 03/08/2022 | 11/27/2024 |  |  | 06/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495906 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced a fever of unknown etiology. actions taken are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13535 | 1220648-2024-24315 | 20799428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799428&pc=OZD | 11/12/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 11/12/2022 | 11/27/2024 |  |  | 09/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022146906 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. during impella placement, the tissue plasminogen activator (tpa) protocol was initiated by the medical staff due to high purge pressures. tpa was infused, resulting in a significant decrease in purge pressure. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13536 | 1220648-2024-24314 | 20799584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799584&pc=OZD | 05/25/2023 | 11/27/2024 |  | 11/27/2024 |  |  | 05/26/2023 | 11/27/2024 |  |  | 08/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131610 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Pump | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. following insertion, a yellow "impella defective" alarm was triggered. the impella 5.0 was removed and replaced with an impella 5.5. the patient expired after the insertion of the impella 5.5. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into the impella has been completed. no product was returned. per the data log analysis show the 5.0 is initially connected to the aic and the raw placement signal is below 0 and out of range, however the signal is not completely out. this results in a ps unreliable alarm. the pump is then disconnected and reconnected, likely to try to fix the ps unreliable alarm. this results in an impella defective alarm. this is likely due to an issue with the differential pressure sensor, as the circuit for the differential pressure sensor and the electrically erasable programmable read-only memory share the same power. without returned product, the exact root cause of the differential pressure sensor issue and unreliable alarm could not be determined. this report is being filed as part of a retrospective review of historical records.
13537 | 1220648-2024-24316 | 20799614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799614&pc=OZD | 12/04/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 12/04/2022 | 11/27/2024 |  |  | 07/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022115381 |  | 004680-AU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. while on support the patient was reported to have sepsis with a fever and experienced febrile episodes throughout support. it was reported that the patient also was suspected to have heparin induced thrombocytopenia resulting in the purge being changed from heparin to sodium bicarbonate with bivalirudin. the patient¿s condition was reported to have deteriorated. computer tomography revealed multiple clots that were surgically removed. reportedly on (b)(6) 2022, the patient underwent video assisted thoracoscopic revision of the central line and right anterior thoracotomy for pleural effusion due to hematoma and blood collection. additionally, on (b)(6) 2022, there were multiple thrombus reported to the spine and aorta that were being managed with upper levels of therapeutic anticoagulation. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13538 | 1220648-2024-24317 | 20799629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799629&pc=OZD | 08/31/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 08/31/2021 | 11/27/2024 |  |  | 08/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1445064 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported the inability to deliver the pump due to patient's anatomy. insertion was aborted and the patient was supported with an intraaortic balloon pump. there was no adverse impact to the patient as a result. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issues was determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
13539 | 1220648-2024-24323 | 20799719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20799719&pc=OZD | 06/20/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 06/20/2022 | 11/27/2024 |  |  | 03/31/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023195881 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, bleeding was noted at the location of the patient¿s chest. the device was explanted, and the procedural outcome was the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13540 | 1220648-2024-24322 | 20800071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20800071&pc=OZD | 04/28/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/28/2021 | 11/27/2024 |  |  | 09/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021018101 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [10 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support the patient coded and experienced ventricular tachycardia and ventricular fibrillation requiring eight rounds of cardio-pulmonary resuscitation. oozing at site after cpr. the team was able to successfully restore the patient¿s vitals. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13541 | 1220648-2024-24324 | 20800378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20800378&pc=OZD | 04/20/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 04/20/2022 | 11/27/2024 |  |  | 09/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022133729 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient admitted in with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was noted that the impella 5.0 was implanted as a bridge-to-lvad, and the procedure was complicated by high purge pressure, despite the patient being treated with argatroban for heparin-induced thrombocytopenia (hit). the cardiologist initiated tissue plasminogen activator (tpa) therapy that afternoon. the patient was considered stable under impella support, although high purge pressure persisted despite the tpa infusion in the purge solution. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13542 | 1220648-2024-24325 | 20800523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20800523&pc=OZD | 07/17/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 07/17/2022 | 11/27/2024 |  |  | 03/31/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023193544 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Death | Yes | No | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was difficult insertion using transesophageal echocardiography [tee], as the impella was unable to pass through the patient¿s vasculature. it was noted that the patient¿s main problem was bleeding. heparin was withheld. upon examination it was noted that the device was to close to the atrioventricular canal and repositioning attempted. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
13543 | 1220648-2024-24326 | 20801016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801016&pc=OZD | 08/03/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 08/02/2022 | 11/27/2024 |  |  | 09/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504534 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Fluid/Blood Leak | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was a low purge pressure alarm that could not be resolved by changing purge system and increasing purge fluid glucose concentration to 10%. reducing pump flow from p9 to p5 made the alarm stop. later, a cerebral hemorrhage was observed and the patient ultimately expired on support after cardiac arrest. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs were reviewed and confirm a purge drop when the p-level was raised. the root cause of the low purge pressure was unable to be determined as no product was returned for evaluation. this report is being filed as part of a retrospective review of historical records.
13544 | 1220648-2024-24328 | 20801303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801303&pc=OZD | 04/25/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 04/25/2021 | 11/27/2024 |  |  | 01/31/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022051475 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced a hematoma. it was reported that a vascular surgery of the anastomoses site was performed with a graft then sutured around the graft and the sheath. the surgeon also reportedly repaired the muscle tear that was bleeding and the groin site. reportedly, the patient received large amounts of blood products during the repair procedure resulting in pressure stabilizing and the patient regaining pulsatility. additionally, the patient subsequently coded and the family made the decision to withdraw care. the survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13545 | 1220648-2024-24327 | 20801326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801326&pc=OZD | 10/02/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 10/02/2022 | 11/27/2024 |  |  | 09/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013231 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. the impella 5.0 was inserted via the right axillary artery, and the patient was transferred to the cardiac care unit for further management. during the course of support, the patient experienced frequent ectopy and runs of ventricular tachycardia, despite adjustment of the pump position to 45 cm and the administration of a lidocaine drip, which did not provide relief. additional information noted the patient expired while still on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13546 | 1220648-2024-24331 | 20801455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801455&pc=OZD | 05/10/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 05/10/2021 | 11/27/2024 |  |  | 05/31/2021 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1428037 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, bleeding from the anastomotic site and disseminated intravascular coagulation occurred, requiring blood to be transfused to the patient. the weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13547 | 1220648-2024-24330 | 20801562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801562&pc=OZD | 11/18/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 11/18/2021 | 11/27/2024 |  |  | 05/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1478768 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported a pump failing to advance due to patient's anatomy. the pump was removed and exchanged for an impella cp. there was no harm to the patient. | [5 empty] | The investigation into the delivery issue has been completed. the impella 5.0 was not received from the customer and therefore, an evaluation of the device by the manufacturer was not possible. based on clinical information the impella 5.0 being unable to advance is most likely due to patient condition as the physician was unable to advance past the graft anastomosis due to the patient anatomy. this report is being filed as part of a retrospective review of historical records.
13548 | 1220648-2024-24333 | 20801604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801604&pc=OZD | 12/05/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 12/05/2021 | 11/27/2024 |  |  | 01/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1464575 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. on insertion to the left ventricle the wire got wrapped around the device and made removing the wire without the device difficult. both the device and the wire were removed and a new impella 5.0 was implanted to support the patient. there was no harm to the patient. | [5 empty] | The investigation into the delivery issue has been completed. the impella 5.0 was not received from the customer and therefore, an evaluation of the device by the manufacturer was not possible. the clinical details provided were insufficient to determine a root cause, also the data logs and product were not returned, therefore a root cause of the guidewire issues was unable to be determined. h.4 device manufacture date is unknown. this report is being filed as part of a retrospective review of historical records.
13549 | 1220648-2024-24329 | 20801640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801640&pc=OZD | 05/01/2021 | 11/27/2024 |  | 11/27/2024 |  |  | 05/01/2021 | 11/27/2024 |  |  | 12/31/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037253 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced a hematoma at the impella insertion site. vascular surgery did not recommend further exploration of the site. the patient was administered blood and platelet transfusions. additionally, there was noted oozing and drainage from the axillary insertion site. later, the clinical staff observed that the patient¿s lower extremities became cold and exhibited signs of mottling. action taken to resolve the issue are unknown. additional information noted care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13550 | 1220648-2024-24332 | 20801684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20801684&pc=OZD | 11/26/2021 | 10/09/2025 |  | 11/27/2024 |  |  | 11/29/2021 | 11/27/2024 |  |  | 06/30/2022 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489795 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.0 placement, the patient experienced high purge pressure requiring initiation of the tissue plasminogen activator protocol. the device was explanted, and the procedural outcome was the patient survived. | The complainant reported that during impella placement, the patient experienced high purge pressure requiring initiation of the tissue plasminogen activator (tpa) protocol. the device was explanted, and the procedural outcome was the patient survived. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Data logs showed the purge pressure gradually increasing and the purge flow decreasing with occurrence several "purge pressure high" alarms. the purge pressure continued to remain elevated until the date of the reported tpa protocol. blood was found in the purge gap. when the pump was connected with returned purge cassette, the purge pressure was elevated but no alarms occurred. the pump was then run, and the purge pressure decreased. in conclusion, it was determined that the cause of the high purge pressure was biomaterial buildup occluding the purge gap.
13551 | 1220648-2024-24334 | 20802060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20802060&pc=OZD | 08/17/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 08/17/2022 | 11/27/2024 |  |  | 04/30/2024 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023206210 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support there was slow bleeding, blood was observed exiting the red plug of the impella. the leakage possibly caused by the complex implantation due to difficult patient anatomy (vessel dimension and tortuosity). no further information was provided therefore explant date and patient outcome are unknown. | [5 empty] | The investigation for the reported product damage (hole in catheter) issues has been completed. the product was returned and blood was confirmed in the red handle. the entry point for the blood was found to be a puncture at the 30 cm mark. the root cause of the product damage was determined to be a hole in the catheter that occurred from improper use. this report is being filed as part of a retrospective review of historical records.
13552 | 1220648-2024-24335 | 20802277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20802277&pc=OZD | 11/16/2022 | 11/27/2024 |  | 11/27/2024 |  |  | 11/16/2022 | 11/27/2024 |  |  | 11/30/2023 |  | 11/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161290 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the physician was unable to cross the graft for left axillary insertion. the physician stated that he looses torque ability on the catheter once inside the graft, so an impella cp has been successfully placed instead. it was further reported that the patient expired and the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was most likely a steep insertion angle. the recommended angle of insertion is 35-40 degree. this report is being filed as part of a retrospective review of historical records.
13553 | 1220648-2024-24176 | 20777080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20777080&pc=OZD | 09/15/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 09/15/2022 | 11/26/2024 |  |  | 07/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022111092 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited an error in the purge system error indicating high purge pressure, the resolution for this issue is unknown. additionally, the patient experienced sepsis. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13554 | 1220648-2024-24172 | 20777726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20777726&pc=OZD | 08/09/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 08/09/2022 | 11/26/2024 |  |  | 05/31/2024 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023211893 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding from the thorax and anticoagulant was restricted. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13555 | 1220648-2024-24177 | 20777737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20777737&pc=OZD | 10/02/2024 | 12/24/2024 |  | 11/26/2024 |  |  | 10/31/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531413 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella cp device used: ¿patient developed marked hematuria and was noted to have low haptoglobin with elevated ldh concerning for hemolysis in the setting of impella. patient s/p bedside impella removal on (b)(6) worsening renal function is due to hemolysis hemoglobin was down trending likely due to hemolysis at postop procedure, patient was free of signs or symptoms of bleeding, she started to have significant hemolysis despite good position of the impella on echo with hematuria and rising ldh and aki. impella was removed bedside (b)(6) 2024. creat plateua at 3s with 5 gram proteinuria.¿. the patient recovered with no sequelae. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute renal failure with a "possible" relationship to the impella device used: ¿patient with up trending creatinine, possibly secondary to hypovolemia and contrast nephropathy. (b)(6) renal ultrasound showed no hydronephrosis nephrology was consulted due to persistently up trending creatinine workup was revealing for worsening renal function is due to hemolysis and atn/contrast induced nephropathy. protein electrophoresis was negative for multiple myeloma, peripheral smear was negative for blasts/dysplasia/schistocytes renal function remained stable in the last 3 days prior to discharge. educated that renal function has plateaued and most likely start to improve. she is making urine and asymptomatic. creatinine baseline 1.2 peak 3.79¿. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and renal failure has been completed. data log were reviewed finding no positioning issue alarms observed. a few suction alarms found on the same day issue reported. there is a motor current (mc) drop with drop in flows when there was no change in p- levels. no clear indications found to confirm the rc. the cause of the hemolysis issue cannot be determined since the compliant devise not returned for analysis and lack of clinical details. the cause of the renal failure issue cannot be determined since the compliant devise not returned for analysis and lack of clinical details.
13556 | 1220648-2024-24185 | 20778367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20778367&pc=OZD | 06/22/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 06/22/2021 | 11/26/2024 |  |  | 01/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022039116 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure this was resolved by changing the purge fluid. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13557 | 1220648-2024-24202 | 20778694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20778694&pc=OZD | 09/28/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 09/28/2022 | 11/26/2024 |  |  | 09/30/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1502909 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Peeled/Delaminated | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported just after delivering the impella 2.5 into the left ventricle, the pump was found to have shifted back across the valve. when attempts to re-cross the valve failed, the device was removed for a complete reinsertion attempt. however, after successfully crossing the valve a second time, the guidewire was caught within the pump and the pump was once again removed. attempts to subsequently backload the pump over a third wire resulted in damage to the peel away sheath and the decision was made to use a new pump. insertion with the new pump was successful without further issue. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the delivery issue has been completed. the clinical details provided were sufficient to determine the root cause. the root cause of the positioning issues was wireless re-access due to improper technique by the end user. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
13558 | 1220648-2024-24193 | 20779325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20779325&pc=OZD | 10/11/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 10/14/2021 | 11/26/2024 |  |  | 10/31/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1450376 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient had a persistent fever and was infected, this was treated with antibiotics. it was also reported that the purge cassette was replaced due to a drop in purge flow. it was reported that the patient survived normal explant. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge leak, high purge pressure, and infection/sepsis has been completed. the product was returned, and biomaterial was found in the outflow and around the purge gap. high purge pressure was reproduced using a lab purge cassette. the pump was soaked in tergazyme and was able to start. as the pump ran it ejected the biomaterial, and the purge pressure dropped to within a normal range. the root cause of the high purge pressure was determined to be ingested biomaterial. the root cause of the infection was determined to have an unknown relationship to the impella as no evidence was provided to suggest a causal relationship between impella and the infection. the cause of the high purge pressure was ingested biomaterial occluding the purge gap. this report is being filed as part of a retrospective review of historical records.
13559 | 1220648-2024-24203 | 20780506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780506&pc=OZD | 10/13/2021 | 05/27/2025 |  | 11/26/2024 |  |  | 10/13/2021 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support purge bypass was completed when checked into unit as transfer to standard configuration was unsuccessful. it was reported that the patient was stable. | [5 empty] | D4, h4: this information cannot be provided, the serial number is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13560 | 1220648-2024-24201 | 20780675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780675&pc=OZD | 02/08/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 02/08/2022 | 11/26/2024 |  |  | 08/31/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1436980 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella for mechanical circulatory support. during support, the patient experienced limb ischemia, it was reported that the patient outcome was recovered. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13561 | 1220648-2024-24204 | 20780755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780755&pc=OZD | 03/18/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 03/18/2022 | 11/26/2024 |  |  | 06/30/2023 |  | 11/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128871 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that due to stenosis in the patient's vasculature, the impella pump would not cross into the aortic valve. it is unknown how the complication was managed. it was reported that the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13562 | 1220648-2024-24200 | 20780785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780785&pc=OZD | 03/07/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 03/07/2022 | 11/26/2024 |  |  | 07/22/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1491150 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was going to be on an impella 5.0 for mechanical circulatory support. during delivery, the physician had difficulty advancing the pump once coming out of the sheath due to patient disease (calcium/stenosis). it was noted that the red lumen was reportedly sticking to the guidewire as it was advancing, and the red lumen was cut. an attempt to deliver the impella 2.5 was made but it would not advance out of the sheath due to calcification ad stenosis of the vessel. the physician decided to abort the implant. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the delivery/removal issue has been completed. the root cause of the delivery issues was determined to be patient condition per clinical detail review as the patient had calcified and stenotic vessel. additionally, it was noted that there was a red lumen removal issue. the cause was the component. this report is being filed as part of a retrospective review of historical records.
13563 | 1220648-2024-24196 | 20780818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780818&pc=OZD | 04/05/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 04/05/2022 | 11/26/2024 |  |  | 06/30/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489795 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device had a loss of placement signal. the team observed this prior to implant and attempted all troubleshooting. when this failed the device was replaced and support proceeded without patient harm. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, the pump plug was opened and no damage or abnormalities were observed within it. therefore, it was determined that the cause of the placement signal issue was an open electrical line stemming from manufacturing. this report is being filed as part of a retrospective review of historical records.
13564 | 1220648-2024-24199 | 20780917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780917&pc=OZD | 03/17/2023 | 11/25/2024 |  | 11/26/2024 |  |  | 03/17/2023 | 11/26/2024 |  |  | 06/30/2024 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023256549 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the pump stopped and was explanted. no further information is available. | [5 empty] | The investigation for the reported display issue has been completed. the impella device was not received from the customer. the data logs and clinical details provided were sufficient to determine a root cause. data logs show a gradual loss in motor current pulsatility, a gradual rise in purge pressure, saturated flows, frequent drops in motor speed in the 24 hours leading up to the pump stop. within 12 hours of the pump stop, there were increasingly frequent motor current spikes until the "impella stopped - motor current high" alarm triggered, and the pump suddenly stopped. in totality, the pump performance metrics show strong evidence of biomaterial interaction with the motor over time. therefore, it was determined that the cause of the pump stop was most likely biomaterial buildup on the motor. this report is being filed as part of a retrospective review of historical records.
13565 | 1220648-2024-24192 | 20780956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20780956&pc=OZD | 09/02/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 09/02/2021 | 11/26/2024 |  |  | 03/31/2023 |  | 11/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022067966 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 for mechanical circulatory support. the automated impella controller (aic) experienced bent pins. no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. patient history showed that after two attempts to insert an impella 5.0, an impella cp was successfully implanted. therefore, the root cause of the delivery issue was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
13566 | 1220648-2024-24195 | 20781066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20781066&pc=OZD | 10/25/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 10/25/2021 | 11/26/2024 |  |  | 03/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022064741 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was going to be implanted with an impella 5.0 device for mechanical circulatory support. the 5.0 pump was unable to pass through the subclavian, got caught up at the motor housing of the 5.0, a 5.5 pump was inserted and was successful, and there was no mention of patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
13567 | 1220648-2024-24197 | 20781911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20781911&pc=OZD | 05/09/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 05/09/2022 | 11/26/2024 |  |  | 06/30/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021009005 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device had a purge leak, purge pressure decreased, and the purge flow increased. it was noted that the staff thought that the purge liquid could not be supplied to the pump. the physician decided to explant the device, the patient was hemodynamically stable and survived explant. | [5 empty] | The investigation for the reported pump leak issues has been completed. the impella device was not received from the customer. therefore, the occurrence and location of the leak is undetermined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
13568 | 1220648-2024-24190 | 20781969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20781969&pc=OZD | 08/30/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 08/30/2021 | 11/26/2024 |  |  | 04/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078669 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported a positioning issue during use of the device that prompted removal and replacement. there was no adverse impact to the patient. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the positioning issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
13569 | 1220648-2024-24189 | 20781988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20781988&pc=OZD | 08/30/2024 | 12/20/2024 |  | 11/26/2024 |  |  | 11/06/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured stroke with a "possible" relationship to the impella cp device used: ¿patient had impella placement on (b)(6)2024 and discharged home on (b)(6)2024. she had a sudden onset of confusion the same evening, ems was called and she was taken to ed. code stroke was called. mri brain showed acute left frontoparietal mca territory infarct measuring 4.1 cm. she experienced worsening aphasia, difficulty speaking and right sided weakness. she continued to decline during her hospitalization with decreased responsiveness. was discharged home with hospice on (b)(6)2024 and expired on (b)(6)2024.¿ | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
13570 | 1220648-2024-24194 | 20782138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20782138&pc=OZD | 10/19/2021 | 11/25/2024 |  | 11/26/2024 |  |  | 10/25/2021 | 11/26/2024 |  |  | 03/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022062391 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 20 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Decrease in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. there was a liquid leak from the pressure transmitter part of the purge cassette, another cassette was utilized. additionally, the patient experienced hemolysis, the resolution for this issue is unknown. it was reported that the patient survived normal explant. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device has not been returned for evaluation and review of the data logs did not identify the root cause of the purge cassette leak at the pressure transmitter. however, the probable cause of the purge leak was a pressure transmitter leak. the cause of the hemolysis was unable to be determined. this report is being filed as part of a retrospective review of historical records.
13571 | 1220648-2024-24191 | 20782677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20782677&pc=OZD | 12/29/2021 | 07/09/2025 |  | 11/26/2024 |  |  | 12/29/2021 | 11/26/2024 |  |  | 01/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1467005 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella implant, a placement signal alarm was received during device priming. it was reported that the flows did not resolve, and a new device was requested to manage the complication and provide a reliable signal. it was reported that the process of impella implantation was aborted. no patient harm was reported. | [5 empty] | Returned data logs show the raw placement signal at 5000 mmhg from the start of the case, indicating an open line. the pump was returned and an open line in both the v+/v- and s+/s- lines was confirmed. the exact cause of the open line was not determined. of note, the ps wiring was caught under the purge kink protector upon opening of the pump handle. when removed from under the kink protector, ps wire damage was noted. added actual investigation details above to h10 (product was returned). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
13572 | 1220648-2024-24198 | 20782722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20782722&pc=OZD | 03/07/2023 | 11/25/2024 |  | 11/26/2024 |  |  | 03/07/2023 | 11/26/2024 |  |  | 05/31/2024 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023226757 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, a placement signal malfunctioning was observed and a loss of blood from the connector between the catheter shaft and red impella plug was noticed by the hospital staff. the device was still utilized and the patient survived normal explant. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the malfunction and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
13573 | 1220648-2024-24205 | 20782932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20782932&pc=OZD | 01/19/2022 | 11/25/2024 |  | 11/26/2024 |  |  | 01/19/2022 | 11/26/2024 |  |  | 10/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022027983 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. after removal, the perclose failed, so angio seal was attempted, up and over balloon tamponade for hemostasis of lymphocyte function-associated antigen (lfa). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13574 | 1220648-2024-24209 | 20783445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783445&pc=OZD | 07/08/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/08/2021 | 11/26/2024 |  |  | 03/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022064709 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. the pump was explanted due to high purge pressure after 6 days of support. the patient was then placed on intra-aortic balloon pump (iabp) but, when the support was not sufficient, impella 5.0 pump was placed again. no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported high purge pressure has been completed. the returned pump was purged without issue. the high purge pressure alarms did not reproduce, and no damage was observed that would have contributed to a high purge pressure alarm. the low purge pressure alarms reproduced at high speeds (greater than p-3). no leaks in the purge system were observed. of note, no purge cassette was returned with the pump for analysis. the root cause of the low purge pressure was most likely due to low purge flow path resistance, as no leaks were found in the purge system and the pump was naturally purging at pressure on the lower end of the specification. this report is being filed as part of a retrospective review of historical records.
13575 | 1220648-2024-24225 | 20783572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783572&pc=OZD | 11/13/2024 | 01/03/2025 |  | 11/26/2024 |  |  | 11/13/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568130 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a patient with cardiac history. the acute myocardial infarction / cardiogenic shock prompted pump placement, and the support was ongoing when after seven days the site had a bleed. the bleed prompted the team to stop heparin. the ooze did not prompt pump explant. pump remains on for support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be patient movement because the patient was agitated and anxious the night before. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13576 | 1220648-2024-24222 | 20783615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783615&pc=OZD | 08/31/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/31/2021 | 11/26/2024 |  |  | 02/28/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044183 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. the patient had an infectious disease causing a cerebral aneurysm. due to the prolongation of the infectious disease, bleeding occurred from the aneurysm, and pump was operating without any problems. the pump stopped 39 days after the start of support which is beyond the recommend instructions for use (ifu). pump was successfully weaned and explanted. | [5 empty] | The investigation for the reported stroke/cva and pump stop has been completed. the impella device has not been returned for evaluation. the reported intracerebral hemorrhage was determined to be unrelated to impella as no product was returned and no evidence was provided to indicate that the event occurred or that it was related to the use of impella. additionally, the site reported that there was a pump stop after 39 days of support which was preceded by a decrease in purge flow and increases in purge pressure and motor current. a specific cause for the pump stop could not be determined through this investigation as no product or data logs were provided for evaluation. however, the pump stop occurred beyond the 28 day indication for use per design trace matrix 0046-9717. this report is being filed as part of a retrospective review of historical records.
13577 | 1220648-2024-24216 | 20783683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783683&pc=OZD | 10/16/2024 | 01/06/2025 |  | 11/26/2024 |  |  | 11/13/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, there were ongoing placement signal not reliable alarms. position was confirmed. audio was disabled. investigation team requested that the automated impella controller (aic) be investigated as it may be responsible for the optical issues. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) loss data has been completed. console logs from the reported day of the event were consistent with the complaint. the logs showed multiple instances of the following error algs suspended opm signal invalid," and "invalid bad snr sample mask" due to signal-to-noise ratio (snr) falling below the minimum threshold. subsequently, several placement signal not reliable alarms and placement signal not reliable alarm were observed. the long-term log data indicated that the unreliable placement signal persisted throughout most of the case, varying in intensity, and this correlated with the low snr. the console was sent back to field service for further inspection. while the reported issue could not be reproduced, intermittent signal distortions (envelope/fringe) were observed in the analog output from the ccd of the optical bench. this localized defect in ccd array caused the optical bench to pass the 5s testing but fail when used with the pump. the fixed test cavity interference fringe, located outside the defect region during the 5s testing, overlaps the defect when the pump sensor is exposed to pressure, resulting in low snr. the root cause of the issue was identified as intermittent distortions in the optical bench signal, likely caused by a malfunction in the optical bench or its components. d9 device return date/information was added. b7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24216. d2 revised common device name since the initial manufacturer device report number 1220648-2024-24216 was submitted in accordance with updated procedures.
13578 | 1220648-2024-24215 | 20783770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783770&pc=OZD | 11/12/2024 | 06/27/2025 |  | 11/26/2024 |  |  | 11/12/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562381 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing for a st elevated myocardial infarction patient admitted in with extensive cardiac history and care escalation from an intra-aortic balloon pump. the patient was sent for coronary artery bypass graft. postoperatively there was a purge filter leak that did not prompt pump replacement. also there was limb swelling and ventricular tachycardia (vt). the vt was treated by multiple shocks. the pump remained on for support until wean and explant. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of vf/vt/af/at and purge leak - pump has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24215.
13579 | 1220648-2024-24228 | 20783780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783780&pc=OZD | 06/16/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 06/16/2022 | 11/26/2024 |  |  | 06/30/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489284 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.0 had low pump flow. ongoing suction alarms were noted for which they administered fluid in order to reduced p level. the ps became unreliable. walked through zeroing the sensor with no effect. explained the suction could be false and to use the estimated flow chart to determine impella flow. csc called in to help silencing suction alarms. the rn called back stating the ps not reliable alarm is no longer present, but the suction alarm has returned. on ic the ps is 160/100. advised the ps is not accurate and giving a false suction alarm. disabled audio for suction. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13580 | 1220648-2024-24223 | 20783814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20783814&pc=OZD | 08/18/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/18/2021 | 11/26/2024 |  |  | 05/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022088321 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. it was noted that the placement signal was unreliable while prepping. another impella 5.0 pump was used. no patient harm was reported. | [5 empty] | The investigation for the reported placement signal issues has been completed. the impella device has not been returned for evaluation. the root cause of the placement signal unreliable alarm was unable to be determined as no product was returned for evaluation and no data logs were provided for analysis. this report is being filed as part of a retrospective review of historical records.
13581 | 1220648-2024-24218 | 20784101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784101&pc=OZD | 10/24/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 10/24/2021 | 11/26/2024 |  |  | 06/30/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1494055 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the wire used to put the pump over was an asaio 0.18 because the impella wire got kinked. it was reported that after pump placed and running patient was cardioverted three times unsuccessfully. he later went into ventricular tachycardia and was cardioverted one more time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13582 | 1220648-2024-24206 | 20784273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784273&pc=PYX | 11/11/2024 | 12/30/2024 |  | 11/26/2024 |  |  | 11/11/2024 | 11/26/2024 |  |  | 07/31/2025 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024388953 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp insertion was aborted due to the patient having an inferior vena cava filter. this information was undisclosed prior to surgery. there was no harm to the patient. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient was having a ivc filter.
13583 | 1220648-2024-24217 | 20784362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784362&pc=OZD | 08/02/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/02/2021 | 11/26/2024 |  |  | 05/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022087117 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. while on support, flows dropped quickly to less than 0.5 l/min, clot suspected. the physician elected to remove the pump, run in basin of sterile saline, and reinsert. the pump restarted, and the flows were appropriate for p level. perforation requiring covered stent and/or balloon tamponade. the patient expired. it is more likely that the obtuse marginal 1 perforation was the cause of the patient's death. however, the immediate low flow at start of support that resulted in explant and re-implant of the impella catheter cannot be excluded as a contributing factor leading to the complication. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13584 | 1220648-2024-24229 | 20784526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784526&pc=OZD | 05/23/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 05/23/2021 | 11/26/2024 |  |  | 12/31/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1457952 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | Failure to Sense |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, systemic heparin was paused due to bleeding complications, placement signal issues, high purge pressure alarm, and positioning issues was noted. the device was removed and a new impella 5.0 was implanted due to the loss of signals. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13585 | 1220648-2024-24213 | 20784661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784661&pc=OZD | 07/21/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/21/2021 | 11/26/2024 |  |  | 09/30/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1449375 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella 5.0 device for mechanical circulatory support. impella 5.0 was inserted via artificial blood vessel. the blood vessel diameter of the delivery was 8mm by prior evaluation. impella 5.0 was inserted but it was difficult to insert due to resistance around the branch when inserting. insertion was performed three times by using a buddy wire, but there was a shaft kink on the third try. the pump was replaced because there was a concern that it would lead to a pump failure. no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported delivery issues has been completed. the catheter shaft kinked upon insertion due to slight vascular stenosis. the root cause of the delivery issues was due to the patient's vascular stenosis. this report is being filed as part of a retrospective review of historical records.
13586 | 1220648-2024-24224 | 20784925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784925&pc=OZD | 03/30/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 03/30/2022 | 11/26/2024 |  |  | 03/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1471782 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that during axillary cutdown, the cardiac surgeon noticed the axillary artery was too calcific and borderline small for insertion. impella insertion was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13587 | 1220648-2024-24234 | 20784984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20784984&pc=OZD | 11/13/2024 | 01/03/2025 |  | 11/26/2024 |  |  | 11/13/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565539 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and just after the implant, the patient accidentally pulled the impella significantly into the aorta. before a repositioning attempt was made the doctor discussed goals of care with the spouse who decided to withdraw care and transition to comfort measures. the patient expired after care was withdrawn. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was use related as the patient accidentally pulled the impella catheter significantly into the aorta.
13588 | 1220648-2024-24212 | 20785543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20785543&pc=OZD | 07/19/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/20/2021 | 11/26/2024 |  |  | 04/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022082497 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.0 device for mechanical circulatory support. impella 5.0 was attempted from below the right clavicle, but the bypass anastomosis at the time of a dissection operation was narrowed which made it difficult to insert. the shaft kinked. impella 5.0 was explanted and an impella cp was used instead. no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported delivery issue has been completed. based on the clinical account, the root cause of the shaft kinking was due to a narrow anastomosis. this report is being filed as part of a retrospective review of historical records.
13589 | 1220648-2024-24211 | 20786161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786161&pc=OZD | 07/12/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/13/2021 | 11/26/2024 |  |  | 02/28/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1412378 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient developed an infection at the impella access site. treatment details were not provided. device was explanted and moved to left ventricular assist device (lvad). patient survived the procedure. | [5 empty] | The investigation for the reported infection/sepsis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the infection/sepsis issue was patient condition. this report is being filed as part of a retrospective review of historical records.
13590 | 1220648-2024-24220 | 20786260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786260&pc=OZD | 01/18/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 01/18/2022 | 11/26/2024 |  |  | 01/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1467005 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. however, the device was unable to pass through the patient's vasculature. insertion was aborted. however, it was reported that the patient survived the procedure. | [5 empty] | Device manufacture date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13591 | 1220648-2024-24226 | 20786332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786332&pc=OZD | 03/31/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 03/31/2022 | 11/26/2024 |  |  | 05/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1484859 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was extremely oozy upon opening of chest. patient taken off bypass and impella flows increased. the patient continued to bleed and required 2 units of packed red blood cells, 2 units of platelets, 2 cryp, and kcentra. bleeding eventually slowed and chest closed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13592 | 1220648-2024-24207 | 20786366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786366&pc=OZD | 11/11/2024 | 01/10/2025 |  | 11/26/2024 |  |  | 11/11/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525653 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support and during support, the pump had slipped into the aorta. it was unclear what made the positioning issues occur. the pump was explanted and the patient was stable. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was patient movement.
13593 | 1220648-2024-24231 | 20786527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786527&pc=OZD | 08/26/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/26/2021 | 11/26/2024 |  |  | 04/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022083132 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported a failed impella insertion with the device. a new impella was used without further issue. there was no adverse impact as a result. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. therefore, the root cause of the delivery issue was unable to be determined. this report is being filed as part of a retrospective review of historical records.
13594 | 1220648-2024-24210 | 20786727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786727&pc=OZD | 07/08/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/09/2021 | 11/26/2024 |  |  | 07/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021019638 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. impella did not pass when inserted via subclavian. therefore, the insertion was switched to the femoral which was also unsuccessful. impella 5.0 insertion was eventually aborted and impella cp was used instead. it was noted that the blood vessel diameter was evaluated preoperatively by echocardiogram, but it did not pass when it was inserted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was likely due to the patient's small vasculature. this report is being filed as part of a retrospective review of historical records.
13595 | 1220648-2024-24227 | 20786949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786949&pc=OZD | 05/28/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 05/28/2022 | 11/26/2024 |  |  | 05/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1478768 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after impella placement, it was noted that the axillary artery appeared small. the medical staff attempted to place the device however, it was reported that the pump could not pass the innominate artery. it was stated that the impella pump was explanted, and an intra-aortic balloon pump was implanted. the procedural outcome was the implant process was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13596 | 1220648-2024-24219 | 20786967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20786967&pc=OZD | 11/13/2024 | 01/15/2025 |  | 11/26/2024 |  |  | 11/13/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563493 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed access site bleeding during impella 5.5 support. blood products were given to the patient and manual pressure was applied to treat the condition. support continued following the intervention. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: general patient care considerations assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2024-24219 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
13597 | 1220648-2024-24230 | 20787106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787106&pc=OZD | 08/27/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/27/2021 | 11/26/2024 |  |  | 12/31/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1463041 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. the impella 5.0 was inserted via the subclavian area. there was push back when inserting and a kink was noted on the sheath. there may have been arteriosclerosis on the central side as it seemed to be slightly narrowed by echocardiogram. an impella cp was opened and inserted instead. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issues was determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
13598 | 1220648-2024-24221 | 20787343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787343&pc=OZD | 08/04/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 08/04/2021 | 11/26/2024 |  |  | 05/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022088776 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported placement signal issues occurring during use of this device. the surgeon did note there was a little resistance past the anastomosis while inserting the impella. after removing the peel away sheath and advancing the repo sheath, placement was checked again and adjusted accordingly. no patient harm was reported. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the placement signal issue could not be determined. this report is being filed as part of a retrospective review of historical records.
13599 | 1220648-2024-24235 | 20787356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787356&pc=OZD | 11/13/2024 | 01/06/2025 |  | 11/26/2024 |  |  | 11/13/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512454 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the patient had a hematoma at the access site. pressure was held, but the impella cp was removed and replaced with an impella 5.5. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
13600 | 1220648-2024-24267 | 20787519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787519&pc=OZD | 07/06/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 07/04/2022 | 11/26/2024 |  |  | 07/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022110409 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced bleeding and a hematoma, a blood transfusion was administered. it was reported that the patient survived normal explant. | [5 empty] | The investigation into the access site bleed issue has been completed. the impella 5.0 was returned for investigation. the case data logs were not returned. the returned pump was run in the laboratory for two hours and sensor drift was reproduced. the cause of the issue was sensor drift. this report is being filed as part of a retrospective review of historical records.
13601 | 1220648-2024-24214 | 20787570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787570&pc=OZD | 07/28/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 07/28/2021 | 11/26/2024 |  |  | 04/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075951 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. there was a failure to deliver the device due to patient anatomy and stenosis. an impella cp was used without issue. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was patient condition. this report is being filed as part of a retrospective review of historical records.
13602 | 1220648-2024-24233 | 20787644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787644&pc=OZD | 09/30/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 09/30/2021 | 11/26/2024 |  |  | 01/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022049639 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the impella 5.0 pump was impossible to advance beyond the subclavian, so an impella cp pump was used. there was no report of patient harm or medical intervention. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the returned pump was visually inspected and no abnormalities were observed. the cause of the delivery issue was stenotic patient vessels. this report is being filed as part of a retrospective review of historical records.
13603 | 1220648-2024-24240 | 20787652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787652&pc=OZD | 11/01/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 11/01/2021 | 11/26/2024 |  |  | 06/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022103908 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. pump had sensor issues after 138 hours of support. impella 5.0 was explanted and another pump was placed without issues. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. it was noted that the sensor was recording approximately 50mmhg in both air and water. biomaterial was found on the sensor, therefore, the sensor was cleaned. post cleaning, the sensor still had a reading of approximately 50mmhg. the sensor was zeroed and continued to read 50mmhg. a cotton swap was used on the sensor to ensure that it was responsive, the pressure read out increased as more pressure was applied to the cotton swap proving that the sensor was functional. the pump was zeroed for a second time and the resulting placement signal was close to zero as expected for a functioning sensor. the placement signal issue was most likely use related as the loss of flow due to increased placement signal occurred after the pump was manually zeroed and the event was repeated on a second 5.0. the placement signal issue was most likely use related as the loss of flow due to increased placement signal occurred after the pump was manually zeroed and the event was repeated on a second 5.0. this report is being filed as part of a retrospective review of historical records.
13604 | 1220648-2024-24258 | 20787748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787748&pc=OZD | 04/12/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 04/12/2022 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489284 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to an aortic valvuloplasty. while on support it was noted that heparin was switched to bicarb in the purge due to case bleeding and hematoma. the impella was successfully weaned and explanted. | [5 empty] | D4: expiration date and unique identifier (udi)# are unknown. h4: device manufacture date are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13605 | 1220648-2024-24250 | 20787836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20787836&pc=OZD | 11/19/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 11/19/2021 | 11/26/2024 |  |  | 01/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022051477 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported a failure to pass the device due to patient's anatomy. a new impella was used. there were no adverse effects endured as a result of attempted delivery of this device. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. based on the clinical account, the root cause of the delivery issue was due to anatomy. this report is being filed as part of a retrospective review of historical records.
13606 | 1220648-2024-24243 | 20788331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20788331&pc=OZD | 01/14/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 01/14/2022 | 11/26/2024 |  |  | 06/30/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120598 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. while on support, it was reported that the peel away sheath was cracked and peeled. the patient then had labs drawn and the labs came back hemolyzed. after 11 days the patient and family have decided on no additional treatment. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13607 | 1220648-2024-24253 | 20788365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20788365&pc=OZD | 04/12/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 04/12/2021 | 11/26/2024 |  |  | 12/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032762 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. while on support, it was reported that the impella was not fitting the groin. the impella was stuck in central position (bifurcation) therefore the case was aborted. the patient expired in procedural area. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13608 | 1220648-2024-24242 | 20788368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20788368&pc=OZD | 11/14/2024 | 01/03/2025 |  | 11/26/2024 |  |  | 11/14/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502441 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced left ventricle perforation after the implant. an effusion was noted and an emergency pericardiocentesis was performed and blood was drained and auto transfused back into the femoral vein. according to the physician, it was not caused by the impella or the impella wire. the patient had tokotsubo and suspects the left ventricle had a spontaneous perforation just prior to insertion. there is not enough information to exclude the impella device as an associated factor. | [5 empty] | The investigation for the ventricle perforation has been completed. clinical mentions that the patient had takotsubo which weakens the walls of the heart. as per the doctors observation the perforation was not caused by impella but before the insertion. the root cause of the ventricle perforation issue was most likely patient condition since the patient had takotsubo and the injury was unknown/unrelated to impella since the spontaneous perforation occurred prior to impella insertion. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13609 | 1220648-2024-24244 | 20788463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20788463&pc=OZD | 03/21/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 03/21/2022 | 11/26/2024 |  |  | 01/31/2024 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023175336 |  | 005066 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. it was reported that the purge pressure rose sharply. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13610 | 1220648-2024-24208 | 20788628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20788628&pc=OZD | 11/11/2024 | 01/03/2025 |  | 11/26/2024 |  |  | 11/11/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547838 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the nightshift nurse noted purge leaking from the purge sidearm of impella 5.5. tegaderm was placed over the crack. there was a crack at the bottom edge of the sidearm closest to the filter. the plan is to leave pump as is and watch purge closely. no purge alarms have been present. the nurse stated that during purge cassette change, multiple nurses could not get the yellow-to-yellow connection loose, so they used a hemostat. the leak was noticed after this. the nurse also noted, patient¿s wife was cleaning the sidearm with baby wipes. the patient is stable and waiting for a heart transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. the pump and purge cassette were returned for investigation. the cannula was cut on the returned product. no physical damage was observed on the returned product. the pump was investigated for purge leak by blocking the purge line at the catheter side. no leak was observed on the returned product. data log shows no sign of purge leak. as per the given clinical details, a purge leak was observed from the sidearm, and a crack was observed on the bottom edge of the sidearm. the rn did not observe any purge-related alarms during the case run. additionally, the rn used hemostat while making the connection to the yellow-to-yellow luer connection, and this is when the leak was noticed. the cause of the purge leak issue was most likely use related since no leak was reproduce on returned product and no purge leak trend was observed on data log.
13611 | 1220648-2024-24264 | 20789011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789011&pc=OZD | 06/17/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 06/17/2022 | 11/26/2024 |  |  | 03/31/2024 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023188177 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support. the complainant reported a failure of the pump to start. a new pump was used without any reported harm to the patient. no further information was provided. | [5 empty] | The investigation into the pump did not start issue has been completed. the impella 5.0 was returned for investigation. the returned impella 5.0 was tested and ran without issue. the root cause of the pump stop was unable to be determined as no data logs and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
13612 | 1220648-2024-24263 | 20789079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789079&pc=OZD | 09/06/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 09/06/2022 | 11/26/2024 |  |  | 10/31/2022 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022026587 |  | 005082 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the patient had oozing at the groin site. the patient received one unit of packed red blood cells and fresh frozen plasma. after 62.28 hours of impella 5.0 support, the family opted to have care withdrawn with the patient expiring, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13613 | 1220648-2024-24245 | 20789413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789413&pc=OZD | 01/11/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 01/11/2022 | 11/26/2024 |  |  | 04/30/2021 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1412264 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. insertion of the impella was unsuccessful due to patient vascular. the procedure was aborted, and intervention was completed without impella support. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13614 | 1220648-2024-24246 | 20789464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789464&pc=OZD | 11/14/2024 | 01/09/2025 |  | 11/26/2024 |  |  | 11/14/2024 | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557126 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the impella sheath was exchanged to reposition the sheath, however the impella was occluding the femoral artery. impella was occluding flow down the leg. the physician did not want to put distal perfusion catheter therefore the impella was weaned and removed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13615 | 1220648-2024-24271 | 20789685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789685&pc=OZD | 12/20/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 12/20/2022 | 11/26/2024 |  |  | 07/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022114015 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the cannula was damaged during insertion. a new impella pump was inserted successfully. no patient harm was reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the product was returned, and the issue was confirmed. kinking was found along the length of the cannula. no other damage or abnormalities were found. per the results of the clinical review and investigation, the root cause was determined to be due to excessive force. this report is being filed as part of a retrospective review of historical records.
13616 | 1220648-2024-24257 | 20789722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789722&pc=OZD | 12/03/2021 | 11/26/2024 |  | 11/26/2024 |  |  | 12/03/2021 | 11/26/2024 |  |  | 03/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022113977 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella ld for mechanical circulatory support. it was reported that the patient did not have doppler pulses in bilateral lower extremities. also noted high purge pressure alarm. alarms persisted despite tissue plasminogen activator (tpa) infusions via the purge. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13617 | 1220648-2024-24237 | 20789774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20789774&pc=OZD | 05/06/2022 | 11/26/2024 |  | 11/26/2024 |  |  | 05/06/2022 | 11/26/2024 |  |  | 12/31/2023 |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023162486 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Device Sensing Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the position waveform had stopped appearing, and an alarm of "there is no signal for position detection" displayed. additional information noted the patient had hemolysis. no intervention was mentioned. the loss of sensor/signal did not cause any harm or injury. cardiac function was restored. | [5 empty] | The investigation for the reported hemolysis and optical signal issues has been completed. the product was returned, and the optical signal issue was confirmed; however, hemolysis was not confirmed. the cause of the optical signal issue was a damaged optical fiber line due to excessive force on the catheter. the reported hemolysis issue was not confirmed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13618 | MW5162930 | 20797691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20797691&pc=OZD | 03/25/2024 | 11/20/2024 |  | 11/26/2024 |  |  |  | 11/26/2024 |  |  |  |  | 11/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | PUMP HEART IMPELLA 5.5 W/SMART ASSIST S2 |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. | [8 empty] | No | [9 empty] | Voluntary report | Not Available | Unknown | [6 empty] | 70 YR |  | Female | 70 KG | Non Hispanic | White | Injury | Yes | Yes | [7 empty] | Obstruction/Occlusion | Reduced Blood Flow | [11 empty] | Hospitalization | Life Threatening | Disability | Other |  |  | ATORVASTATIN 20 MG TABLET QD | ELIQUIS 2.5 MG TABLET BID | FARXIGA 10 MG QD | IMPLANTED DEFIBRILLATOR | LOSARTAN 12.5 MG TABLET QD | MAGNESIUM | METOPROLOL SUCCINATE XL 25 MG 24 HR TABLET | MULTI-VITAMIN B-COMPLEX | SPIRONOLACTONE 25 MG TABLET QD | [14 empty] | Initial submission | [9 empty] | While she was on ecmo(extracorporeal membrane oxygenation) she had a femoral impella that caused loss of blood flow to right leg that was irreparable and caused amputation above the knee in order to save her life from rhabdomyolysis. test peaked on (b)(6) 2024 at 39.103. her amputation was the evening of (b)(6) 2024. her next value was on (b)(6) 2024 at 6.669.
13619 | 1220648-2024-24164 | 20769010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20769010&pc=OZD | 01/27/2022 | 10/20/2025 |  | 11/25/2024 |  |  | 01/27/2022 | 11/25/2024 |  |  | 06/30/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED. INC. |  |  | US |  | 2022103100 |  | 004680-AU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella for mechanical circulatory support. during support, the patient experienced limb ischemia. the resolution for this issue is unknown. it was reported that the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13620 | 1220648-2024-24165 | 20772145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772145&pc=OZD | 05/18/2021 | 11/25/2024 |  | 11/25/2024 |  |  | 05/18/2021 | 11/25/2024 |  |  | 12/31/2022 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022044149 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemolysis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella cp support that the patient's laboratory samples were hemolyzed. the impella position was verified however there was no information provided on if hemolysis was resolved or any interventions performed. it was also noted that the medical staff requested abiomed support staff come onsite. they stated that the patient was dying and required an increased hemodynamic support. the impella 2.5 was already operating at the maximum flow level, and advised the patient would need to be escalated to a larger pump. it was noted that the patient later expired while on support. the relationship between impella and patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13621 | 1220648-2024-24166 | 20772353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772353&pc=OZD | 06/21/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 06/21/2022 | 11/25/2024 |  |  | 07/31/2022 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1496053 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported after the successful wean and explant of the impella 2.5, a covered stent was placed to achieve hemostasis at the impella access site due to damage caused by the perclose device that was deployed at the start of the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13622 | 1220648-2024-24170 | 20772508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772508&pc=OZD | 03/19/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 03/19/2022 | 11/25/2024 |  |  | 01/31/2024 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023175658 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/19/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure. the physician decided to explant the device and replace with a new impella 5.0. additional information noted that the patient died during support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13623 | 1220648-2024-24173 | 20772865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772865&pc=OZD | 08/09/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 08/09/2022 | 11/25/2024 |  |  | 09/30/2022 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2021013263 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. it was reported that the patient became coagulopathic and experienced profuse bleeding. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13624 | 1220648-2024-24167 | 20772938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772938&pc=OZD | 10/17/2024 | 12/18/2024 |  | 11/25/2024 |  |  | 10/30/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531998 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Renal Failure | Thrombocytopenia | [10 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. during procedure, .018 wire kinked during pushing. while on support, the patient was resuscitated with blood products due to bleeding into chest tubes. patient also had a thrombectomy of lower left leg (lle). the patient developed coagulopathy requiring massive transfusion of 4 units of packed red blood cells, 2 units of platelets, 2 units cyro, 2 units of fresh frozen plasma. coagulopathy improved. platelets continued to downtrend with the impella. patient received 1 unit of platelets. care was withdrawn. patient expired. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the thrombocytopenia, limb ischemia has been completed. no product was returned for analysis. clinical details mention significant difficulty passing 5.5 on .018" wire due to both anastomosis and angle of aorta to aortic root from right axillary. the root cause of the delivery issues was most likely patient condition related as evidenced by clinical details. the root cause of the thrombocytopenia and limb ischemia could not be determined without additional clinical details. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. corrections to the initial mfr report: b1-should have had product problem selected in addition to adverse event h6 medical dev prob code added 2682 h6 clinical code 2041 entered in error.
13625 | 1220648-2024-24188 | 20772943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20772943&pc=OZD | 10/15/2024 | 01/03/2025 |  | 11/25/2024 |  |  | 11/04/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534800 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had renal failure and continuous renal replacement therapy (crrt) was started. it is unknown if the impella contributed or caused the renal failure. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding at impella insertion site with a definite relationship to the impella device used: patient developed a hematoma several hours postoperatively at the lmpella insertion site requiring open decompression at the bedside with placement of a durable wound vac. she has continued to bleed from the wound vac site and critical care concerns, left us to return the patient to the operating room for re-exploration of the right infraclavicular incision. she got two units of packed red blood cells (prbc) on 10/17/2024. had oozing from femoral access and got two units of prbc 10/20. it was reported that the patient/event has not recovered/not resolved. the patient was on impella support for 532.70 hours before the patient expired after the family decided to withdraw care. | [5 empty] | The investigation for the renal failure and access site bleed has been completed. data logs showed no motor current (mc) spike or flow drop observed outside of p-levels change. no placement signal trend was observed. too many suction alarms occurred during the case. many impella position unknown alarms occurred. the cause of the access site bleed and renal failure could not be determined since the complaint device not returned for analysis and lack of clinical details. h.6 health effect - impact codes 1802 and 4617 were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24188 and were added. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13626 | 1220648-2024-24182 | 20773499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20773499&pc=OZD | 10/13/2022 | 10/09/2025 |  | 11/25/2024 |  |  | 10/13/2022 | 11/25/2024 |  |  | 07/31/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022114202 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced sepsis and was administered antibiotics. the patient died due to septic shock. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that during impella placement, the patient developed sepsis. it was stated that the issue was treated with antiobiotics in an effort to manage the complication. it was stated that the patient expired on support due to septic shock. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13627 | 1220648-2024-24181 | 20774028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20774028&pc=OZD | 08/12/2021 | 11/25/2024 |  | 11/25/2024 |  |  | 08/12/2021 | 11/25/2024 |  |  | 08/31/2021 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1444654 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited low flow, the resolution for this issue is unknown. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13628 | 1220648-2024-24175 | 20774039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20774039&pc=OZD | 08/13/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 08/13/2022 | 11/25/2024 |  |  | 11/30/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023167631 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombus and a stroke. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13629 | 1220648-2024-24168 | 20774675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20774675&pc=OZD | 01/31/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 01/31/2022 | 11/25/2024 |  |  | 04/30/2025 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022077227 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced heavy blood loss and 3 hours later the patient died. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13630 | 1220648-2024-24187 | 20775139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20775139&pc=OZD | 11/02/2024 | 12/30/2024 |  | 11/25/2024 |  |  | 11/03/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565649 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient developed hemolysis. the physician performed echocardiogram and repositioned, but ultimately decided to remove impella. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the hemolysis could not be determined. g.1 revised reporting contact email as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23970.
13631 | 1220648-2024-24179 | 20775169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20775169&pc=OZD | 10/05/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 10/05/2022 | 11/25/2024 |  |  | 08/31/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136283 |  | 005062-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited an air in the purge system error, the resolution for this issue was unknown. additionally, the patient experienced bleeding at multiple sites including a nosebleed that was treated with packing and heparin was withheld. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13632 | 1220648-2024-24180 | 20775173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20775173&pc=OZD | 10/06/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 10/06/2022 | 11/25/2024 |  |  | 01/31/2024 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023202538 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device migrated into the aorta, the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13633 | 1220648-2024-24184 | 20775885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20775885&pc=OZD | 05/28/2021 | 11/25/2024 |  | 11/25/2024 |  |  | 05/28/2021 | 11/25/2024 |  |  | 03/31/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022060123 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited a purge leak in the cassette, the resolution for this issue is unknown. additionally, there were concerns about bleeding from the chest tube and heparin was withheld. it was reported that the patient survived normal explant and the procedure. | [5 empty] | He impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13634 | 1220648-2024-24169 | 20776097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776097&pc=OZD | 10/31/2024 | 11/25/2024 |  | 11/25/2024 |  |  | 10/31/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544906 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had non-sustained ventricular tachycardias requiring no shock broke with lidocaine push. magnesium and potassium hemolyzed on last draw. there were no other labs hemolyzed and urine output is yellow. later, the patient had ongoing arrhythmias requiring defibrillation. additionally, the patient had gastrointestinal bleed. patient was not on anticoagulants and received two packed red blood cells and one fresh frozen plasma. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13635 | 1220648-2024-24178 | 20776421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776421&pc=OZD | 09/26/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 09/26/2022 | 11/25/2024 |  |  | 11/30/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161133 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the physician put in temporary pacemaker due to complete heart block. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13636 | 1220648-2024-24183 | 20776451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776451&pc=OZD | 05/27/2021 | 11/25/2024 |  | 11/25/2024 |  |  | 05/27/2021 | 11/25/2024 |  |  | 04/30/2022 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1476519 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombocytopenia and was administered sodium bicarbonate. additionally, the patient experienced a retroperitoneal bleed and heparin was discontinued. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13637 | 1220648-2024-24186 | 20776475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776475&pc=OZD | 10/07/2024 | 12/24/2024 |  | 11/25/2024 |  |  | 11/01/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545153 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Coagulation Disorder | Ischemia | Vascular Dissection | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in acute myocardial infarction and cardiogenic shock with st elevation myocardial infarction and ventricular septal defect. the pump support was ongoing when the abiomed¿s long-term outcome and quality indicator database reported upon adverse events occurring and attributed to impella use. there was thrombus present in the impella limb- the pulses were lost and thrombectomy was done. additionally a vessel suture wad done to vessel with a 5.0 prolene suture. there was a "possible" coagulopathy which did prompt blood products. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and vascular damage could not be determined.
13638 | 1220648-2024-24174 | 20776545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776545&pc=OZD | 08/09/2024 | 12/30/2024 |  | 11/25/2024 |  |  | 10/31/2024 | 11/25/2024 |  |  |  |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442878 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 77 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella 5.5 registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿gross hemolysis noted due to "failing circuit." required transfusions for correction.¿ additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿crrt started. cr at admission = 1.5. peak cr at 2.83.¿ it was reported that the patient/event has not recovered/not resolved. it was further reported that care was withdraw and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-24174. d6b explant date. g2 report source; study missing.
13639 | 1220648-2024-24171 | 20776623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20776623&pc=OZD | 03/20/2022 | 11/25/2024 |  | 11/25/2024 |  |  | 03/20/2022 | 11/25/2024 |  |  | 07/21/2023 |  | 11/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022111093 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Vascular Problem | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/20/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced limb ischemia and this led to malperfusion in legs. the physician decided to explant the device. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13640 | 1220648-2024-24130 | 20767257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767257&pc=OZD | 11/15/2024 | 12/19/2024 |  | 11/24/2024 |  |  | 11/15/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786102 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) had a broken purge disc. this issue was not discovered during a clinical procedure. there was no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
13641 | 1220648-2024-24131 | 20767261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767261&pc=OZD | 11/14/2024 | 01/22/2025 |  | 11/24/2024 |  |  | 11/14/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570415 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, ischemia developed in the patient¿s leg. pulsatility was difficult to verify on both sides due to severe peripheral arterial occlusive disease (pavk). the patient's condition worsened, and it was noted that the patient expired from multiple organ failure. the relationship between the impella and the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device was returned for evaluation and no abnormalities were found. there was no sufficient clinical information provided. therefore, the root cause of the limb ischemia could not be determined. added d9 device return date.
13642 | 1220648-2024-24132 | 20767315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767315&pc=PYX | 11/12/2024 | 01/02/2025 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  | 12/31/2024 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024320073 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Defective Device | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following insertion of a 23 french introducer during impella rp implant, it was discovered that a piece of the tip of the introducer had broken off outside of the patient. the dilator was removed and a new introducer was placed without issue. roughly five hours after the pump was implanted, the patient arrived in the intensive care unit. at that point in time, suction issues began and the p level dropped. an x-ray confirmed that the pump had backed out of the right ventricle and was no longer supporting the patient. the pump was explanted bedside and discussions were made about alternate forms of support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue and introducer damage has been completed. the pump was not returned for investigation. a data log review was not required. as per the clinical details, the dilator was found broken on the tip side. the dilator was changed from a new impella rp kit, and the doctor successfully placed the impella rp. the pump was found in the right ventricle after the initial placement. the doctor pulled the pump back into the inferior vena cava and then safely removed it. the cause of the introducer damage issue was not determined since the product was not returned and sufficient clinical information was not provided. the cause of the positioning issue was most likely use related, based on given clinical details.
13643 | 1220648-2024-24133 | 20767375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767375&pc=OZD | 11/12/2024 | 12/11/2024 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  | 07/31/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442857 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the pump had suction alarms when running at p-2. the pump was repositioned multiple times and did not resolve. the medical team made a plan to explant the pump due to the low pump flow. the patient was successfully weaned and explanted. there was no harm to the patient reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | H6 health effect - impact code 4628 was missing on manufacturer device report 1220648-2024-24133.
13644 | 1220648-2024-24134 | 20767447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767447&pc=OZD | 11/12/2024 | 01/02/2025 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544903 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella 5.5 device for mechanical circulatory support. it was reported while attempting to implant the impella 5.5, several attempts were made to pass the impella 5.5 past the coiled innominate artery and all were unsuccessful. the impella 5.5 insertion was aborted and an impella cp was used instead. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had coiled innominate artery.
13645 | 1220648-2024-24135 | 20767523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767523&pc=PYX | 11/10/2024 | 01/07/2025 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  | 10/31/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024423966 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella rp and impella cp device for mechanical circulatory support. it was reported while on impella rp support, the patient experienced hemolysis that was confirmed by elevated plasma free hemoglobin levels. it was later reported the patient had an acute kidney injury that occurred from the hemolysis that occurred during impella rp support. the patient was placed on continuous renal replacement therapy as a result of the kidney injury. the physician stated that they were unsure if the hemolysis was due to the impella rp or the impella cp. it was also reported the next day that the impella rp became malpositioned. due to the malposition, the physician decided to explant the pump due to the patient not requiring right ventricular support. the patient was successfully weaned and the impella rp was explanted. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the hemolysis, renal failure, and positioning issue has been completed. the pump was not returned for investigation. according to the clinical notes, the patient was reported with hemolysis diagnosed by elevated plasma free hemoglobin (pghb). this is when the cp pump was placed for left side support. the doctor reported that the hemolysis was resolved after the rp removal, while the cp pump was still in use. the data log shows that the rp had placement signal trends and motor current trends observed during the hemolysis event. also, at the time of the hemolysis event, several suction alarms and the trending placement signal and motor current were reported on the cp pump. also, the patient was on and off crrt during the impella support. based on the data log review, the cause of the hemolysis issue was related to the patient condition. the cause of the renal failure issue was related to the patient's condition, as there is no connection between hemolysis and the device's performance. rp pump was malposition on (b)(6). the cause of the positioning issue was most likely use related.
13646 | 1220648-2024-24136 | 20767530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767530&pc=OZD | 11/12/2024 | 01/13/2025 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521263 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited lower than expected flows. the device was repositioned and volume was given, but the low flows persisted. with no indication of improvement, the pump was explanted the following morning. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. returned complaint cp pump was visually inspected. biomaterial observed in purge gap and a big chunk observed at inflow cage. no broken blade found. no cannula kink observed. no other abnormality observed. the rc of the low pump flow was biomaterial ingestion.
13647 | 1220648-2024-24137 | 20767561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767561&pc=OZD | 11/15/2024 | 01/22/2025 |  | 11/24/2024 |  |  | 11/15/2024 | 11/24/2024 |  |  | 03/31/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359588 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Death | Yes | Yes | Difficult to Remove | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that while attempting to separate the two sides of the peel-away introducer, it broke and split as it should but then left the perforations and needed to be cut apart. it was noted that the procedure was successful. additional information noted the patient was going to be made a do not resuscitate and care was withdrawn. | Additional information the patient hemorrhaging into chest cavity, volume very low and cannot be maintained. |  |  |  |  | B2: the date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation introducer damage ¿ mechanical and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention . section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patient with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b5 additional information that was inadvertently omitted from manufacturer device report 1220648-2024-24137. g1 reporting contact email revised on manufacturer device report 1220648-2024-24137 in accordance with updated procedures. h6 health effect - clinical code 1888 was inadvertently omitted from manufacturer device report 1220648-2024-24137. h10 instructions for use were inadvertently omitted from manufacturer device report 1220648-2024-24137.
13648 | 1220648-2024-24138 | 20767565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767565&pc=OZD | 05/10/2024 | 05/15/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494432 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  | MW5169721,MW5161552 | Yes | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient's grandson reported that there was a pump stop during the procedure and the impella was explanted, and a balloon pump was implanted to continue support. the grandson also reported that he reviewed the medical records, and the attending nurse recorded in their notes "thrombocytopenia likely related to impella". he also noted that his grandmother was sent back to surgery soon after the initial surgery, the reason for the additional surgery is unknown. the patient expired two days after the initial surgery. | Additionally, the medwatch report explained that the patient a leak around the impella in the ascending aorta. the patient was losing a lot of blood. the impella was removed. the patient continued to bleed. there was epicardial bleeding on the anterior surface of the heart and there was also bleeding form the proximal anastomosis vein graft to the aorta. the bleeding was controlled and the patient was sent back to the icu. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the thrombocytopenia and pump stop has been completed. the pump was not returned. the root cause of the thrombocytopenia cannot be determined due to insufficient clinical information provided. the data logs were observed and there was an interaction with the impeller as there were suction alarms with an increase in motor current and pump flow on (b)(6) at 14:00. the pump stopped an hour later during cabg. the pump stop alarm that was seen was did not indicate an electrical failure of the device but rather a mechanical interaction with the impeller. the root cause of the pump stop cannot be determined as the product was not returned to confirm the interaction. corrections to the initial mfr report: g1 mdr reporting contact email. | H9 medwatch report 5169721 added. medwatch report 5169721 added. thrombocytopenia. the root cause of the thrombocytopenia cannot be determined due to insufficient clinical information provided. pump stop. the pump was not returned. the data logs were observed and there was an interaction with the impeller as there were suction alarms with an increase in motor current and pump flow on 5/10 at 14:00. the pump stopped an hour later during cabg. the pump stop alarm that was seen was did not indicate an electrical failure of the device but rather a mechanical interaction with the impeller. the root cause of the pump stop cannot be determined as the product was not returned to confirm the interaction. vascular damage - great vessel. due to the limited clinical information provided about what was occurring during this time frame, the root cause of the vascular damage - great vessel cannot be determined. b5 updated with additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24138. e4 should have been yes on the initial manufacturer device report number 1220648-2024-24138. h6 health effect - clinical code 1888 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24138. h6 health effect - impact code 4629 was inadvertently duplicated on the initial manufacturer device report number 1220648-2024-24138. h9 mw5161552 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24138. mw5161552 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24138.
13649 | 1220648-2024-24139 | 20767566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767566&pc=OZD | 11/11/2024 | 12/23/2024 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Limb Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in with acute myocardial infarction / cardiogenic shock and complex medical history, inclusive of guillain-barre syndrome, coronary artery disease, and hypertension. the pump was placed to allow for percutaneous coronary intervention of the coronaries. the pump was placed but there were harms associated, the patient suffered limb ischemia from vessel occlusion that was ballooned and nerve pain/numbness. the pump was explanted after just over 1 hour. patient awaiting coronary artery bypass graft. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation for the limb ischemia and peripheral nerve damage has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and peripheral nerve damage could not be determined as the device was not returned nor sufficient clinical details.
13650 | 1220648-2024-24140 | 20767577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767577&pc=OZD | 11/11/2024 | 01/02/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492228 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. following insertion, it was documented that the pump would not advance out of the subclavian into the arch. attempts to force the pump forward resulted in the device kinking and the decision was made to abort the implant. an impella cp pump was later placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
13651 | 1220648-2024-24141 | 20767598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767598&pc=OZD | 11/11/2024 | 02/11/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515051 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella 5.5 for mechanical circulatory support. it was reported during impella 5.5 insertion, the tip of the impella would not advance out of the subclavian into the aortic arch. the physician reported there was resistance. multiple attempts were made but were not successful. the impella 5.5 delivery was aborted after the catheter kinked and was no longer able to be pushed. the surgeon reattempted with a new impella 5.5. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the delivery issue and product damage has been completed. the product damage (cannula kink) was as a result of delivery issues. no product was returned for analysis. clinical details mention pump was unable to pass through vasculature. there was resistance at innominate artery and cannula kink was noted after physician pushed uncomfortably against resistance. no vessel conditions were noted but reason for aborted delivery was noted to be due to vascular anatomy. the root cause of the delivery issue was most likely due to patient condition.
13652 | 1220648-2024-24142 | 20767604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767604&pc=OZD | 11/11/2024 | 01/03/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546065 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, upon getting the patient ready to go to icu, the physician attempted to reposition impella out from under papillary muscle and dislodged the pump out of the atrioventricular valve. the physician attempted multiple techniques to get back across the valve, but unsuccessful. as a result, the impella was weaned and explanted. the patient was stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
13653 | 1220648-2024-24144 | 20767747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767747&pc=OZD | 11/11/2024 | 01/09/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  | 01/31/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315396 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that the pump could not pass through the peel away sheath and the implant was aborted due to the patient's anatomy. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. | The investigation of delivery issues and difficulty inserting/removing pump through introducer has been completed. no pump was returned. the pump was unable to pass the vasculature through introducer due to tortuous vessel condition and aborted insertion. the cause of the delivery issue was patient condition related due to tortuous vessel condition.
13654 | 1220648-2024-24145 | 20767751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767751&pc=OZD | 11/11/2024 | 01/06/2025 |  | 11/24/2024 |  |  | 10/28/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519150 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support there was a large hematoma noted at the access site. the impella was weaned and removed to gain hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was reported incorrectly on manufacturer device report 1220648-2024-24145.
13655 | 1220648-2024-24146 | 20767752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767752&pc=OZD | 11/11/2024 | 12/30/2024 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530652 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had an episode of supraventricular tachycardia. the p level was increased from p7 to p9, and pressors were stopped. additionally, the patient was presented with a septic workup. it was reported that the patient had additional episodes of ventricular tachycardia and required defibrillation. the patient suffered another ventricular tachycardia arrest and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of vf/vt/af/at and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13656 | 1220648-2024-24147 | 20767764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767764&pc=OZD | 11/08/2024 | 01/31/2025 |  | 11/24/2024 |  |  | 11/08/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1868829 |  | 0042-4201 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Display or Visual Feedback Problem | Application Program Problem | [5 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while powering on the automated impella controller (aic), the screen started flickering. aic booted up and appeared to be working so the physician proceeded with the implant. once the pump was connected and ran in auto mode, yellow alarm ¿controller failure¿ alarm was displayed. a pulsatile motor current was still noted as well as impella flows. the patient was actively coding, and the code was called before consoles could be swapped. once the code was called, the nurse called stating they couldn¿t get the console to power off. the console was eventually turned off and the patient expired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) and aic - display issue has been completed. according to device and data analysis the root cause of the controller error was due to an aic software issue. this was entered into jama with defect/ bug. the root cause for the flickering screen was not determined due to the inability to reproduce. d2 common device name was revised on manufacturer device report 1220648-2024-24147 in accordance with updated procedures. h5 labeled for single use was inadvertently reported incorrectly on manufacturer device report 1220648-2024-24147. h8 usage of device was inadvertently reported incorrectly on manufacturer device report 1220648-2024-24147.
13657 | 1220648-2024-24008 | 20767872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767872&pc=OZD | 11/13/2024 | 01/06/2025 |  | 11/24/2024 |  |  | 11/13/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561402 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ulcer | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and escalated care due to poor condition and underlying comorbidities. extracorporeal membrane oxygenation was also placed. the pump was placed but there was a gastric bleed that prompted the purge to be changed to sodium bicarb. the team treated with multiple units of blood and a surgical intervention. |  |  |  |  | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. corrections to the initial mfr report: b2-selected death and added date of death. b5-mso review. added d6b. g1 mdr reporting contact email. h1-changed to death. h6 impact code added 1802. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13658 | 1220648-2024-24143 | 20767874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767874&pc=OZD | 11/11/2024 | 01/02/2025 |  | 11/24/2024 |  |  | 11/11/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558973 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during an automated impella controller (aic) to aic transfer, the pump was not recognized when plugged into the second aic. after multiple failed attempts to troubleshoot the issue, the pump was explanted. there was no patient harm resulting from the failure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the pump not detected has been completed. data logs were reviewed and the complication was confirmed. the pump was disconnected from console while the pump was still running, which explains the trigger of alarm 115. at this time, the console sends an updated usage file to the eeprom, but there is no way to confirm whether the eeprom received and stored the file at this time. they then shut down the console, and start up a different one, but logs show them coming back to the first console before doing anything on any others. after restarting, they insert the purge cassette and get to case start step #3, when the logs show that the epm is reset and the message "connect cable: connect grey connectors" was displayed on the screen. case start is then exited. about 40 minutes later, a second console is started and the same exact sequence occurs, so they shut down the console. the only returned product was a section of the cut catheter/cannula. for this reason, there was no way to connect the pump to a console and check its functionality. there was biomaterial wrapped around the impeller, in the inflow cage, and blood ingress into the purge line, but again, since the catheter was cut there was no way to test if this was related to the failure. therefore, the root cause of the pump not detected could not be determined. added h6 impact code 4627.
13659 | 1220648-2024-24148 | 20767945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767945&pc=OZD | 11/11/2024 | 02/05/2025 |  | 11/24/2024 |  |  | 11/12/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533465 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced groan bleeding. this required surgical intervention, a suture was applied around the repositioning sheath. additionally, two units of blood products were given to the patient. the patient continued on impella support and ultimately survived. | Additional information was received that while the patient was on support, a pump stop occurred. the pump was restarted and support was continued despite the staff advising to replace the pump. furthermore, the pump eventually stopped again and it was unable to be restarted. the impella cp was pulled back and the patient was stable. |  |  |  |  | D6b: the date of explant is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of pump stop and access site bleeding has been completed. the pump was returned and evaluated. no logs were provided. no abnormalities were found with the returned repositioning sheath. based on the clinical details, the patient had bad, calcified vessels. the cause of the bleeding was most likely patient condition related due to reported calcified vessels. a large purge gap indicating bearing wear was found on the returned pump. no other abnormalities were found. the root cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp indication for use per design trace matrix. | B.1 additional information was received. therefore, adverse event or product problem was updated to reflect both. b.5 additional information was received and added. d.6b explant date was added. d.9 updated as the pump was returned. h.6 new codes were added due to new information that was received. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ e.1 facility name and city were incorrect on initial manufacturer device report number 1220648-2024-24148 and were updated.
13660 | 1220648-2024-24149 | 20768012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768012&pc=OZD | 10/30/2024 | 01/21/2025 |  | 11/24/2024 |  |  | 10/30/2024 | 11/24/2024 |  |  | 04/30/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347699 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Ischemia | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a code stroke (mca occlusion) overnight which was not treated. it was also reported the patient experienced red urine and laboratory samples were reported as hemolyzed. the patient also had no pulses in feet, legs cold, the right leg r leg has no flow, rigor has set in. the patient's family ultimately decided to withdraw care, and the patient expired. the relationship between the patient's outcome and impella use is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ | The investigation of hemolysis, limb ischemia, and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13661 | 1220648-2024-24150 | 20768024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768024&pc=OZD | 05/12/2022 | 11/24/2024 |  | 11/24/2024 |  |  | 05/12/2022 | 11/24/2024 |  |  | 08/31/2022 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503737 |  | 005062 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombocytopenia, and the staff continued with sodium bicarb in the purge system to help resolve this issue. it was reported that the patient survived normal explant and procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13662 | 1220648-2024-24151 | 20768430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768430&pc=OZD | 11/03/2024 | 01/06/2025 |  | 11/24/2024 |  |  | 11/14/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566677 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | Cardiogenic Shock | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported via notification from abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured cardiogenic shock during the procedure, which ended with a demised outcome, and left forearm compartment syndrome following the procedure both with a "possible" relationship to the impella cp device used. six days prior, the patient underwent percutaneous coronary intervention to the left anterior descending artery. three days later, the patient was readmitted and underwent ballooning to the right coronary artery. three days thereafter the last intervention, now day of impella implant, the patient was readmitted through the emergency room as a st-elevation myocardial infarction and taken to the catheterization lab for emergent placement of impella cp device and pacer. percutaneous coronary intervention was performed on the left anterior descending artery. during the procedure patient coded, underwent cardiopulmonary resuscitation and received pressors. at one point during the procedure discussion was made regarding adding extracorporeal membrane oxygenation (ecmo) or impella rp support or continuing the course. it was noted the patient was not an ecmo candidate and the interventional cardiologist decided not to implant an impella rp at such said time. pertaining to the left forearm compartment syndrome, the patient experienced left forearm swelling/turgor/induration following multiple arterial access attempts during code earlier in the day following angioplasty/thrombectomy of recently placed lad stent, an impella device was placed, and the patient was subsequently therapeutically anticoagulated. the patient was noted to have progressive edema and turgor of the left forearm, extending to above the antecubical fossa, and consult was called. the patient underwent left forearm compartment release, left carpal tunnel release, and dorsal forearm wound vacuum was placed. the patient recovered from this adverse event with no sequelae. continuing and while on impella cp therapy, lactic and creatinine continued to climb as well as patient having an increased requirement of inotropic support. the decision was made to escalate the patient to an impella 5.5 device which was therefore completed three days after start of support on the impella cp. the patient remained on impella 5.5 report for approximately five days before care was withdrawn. the patient was reintubated the following day after the start of support impella 5.5 implant, and the next day possible pneumonia was noted given positive sputum cultures. palliative care consulted, care was withdrawn and the patient expired. the patient was very ill going into the impella mechanical circulatory support and noted not a candidate for ecmo nor left ventricular assist device supports. however, given the cardiogenic shock event during the procedure with the impella cp device, it cannot be said this event did not contribute to the patient¿s demise even though after the patient was escalated to the impella 5.5. further, the registry noted the cardiogenic shock event with death outcome to be possibly related to the impella cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of vf/vt/af/at, nerve compression, and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-24151.
13663 | 1220648-2024-24152 | 20768458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768458&pc=OZD | 10/24/2024 | 12/23/2024 |  | 11/24/2024 |  |  | 10/24/2024 | 11/24/2024 |  |  | 11/30/2025 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2024422144 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient was taken to the operating room for a ventricular septal defect repair (this issue was identified prior to impella placement). the heart was continuing to beat, and it was noted the cross clamp was not providing enough seal. the physician questioned the placement of the pump, where they were able to palpate, what they believe to be the motor of the impella device. they decided not to move the impella up out of the ventricle because it would not be appropriate for a good seal or cross clamp. the decision was made to push the catheter into the heart enough to cross-clamp and deliver the cardioplegia. that did work and the surgeon was reminded of the blunt tip and noted where it was within the ventricle. the ventricle was perforated. once the left ventriculotomy was completed, it was noted that the impella was in the trebulae and the heart was empty. the procedure was completed without any issue. when going to place the graft, it was noted the impella was underneath the heart, and the pump was manipulated into the heart and patched. post-procedure, the physician noted that the patient's aorta was short, and the only option was to either remove the impella or advance it further (advancing further was the chosen method). it was reported that the procedure was successful. support continued with the implanted pump following the intervention. | [5 empty] | D6b: the date of device explant is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the ventricle perforation has been completed. no product was returned. as per clinical information patient had vsd before impella use and patient had short aorta anatomy. patient anatomical issues lead to likely ventricle perforation during support. the cause of the ventricle perforation issue was most likely patient condition due to patient anatomy issues per clinical information. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13664 | 1220648-2024-24153 | 20768624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768624&pc=OZD | 11/15/2024 | 12/26/2024 |  | 11/24/2024 |  |  | 11/19/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2025488324 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the pump was weaned and explanted, the patient experienced a vasovagal response from a femoral stop that had been applied to the impella access site closure. when removal of the femoral stop was attempted, the groin site began to bleed, and the femoral stop was reapplied. the patient had another vasovagal episode, requiring a fluid bolus and the administration of atropine. the systems reportedly resolved, and the femoral stop was removed with no report of recurrent bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer/; ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the vascular damage has been completed. no product was returned for analysis. the clinical details do not mention any excess force or relevant patient condition issue. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided.
13665 | 1220648-2024-24154 | 20768761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768761&pc=OZD | 08/22/2024 | 03/07/2025 |  | 11/24/2024 |  |  | 08/22/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2025513974 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. it was reported that during explant, the physician experienced resistance removing the pump and the impella snapped away from the catheter. the physician was able to retrieve the pump through the graph with hemostats. it was reported that the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the cannula/pigtail detachment has been completed. the pump was not returned for investigation. according to the provided clinical information, the pump cannula was detached during the removal of the pump, and the doctor applied excessive force during the removal. no image of the damaged pump was provided. the cause of the issue product damage (detached inflow/outflow - peripheral vessel) was not determined since the product was not returned and sufficient clinical information was not provided.
13666 | 1220648-2024-24155 | 20768778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768778&pc=OZD | 11/04/2024 | 01/02/2025 |  | 11/24/2024 |  |  | 11/05/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2025562943 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. during support, the patient went into ventricular tachycardia requiring cardiopulmonary resuscitation. the physician decided to preventatively reposition the device. it was reported that the patient was hemodynamically stable and survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. a.1 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24155. b.3 date of event was revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24155. e.1 revised first and last name as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24155. h.6 health effect - impact code 4617 was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-24155.
13667 | 1220648-2024-24156 | 20768779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768779&pc=OZD | 11/06/2024 | 01/13/2025 |  | 11/24/2024 |  |  | 11/06/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561326 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and while on support developed bleeding at the access site, device was explanted and the patient expired. it was noted that the femostop system was placed in the wrong area resulting in the bleed. the positioning of the femostop was corrected and two units of packed red blood cells (prbcs) were given to the patient to resolve the bleed event. two days later, it was noted there was no signs of recovery, the physician, therefore, explanted the device, and patient expired one hour later. additional information noted the patient was showing signs of multi-organ system failure and considered for the decision to explant the device. however, the bleeding event that required two units of prbcs cannot be excluded from contributing to the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related. a wrong placed femostop system which caused a medium bleeding at the groin site was notas per the clinical description provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.6 code 4627 was inadvertently omitted from initial manufacturer device report number 1220648-2024-24156 and was added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24156 and was added.
13668 | 1220648-2024-24157 | 20768828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768828&pc=OZD | 08/04/2024 | 02/19/2025 |  | 11/24/2024 |  |  | 08/04/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523891 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced small bleeding from groin impella access site. the physician placed a purse string suture to resolve the bleeding. it was also noted that a distal perfusion catheter was placed bilaterally in the groins for flow past the impella. the patient was withdrawn from care and expired. the relationship between impella use and the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. the root cause of access site bleeding is determined to be use issue because the bleeding was stopped by applying the purse string sutures are the bleeding site.
13669 | 1220648-2024-24158 | 20768838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768838&pc=OZD | 09/07/2024 | 02/14/2025 |  | 11/24/2024 |  |  | 09/07/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2025521508 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp for mechanical circulatory support. it was reported that during the placement of the impella cp that placement in the left ventricle was not possible due to the patient's anatomy. the procedure was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition as the insertion was aborted due to patient anatomy. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ warnings & cautions: cautions. never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
13670 | 1220648-2024-24159 | 20768870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768870&pc=OZD | 09/24/2024 | 02/10/2025 |  | 11/24/2024 |  |  | 09/24/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0045 |  |  | ABIOMED. INC. |  |  | US |  | 2025520993 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding at the access site, integrilin was administered, and heparin was paused. the physician added a suture and that resolved the bleeding. additionally, it was reported that the patient experienced runs of ventricular tachycardia and was administered lidocaine. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and ventricular tachycardia (vt) has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since adding additional suture resolved the bleeding. the root cause of the vt could not be determined without additional details.
13671 | 1220648-2024-24160 | 20768905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768905&pc=OZD | 10/26/2024 | 09/18/2025 |  | 11/24/2024 |  |  | 11/07/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced intermittent episodes of sustained ventricular tachycardia. on (b)(6) 2024 the patient experienced blood in stool and esophagogastroduodenoscopy was performed to reveal duodenal ulcers actively bleeding. it is unknown the relationship of impella to the bleeding, however the patient was heparinized for impella use. on (b)(6) 2024 it was reported the patient had a code blue eent after a sudden increase in st elevation. cardiopulmonary resuscitation was performed for two minutes after the impella was placed on p-2. return of spontaneous circulation was achieved. impella placement was verified with stat echocardiography, and the pump was slowly elevated to p level 6. the patient received multiple pressors due to extreme hypotension, which may have been related to liver shock. the patient descended into multi organ failure requiring increasing chemical support to sustain adequate hemodynamics, also necessitating renal replacement therapy.the family then decided to withdraw care. care was withdrawn from the patient, and the patient expired. the relationship between impella cp and the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13672 | 1220648-2024-24161 | 20768908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20768908&pc=OZD | 08/17/2022 | 02/24/2025 |  | 11/24/2024 |  |  | 08/17/2022 | 11/24/2024 |  |  | 04/30/2024 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023212696 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support and experienced a hemorrhagic stroke with a demised outcome. the patient was escalated from a impella cp to the impella 5.5 for hemodynamic support. six days into support, the patient remained intubated and sedated but followed commands, and the patient was noted to have experienced a subarachnoid bleed. heparin was put back into purge after repeat head computed tomography scan. of note, unrelated, it was documented that the patient was being treated for a fungal infection in lungs. care was eventually withdrawn four days later, and the patient expired. no further information was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13673 | 1220648-2024-24162 | 20769001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20769001&pc=OZD | 03/30/2022 | 11/24/2024 |  | 11/24/2024 |  |  | 03/30/2022 | 11/24/2024 |  |  | 11/30/2023 |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166225 |  | 0550-0008 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support it was noted that there was bleeding at the access site that was being managed with wound vac and transfusions. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13674 | 1220648-2024-24163 | 20769006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20769006&pc=OZD | 11/03/2024 | 01/03/2025 |  | 11/24/2024 |  |  | 11/03/2024 | 11/24/2024 |  |  |  |  | 11/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519261 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and expired after explant of the device in the cardiac catheterization laboratory (ccl). the impella cp was placed using the left femoral artery access. after having a percutaneous coronary intervention to the right and left anterior descending arteries the patient was transferred to another hospital. the right coronary artery stent had occluded during that hospitalization. the patient was not intubated but requiring several inotropes and vasopressors. in the setting of rca complete occlusion, the right ventricle was akinetic and the sinoatrial and atrioventricular nodes were not conducting properly. a temporary venous pacer was placed. the patient was weaned although on several inotropes and vasopressors to p-2 at which time the physician decided to transfer to the ccl and explant the impella cp, which was completed without issues and patient remained hemodynamically stable. the patient had been on support for approximately eight days. post explant the access site was observed to be at the bifurcation of the femoral artery. there were no preclosure devices implanted at time of impella cp placement. minimal to no flow was appreciated past the access site at the bifurcation of the femoral artery. a mod hook was placed through the right femoral artery and an angiogram was taken of the left iliac artery revealing a relatively large clot burden. an indigo retrieval device was inserted into the left femoral artery and retrieval of the clot was obtained with repeat angiogram revealing good distal runoff. attention was then turned to the left femoral access sheath, which was sequentially downsized from 14fr, 12fr, 10fr and ultimately 9fr with manual pressure being held for 25 minutes after each sheath exchange. during the procedure the patient¿s blood pressure started to slowly trend downward despite escalation of inotropes and vasopressors during the case. blood loss was appreciated during the sheath exchanges and an istat hemoglobin revealed a single point drop in hemoglobin from 8.2 to ~7.2 which prompted a unit of packed red blood cells during the case. the patient was also observed to have rales and rhonchi upon respiration which prompted lasix 40mg iv x 1. despite maximum medical management the patient¿s hemodynamics continued to decline and despite pacing the patient¿s heart rate began to drop in to the 30¿s, respirations became agonal, and the patient became unresponsive. a code/rapid response was called nearly two hours after successful removal of the impella cp from the left femoral artery and the patient was intubated and advanced cardiovascular life support (acls) was begun. despite three rounds of acls, return of spontaneous circulation was unable to be obtained and the patient was pronounced deceased. it was noted, in closing, the physician commented that there were not concerns or issues with the impella cp device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13675 | 1220648-2024-24010 | 20766022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766022&pc=OZD | 11/13/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551193 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed oozing from the impella access site. heparin was paused. in addition, the patient had multiple episodes of vomiting blood/clots, with blood in the patients stool and received a blood transfusion. the patient was also taken for a ct scan due to stroke-like symptoms; however, the stroke was not confirmed. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient condition because of the bleeding from multiple locations along with the impella access site. the root cause of the stroke could not be determined without additional information.
13676 | 1220648-2024-24015 | 20766042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766042&pc=OZD | 11/14/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/14/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564761 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was provided which states staff was unable to doppler bilateral pedal pulses. known peripheral vascular disease (pvd). no details were provided on the treatment of the loss of pedal pulses. however, it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the hemostatic valve separated from the introducer while removing the dilator. the physician stated, ¿it was during an emergency and i might have been pulling hard.¿ the sheath was exchanged for a 13cm. there was no reported harm or injury and the patient remained on support. |  |  |  |  | The investigation for the introducer damage and limb ischemia has been completed. there was no product returned. clinical details were sufficient to determine root cause. based on the clinical details provided, the root cause of the introducer damage -mechanical is most likely due to use. the root cause of the limb ischemia could not be determined without additional information. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added the following information: b1 selected adverse event. b2-selected death and date of death. b5-mso review. d6b-explant date. h6 clinical code 1942. h6 impact code 1802 corrections to the initial mfr report. g1 mdr reporting contact email. h1-updated to death. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13677 | 1220648-2024-24017 | 20766064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766064&pc=OZD | 11/14/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/14/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549916 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support due to cardiomyopathy. while on support for 2 weeks, the pump was seen to have high purge pressure. the staff treated this by administering sodium bicarbonate and starting the tissue plasminogen activator procedure. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ g1 reporting contact email revised on manufacturer device report 1220648-2024-24017 in accordance with updated procedures. h10 instructions for use were missing on manufacturer device report 1220648-2024-24017.
13678 | 1220648-2024-24020 | 20766342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766342&pc=OZD | 11/03/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2025553665 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an attempted implant of an impella cp device for mechanical circulatory support to an unknown age patient with cardiogenic shock from an acute myocardial infarction was unsuccessful. details of the situation and fluoroscopic images were confirmed, and after inserting the impella cp into the left ventricle, the stiff/floppy switch part of the placement guidewire was found to be stuck near the discharge hole. although attempts were made to remove the placement guidewire by pushing and pulling it several times, the shaft of the impella kinked, and the shaft twisted violently inside the blood vessel and made a full circle. at this time, arrhythmia was also occurring. the impella cp and placement guidewire were quickly pushed and pulled to quickly remove the kinks and twists and lowered to the descending aorta. the placement guidewire came off the impella just above the sheath of the descending aorta, and the impella cp and placement guidewire were removed from the body. due to concerns about the effect of vascular tortuosity, a new pump was not used, and an intra-aortic balloon pump was used instead. the patient's outcome following the event is unknown. however, there was no report or indication of adverse effects to the patient following the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation has been completed. the impella was returned for evaluation. visual inspection found a kink on the guidewire (gw) coil and cannula about 15 millimeters from the outflow cage and cannula bonding site. no other issues were found on the rest of the pump. as the patient had vascular tortuosity, both cannula and the gw were kinked during the difficult placement resulted in difficult removal of gw. the root cause of delivery issue was most likely patient condition. as the necessary information was not provided, the root cause of the arrhythmia was not determined.
13679 | 1220648-2024-24021 | 20766345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766345&pc=OZD | 11/06/2024 | 12/26/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554064 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and right after the impella cp device started in auto mode, flows remained low at approximately 1 liter. the position was verified and the patient had adequate volume. it was noted that the team heard a humming sound coming from the pump with concern for a clot entrapment. the pump was explanted and replaced. there was no report of adverse effects to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation low pump flow has been completed. returned complaint pump was visually inspected a cannula kink near outlet observed. no broken blade found. no biomaterial observed. pump run for 10 minutes in investigation lab after connected to aic. no unusual sound noticed during impella run. data log reviewed: no positioning issue/ no mc spike observed. suctions occurred followed by drop in flows without changing the p-level. the root cause of the low pump flow issue was cannula kink due to improper technique.
13680 | 1220648-2024-24022 | 20766348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766348&pc=OZD | 11/06/2024 | 08/22/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) experienced a controller error alarm. the pump suddenly stopped and a controller failure alarm occurred. the aic was restarted and support was able to continue with the implanted pump. there was no patient harm. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental manufacturer device report will be filed. | Corrections to the initial mfr report: g1: mdr reporting contact email. d2: device name has been updated. h6: type of investigation code added.
13681 | 1220648-2024-24023 | 20766352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766352&pc=PYX | 11/06/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560957 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with impella rp flex device post implant of an impella cp for mechanical circulatory support. while on support, the patient experienced "significant" bleeding from the impella rp flex access site. during the procedure, the representative recommended placing a hemostatic suture prior to dilation up to the 23 french peel away. however, the physician opted to place stitch later in order to get support more quickly to the unstable patient. hemostatic suture was placed once repositioning sheath was advanced. however, bleeding continued at the insertion site for the duration of the night. manual pressure was held and dressing was changed multiple times. multiple anticoagulation modalities were administered in the catheterization lab resulting in elevated activated clotting time/partial thromboplastin time. however, the bleeding eventually began to slow down once coagulation normalized. the patient was reported to be stable following the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella rp flex with smartassist instructions for use & clinical reference manual states the following: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
13682 | 1220648-2024-24024 | 20766357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766357&pc=OZD | 11/06/2024 | 12/31/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524958 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was undergoing a protected high-risk percutaneous coronary intervention (hrpci) and during the impella cp insertion in the right femoral artery, the impella cp passed into the sheath. however, it was unable to cross the top of the iliac into the aorta. the physician attempted to go into the left femoral artery. the impella cp passed into the sheath and again was unable to cross the top of the iliac into the aorta. the impella cp insertion was aborted. the decision was made to proceed with the hrpci case unprotected. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
13683 | 1220648-2024-24025 | 20766363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766363&pc=OZD | 11/06/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563120 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported delivery issues with an impella 5.5. the impella could not be advanced due to patient anatomy. the decision was made to abort insertion of impella 5.5. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature.
13684 | 1220648-2024-24026 | 20766366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766366&pc=OZD | 11/06/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523836 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported difficulty was encountered during an attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance through the patient's tortuous aortoiliac and horizontal aorta anatomy. the aortic valve was found to be calcified, and the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13685 | 1220648-2024-24027 | 20766370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766370&pc=OZD | 11/06/2024 | 12/20/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556257 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device and experienced access site bleeding around the sheath. further information noted that there was continuous oozing from the impella sheath insertion site for the entire case and the physician attempted to reposition and re-angle the sheath without success. the physician believes that on insertion of the dilators, one of them created a slight tear to the arteriotomy due to the insertion angle. at the end of the case, the impella cp device was successfully explanted, and it was reported the groin looked "okay" with a small hematoma not requiring blood product. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. | The investigation for the vascular damage has been completed. no product was returned for analysis. the root cause of the vascular damage - peripheral vessel was not determined as no product was returned and limited clinical information was provided. corrections to the initial mfr report: g1 mdr reporting contact email.
13686 | 1220648-2024-24028 | 20766373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766373&pc=OZD | 11/13/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487620 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was needing an impella 5.5 device for mechanical circulatory support due to cardiomyopathy. the surgeon placed a jp drain at the surgical site due to continued oozing post impella 5.5 implant. it was reported that one liter of blood was lost overnight. the drain was removed and pressure was being held. the patient was given one unit of packed red blood cells. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13687 | 1220648-2024-24029 | 20766377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766377&pc=OZD | 11/12/2024 | 12/23/2024 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1701427 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was needing an impella 5.5 device for mechanical circulatory support. a rountine purge cassette changed was attempted. when the new purge cassette was inserted, the disc would not click into place. after several attempts, the cassette was placed. however, the automated impella controller (aic) screen was unable to advance from ¿insert new cassette then close door¿ screen. the aic was exchanged and issue was resolved. there was no reported harm or injury to the patient. | [5 empty] | The investigation automated imeplla controller (aic) - purge disc/flag issues has been completed. the data review showed there were numerous purge disc not detected alarms. the purge disc detection issue was unable to be reproduced. the purge disc flag and retainer were both intact and functioning as normal. after further examination of the aic, there were evident signs of ingress residue on the purge flag and all around the purge pressure sensor assembly. the root cause of this issue was likely contamination of the purge flag¿s optical sensor. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13688 | 1220648-2024-24030 | 20766379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766379&pc=OZD | 11/13/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547851 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp pump unexpectedly stopped. attempts to restart the pump were unsuccessful and the decision was made to remove the pump. there was no reported harm or injury to the patient. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon completion of our analysis, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the pump was returned for investigation. data log review shows the pump stopped at the end of the case run after 20 hours of operation, with observed alarm 115 - pump stop. the returned product has a catheter kink near the kink protector. the red handle was opened, and no open line was observed inside it. further investigation revealed the starched open line inside catheter at kink protector. the cause of the pump stop issue was an open electrical line inside the catheter near the kink protector.
13689 | 1220648-2024-24031 | 20766383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766383&pc=OZD | 11/13/2024 | 02/07/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541271 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that roughly one day into support, the impella cp pump unexpectedly stopped. the pump was attempted to be restarted. after the secone attempt to restart, a controller failure occurred and the decision was made to remove the pump. there was no reported harm or injury to the patient. the pump issue was reported in manufacturer device report number 1220648-2024-24030. this report reflects the automated impella controller issue. | [5 empty] | The console was returned, and an evaluation is underway. upon completion of our analysis, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console was tested. the reported failure modes were unable to be reproduced naturally (either pump stop or controller error). forcefully reproduce the controller error by unplugging the lamp connector from the optical bench. there was no issue with the envelope (without connecting the optical test pump) and the fringe pattern from the optical bench (with the optical test pump connected). this confirms no issue with the optical bench. according to the data and device analysis, the root cause for the intermittent ¿controller error¿ was most likely a defective lamp. d.2 revised common device name since the intial manufacturer device report number 1220648-2024-24031 was submitted in accordance with updated procedures.
13690 | 1220648-2024-24032 | 20766389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766389&pc=OZD | 11/13/2024 | 01/10/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1637336 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that during morning inspection, a control device failure occurred on the automated impella controller (aic). "self-diagnosis failed" message was displayed. the aic was re-started but the issue did not resolve. no patient involvement was reported. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console log showed a communication issue with the pbm (power battery manager) during the initial bootup, with error messages. the reported failure mode was reproduced during testing. visual inspection confirmed the pbm contamination. the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
13691 | 1220648-2024-24033 | 20766393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766393&pc=OZD | 11/12/2024 | 12/19/2024 |  | 11/23/2024 |  |  | 11/12/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544897 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Peripheral Nervous Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced a hematoma at the impella insertion site resulting in heparin being removed. it was reported that on day two of support, there was nerve pain in the right arm that prompted the team to treat the patient with gabapentin. the patient remained on impella support as a bridge to transplant and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the access site bleed and peripheral nerve injury has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and peripheral nerve injury could not be determined.
13692 | 1220648-2024-24034 | 20766405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766405&pc=OZD | 11/12/2024 | 12/23/2024 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) was unable to prime the impella. it was reported that a purge system failure would pop up when attempting to prime the pump. the purge cassettes were switched and the aic was turned off and on again but the issue was not resolved. this aic was replaced with another aic and the procedure was successfully completed. | [5 empty] | The investigation for the console (aic) purge issue has been completed. after reviewing the console logs, the reported purge system failure was confirmed. there were two instances of purge system failure (piston blocked) [event set #417] alarms observed in the imc logs from the reported complaint date. the console was returned for evaluation. the reported purge system failure issue was unable to be reproduced. the aic was able to prime and run with a pump without any issues. after inspecting the aic, ingress residue was observed on the purge piston and purge motor housing. the reported purge system failure issue was likely caused by contamination of the purge motor piston and purge motor housing which may have resulted in the purge motor piston getting stuck to the purge motor housing. if the purge motor piston can't turn, the aic is unable to prime the purge system. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13693 | 1220648-2024-24035 | 20766421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766421&pc=OZD | 11/12/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the procedure was completed and then the impella cp was successfully weaned and explanted without any issues. the day after the impella was removed, the nurse informed that post explant the patient complained of increased pain in the right leg. the patient had no pulse in their right lower extremity upon doppler. vascular surgeon was notified and a computed tomography angiography verified a thrombus occlusion of the right superficial femoral artery. the clot was surgically removed and a patch was placed to repair the artery. the patient was stable following the intervention. no further issues were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13694 | 1220648-2024-24036 | 20766442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766442&pc=OZD | 11/09/2024 | 01/10/2025 |  | 11/23/2024 |  |  | 11/09/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519663 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Malposition of Device | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the device needed to be repositioned several times. the physician stated that if he moved the impella by 5 millimeters, the device exhibited an impella in ventricle or impella in aorta alarm. it was reported that the values were in range and the alarm was deactivated. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13695 | 1220648-2024-24037 | 20766459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766459&pc=PYX | 11/09/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/09/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551222 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella rp flex for mechanical circulatory support. it was reported that shortly after initiation of impella flows, there was a pump motor stop which resolved after restarting the pump twice. the flows returned back to p7 without any further reported issues. it was reported that the patient survived the normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the root cause of the temporary pump stop issue was most likely biomaterial.
13696 | 1220648-2024-24038 | 20766480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766480&pc=OZD | 11/08/2024 | 06/18/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563693 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.5 for mechanical circulatory support. during support, the patient experienced runs of supraventricular tachycardia. the device was repositioned. it was reported that the patient survived the procedure. | It was reported that on (b)(6) 2024 the patient has had progressively lowering purge flow baseline flow was 7ml/h when purge flow dropped to 2.8ml/h tissue plasminogen activator (tpa) was put in the purge. |  |  |  |  | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | High pressure purge: pump was returned and investigated. chunk of biomaterial was found on the impeller and also the purge gap had traces of biomaterial. the high purge pressure could not be reproduced since the pump was returned cut from the catheter. data logs were investigated and a steady rise in purge pressure was observed from (b)(6) 2024 when tissue plasminogen activator (tpa) was added and purge flow was observed to be reducing. no motor spikes observed. therefore, as per the product investigation, data logs review and clinical information, the root cause of the high purge pressure issue was gradual biomaterial buildup with was resolved by adding tpa to the purge. the investigation into vt, there was no change to the previously reported investigation results. d9 device returned to manufacturer date added. instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24038. b5 additional information was inadvertently omitted on manufacturer device report number 1220648-2024-24038-1. h6 medical device problem code 1491 omitted on the initial manufacturer device report number 1220648-2024-24038. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24038.
13697 | 1220648-2024-24039 | 20766490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766490&pc=OZD | 11/08/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563868 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during delivery, the impella went deep into the patient's ventricle and kinked. the physician repositioned the device and the patient went into pulseless electrical activity (pea) and required cardiopulmonary resuscitation. the impella was moved during compressions and the pump migrated out of the left ventricle. the physician decided to explant the device. additional information noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B.2 the exact date of death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of the positioning issue was most likely improper technique supporting the catheter while advancing the repo unit since the pump went deep at the time of advancement.
13698 | 1220648-2024-24041 | 20766523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766523&pc=OZD | 11/08/2024 | 01/15/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559106 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced heavy bleeding from the groin impella access site. the blue fixation of the repo-sheath was inserted too deeply, resulting in vascular bleeding. no perforation or dissection noted. the position of the sheath was optimized, fem-stop was placed above the sheath and the surgeon placed a vascular suture. blood products were given. additionally on 13-nov-2024, a motor current high pump stop was reported. the pump ran on p1, when the p level was set higher, the pump would abruptly stop. the pump flow average was lower than expected at 2.5l/min. due to the pump stop issue, the pump was pulled back into the aorta, and then explanted. the patient was stable and remained on extracorporeal membrane oxygenation support. | [5 empty] | The investigation for the access site bleed and pump stop has been completed. mpella cp was returned for evaluation. upon visual inspection, no damage to pump was observed. dried blood was present in the cannula and mold was present along pump and on impeller. electrical testing was performed and all motor cable lines were intact. pump was attempted to be run in the lab at p-9 and had immediate pump stop. pump was attempted to be run again at p-5 and was able to start. dried blood and mold were kicked off the pump when running and pump was able to be turned up to p-9 and run without issue. no residual biomaterial was present on impeller after running. no data logs were returned for evaluation. clinical details noted that pump stop occurred and pump was able to be turned back on to p-1 and run before being explanted. the root cause of the pump stop cannot be definitively determined without data logs returned. clinical details noted that heavy bleeding from the groin area was observed on day 1 of support. it was found that the blue butterfly of the repositioning sheath was inserted too deeply in the access site. the position of the blue butterfly and sheath were optimized and the site was dressed with a pressure dressing. as the customer reported the access site bleeding against the 14fr introducer sheath but butterfly is a component of the repositioning sheath connected to the pump and the complaint description reported that the bleeding was occurring when the repositioning sheath was inserted. the root cause of the access site bleeding - major was most likely a use related issue as the blue butterfly was inserted into the skin. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
13699 | 1220648-2024-24040 | 20766524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766524&pc=OZD | 11/08/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed via the 14 french sheath but the team did not follow recommendations and peel away the sheath. the peel away was left in and the leg became ischemic. the team placed a perfusion catheter for the limb. vascular surgery was consulted. the team explanted after weaning the pump 66 hours. the patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13700 | 1220648-2024-24042 | 20766544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766544&pc=OZD | 11/08/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560892 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was going to be on an impella 5.5 for mechanical circulatory support. it was reported that during delivery, the 0.018 wire was advanced into the left ventricle. the impella 5.5 was opened and prepped for insertion. the impella 5.5 was advanced through the graft, but it was getting caught up near the anastomosis. the impella 5.5 was removed and the physician tried to implant the device multiple times with different wires, angles, and inflation of an 8 millimeter mustang peripheral balloon prior to insertion. the insertion was unsuccessful. the physician was unable to advance the device. support was aborted due to the patient's inadequate vessel size. the vessel was repaired and the site was closed. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella insertion aborted due to inadequate vessel size.
13701 | 1220648-2024-24043 | 20766572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766572&pc=OZD | 11/08/2024 | 12/31/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during delivery, the device recrossed the aortic valve. an echocardiogram was done because the physician was concerned about placement. the physician tried to reposition the device and the impella was pulled into the aorta. the physician decided to explant and replace the device. it was reported that the patient survived explant. no additional information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
13702 | 1220648-2024-24044 | 20766583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766583&pc=OZD | 10/03/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  | 06/30/2024 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | A st-segment elevation myocardial infarction door-to-unload study reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis which resolved by removing the impella cp. the procedure outcome was that the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-24044. b3 date of event, h1 type of reportable event.
13703 | 1220648-2024-24045 | 20766613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766613&pc=OZD | 10/31/2024 | 12/31/2024 |  | 11/23/2024 |  |  | 10/31/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554538 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a left ventricle thrombus and a bilateral cardiovascular accident confirmed by computed tomography. the physician stated that the left ventricle thrombus was a result of the left anterior descending coronary infarction which the patient was originally admitted for. due to the left ventricle thrombus and the stroke, the pump was explanted. the patient was successfully weaned and explanted. the patient remained on extracorporeal membrane oxygenation support. | [5 empty] | The investigation of thrombus and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13704 | 1220648-2024-24046 | 20766651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766651&pc=OZD | 05/02/2022 | 10/24/2025 |  | 11/23/2024 |  |  | 05/02/2022 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 2.5 for mechanical circulatory support. it was reported during impella support, the automated impella controller (aic) had a battery red alarm/vga issue. it was reported there was no red color on display. the issues with the aic prompted the aic to be removed from patient use. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
13705 | 1220648-2024-24047 | 20766654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766654&pc=OZD | 11/06/2024 | 02/24/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1752054 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Packaging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility returned the automated impella controller (aic) due to shipping damage/wet packaging. the aic was removed from the user facility and replaced. there was no patient involvement. | [5 empty] | E.1 initial reporter details are unknown. if this information becomes available, it will be processed and a supplemental manufacturer device report will be filed accordingly. the automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the packaging issues has been completed. the complaint could be neither confirmed nor denied. there are no logs to review for this complaint. the console was returned and functioned as expected during testing. the cause of the issue was not determined since console was investigated outside of packaging.
13706 | 1220648-2024-24048 | 20766678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766678&pc=OZD | 11/06/2024 | 12/27/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560158 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and experienced bleeding and a stroke. the patient exsanguinated a large amount of blood during procedure. the physician is unsure of where the issue occurred (peel away membrane or sheath itself). significant resistance and/or catching into vasculature was noted. an attempt was made to use short 14 french dilator on its own but then transitioned to serial dilations after resistance was met. the patient was given six units of packed red blood cells. the following day, hemoglobin and hematocrit had stabilized post the transfusion. additionally, it was noted, a computed tomography scan of the head revealed possible infratentorial hemorrhage versus left old infarct and neurology consult was made with a plan to hold systemic anticoagulation. assessment with sedation was paused. the patient was reported to be stable following the event. one day prior to planned explant of the impella cp, the patient was noted to be doing well. the patient was successfully weaned and the device was explanted without incident. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding and stroke/cva has been completed. data logs are not relevant to the complication. the 7 fr companion sheath was not returned for investigation. the returned 14 fr introducer was investigated and there were no visual abnormalities. additionally, it was put through a leak reproduction test and held pressure up to 275 mmhg without leaking. based on limited clinical details and no abnormalities with returned product, the root cause of the access site bleed could not be determined. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
13707 | 1220648-2024-24049 | 20766718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766718&pc=OZD | 11/06/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553004 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with persistent ventricular tachycardia (vt) with repeated implantable cardioverter defibrillator shocks went for vt ablation and impella cp device implanted post ablation and experienced limb ischemia, bleeding and hematoma adverse events. the patient required multiple vasopressors to maintain adequate blood pressure. after impella cp was placed, side port injection was done, and flow was limited down the leg post impella sheath. the patient lost pulses at the limb and a femorofemoral bypass was completed. further, hematoma and bleeding developed at the impella site. the patient received two units of packed red blood cells. the patient continued on support and on the third day, critical care stated futility in escalation due to the patient¿s declining clinical status and multi organ failure. a discussion was made with the family with plans to transition towards comfort care. the patient expired the same day in the afternoon. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details.
13708 | 1220648-2024-24050 | 20766736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766736&pc=OZD | 11/06/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with impella devices for mechanical circulatory support. an impella cp device was placed in right groin prior to the percutaneous coronary intervention to the left anterior descending artery. however, the patient continued to decompensate after revascularization and the decision was made to implant the impella rp flex as well. the patients hemodynamics recovered post impella rp flex insertion and pressors were able to be weaned. the patient was transported to the unit on support. the patient initially was hemodynamically stabilized after impella rp flex insertion. however after arrival to the unit, the patient began to decompensate again. both pumps were running at p9 with no active alarms. partial thromboplastin time remained elevated. there was significant bleeding from both impella sites. femostop was applied at low pressure over the impella cp insertion site. both dressings were changed and sites were assessed. the bleeding eventually began to slow down once coagulation was normalized. the next day, the patient continued to decompensate and the decision was made to place extracorporeal membrane oxygenation and remove the impella rp flex. the impella cp was maintained for left ventricular decompression. approximately seven days later, the physician noted the patient decompensation was due to septic and the patient was vasoplegic. the impella cp was decreased due to patient having hemolysis requiring transfusion. the next day it was noted the impella cp was explanted with no further details. the status of the patient following the event is unknown. however, the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this is one of two medical device reports associated with the reported event. the associated impella rp flex device was reported under manufacturing report number 1220648-2024-24023. | The investigation of access site bleeding, hemolysis, and infection/sepsis has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely due to anticoagulation management related as the bleeding improved when activated clotting time was adjusted to normal range. the root cause of hemolysis and infection was unable to be determined as limited clinical details were provided.
13709 | 1220648-2024-24051 | 20766751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766751&pc=OZD | 11/08/2024 | 12/23/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511519 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. during support, it was reported that the patient was bleeding from the axillary cutdown site post procedure. it was discovered that this was a muscle bleed, and the staff transfused two units of fresh frozen plasma. it was reported that the patient survived the explant. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13710 | 1220648-2024-24052 | 20766763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766763&pc=OZD | 11/03/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544107 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. during support, the patient experienced ectopy, and it was reported that the device was malpositioned. the resolution to the issue is unknown. additionally, the patient also experienced supraventricular tachycardia requiring defibrillation twice and was administered amiodarone. it was reported that the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related. as the necessary information was not provided, the root cause of the vt/vf was not determined.
13711 | 1220648-2024-24053 | 20766778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766778&pc=PYX | 11/06/2024 | 11/23/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025546029 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was placed on impella cp and impella rp flex for mechanical circulatory support and during support the impella rp flex purge pressure rose. on unit same day of bipella implant, the impella rp flex purge pressure rose, and purge flow decreased. tissue plasma activator was added to the purge and the staff monitored. approximately four days later, after the patient was repositioned, the automated impella controller triggered an alarm for purge pressure high, then purge system blocked. per the nurse, the critical patient remained hemodynamically stable during this time. the team discussed concern for a possible pump stop and the decision was made to explant the impella rp flex device, which was successfully completed later this same day. it was noted the patient tolerated the explant without an issue and remained hemodynamically stable following the event. | [5 empty] | The investigation has been completed. the pump was inspected and there were no physical abnormalities. a large amount of biomaterial was found in the inlet cage and around the impeller as well as in the cannula. the data logs confirmed high purge pressure alarms with a decrease in purge flow. the data logs showed no motor current spikes before the high purge pressure event. there was a gradual rise in purge pressure for over an hour before alarms were triggered. the first high purge pressure alarm was triggered and there was high purge pressure for about five hours. the next day, there was a motor current spike after which the purge pressures started decreasing towards normal values with an increase in purge flow. clinical information reported tissue plaminogen was started the day of the first high purge pressure alarm and ended the next day. the data logs do not show any other high purge pressure occurrence for the remainder of the case. the root cause of high purge pressure was determined to be biomaterial.
13712 | 1220648-2024-24054 | 20766784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766784&pc=OZD | 11/06/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527484 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unsuccessful attempt to implant an impella cp device in a patient for mechanical circulatory support. the patient presented to the catheterization lab as st-elevation myocardial infarction. the radial access was gained and a left main dissection was noted with elevated end-diastolic pressure. the decision was made to implant an impella cp which was prepped accordingly. access was gained to right femoral artery and the 14 french sheath was inserted. the pigtail catheter was inserted and was unable to cross into the left ventricle. it became known that the left main dissection extended to the descending aorta. the physician aborted the impella cp implant and the patient was taken to the operating room. the patient's status following this event is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the vascular damage and the access site bleed has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause was changed from delivery issue to vascular damage since aortic dissection was reported. the root cause of aortic dissection is determined to be use issue because the dissection had resulted during the insertion as extended from lm to descending aorta where the resistance is being felt during the insertion. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature and unable to cross the lv.
13713 | 1220648-2024-24055 | 20766803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766803&pc=OZD | 11/08/2024 | 12/19/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding. the physician decided to explant the device. manual pressure was held and hemostasis was achieved with positive distal pulses. it was reported that the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13714 | 1220648-2024-24056 | 20766814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766814&pc=OZD | 11/08/2024 | 01/13/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537458 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella 5.5 for mechanical circulatory support. it was reported that the aorta was incised and the impella 5.5 was inserted under direct vision. guidewires and fluoroscopy were not used. the impella 5.5 was inserted after confirmation with echocardiography but due to ventricular collapse, the echocardiogram was difficult to see and considering the risk of perforation, the impella 5.5 was only inserted to a shallow position. after the aorta was sutured and the volume was filled in the left ventricle, it was discovered that the impella 5.5 had perforated the posterior wall of the left ventricle when adjusting its position. the procedure was arrested again. the perforated area was reinforced with a patch and it was confirmed that there was no bleeding. the impella 5.5 was restarted and it was confirmed that there were no problems. it was reported that the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the ventricle perforation has been completed. the pump was not returned for evaluation. based on the clinical details provided, the root cause of the ventricle perforation is most likely due to use as the perforation occurred when adjusting the position of the impella.
13715 | 1220648-2024-24057 | 20766873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766873&pc=OZD | 11/30/2023 | 01/02/2025 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  | 09/30/2025 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395651 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, mild oozing was noted at the right axillary impella insertion site. blood transfusion was required. oozing was noted to be improved days later. care was later withdrawn and the patient expired. | [5 empty] | The investigation of access site bleeding ¿ major has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event reported incorrectly on manufacturer device report 1220648-2024-24057. d6a and d6b revised from manufacturer device report 1220648-2024-24057 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-24057 in accordance with updated procedures. | B2: the date of patient death is unknown. the impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
13716 | 1220648-2024-24058 | 20766891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766891&pc=OZD | 11/04/2024 | 11/23/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485803 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella cp support. while on support, the physician was informed that the device sensor may not function. the physician elected to proceed with pump insertion via peel away sheath and started on support. aorta (ao) pressure reading was non-pulsatile and continued to read 35/35(35), pulsatile motor current and left ventricle waveform and flows 3.2. the decision was made to remove the pump and place new device. the patient survived the event. | [5 empty] | The impella device was returned and evaluated. the root cause of the optical signal issue was not determined as no data logs were returned for analysis.
13717 | 1220648-2024-24059 | 20766896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766896&pc=OZD | 11/04/2024 | 12/27/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568535 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 insertion, the 5.5 could not pass the innominate artery, despite multiple attempts. impella cp was used instead. the patient had a successful cp implant and moved to the intensive care unit. no further information is available. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. corrections to the initial mfr report: g1 mdr reporting contact email.
13718 | 1220648-2024-24060 | 20766897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766897&pc=OZD | 08/20/2024 | 01/09/2025 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527916 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp support initiated for the high-risk percutaneous coronary intervention patient in (b)(6) 2024. the support was successful for just over 2 hours and was then explanted. later in (b)(6) 2024 the protect iv study documentation noted a bleed with possible relationship to the impella. the patient did get transfused. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, due to prolonged pci procedure (cto) increased blood loss through catheter access site without access site complication as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email.
13719 | 1220648-2024-24061 | 20766901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766901&pc=OZD | 11/04/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the implant, the patient had oozing at the access site. the sutures were clinched to mitigate oozing. a hematoma formed and staff held pressure. doppler pluses were absent from distal extremities. a do-not-resuscitate was made for the patient and the patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and access site bleeding ¿ major has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13720 | 1220648-2024-24062 | 20766902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766902&pc=OZD | 11/04/2024 | 02/11/2025 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285933 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella purge cassette was not detected during implla 5.5 support. the cassette was exchanged without success. the automated impella controller (aic) was exchanged, and purge cassette was detected on the new aic. there was no harm to the patient. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the pump not detected issue has been completed. console logs from the reported day of event revealed that purge cassette was detected, but purge disk detection was intermittent, resulting in multiple ¿purge disk not detected¿ alarms, which presumably prompted the staff to repeatedly remove then reinstall the purge cassette and disk in attempts to resolve the issue. the console was returned for evaluation and inspected, and damage to the purge disk detect flag and disk retainer were observed. the cause of the issue was a defective component. d.9 device return date/information was added. a.3 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24062. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-24062 was submitted in accordance with updated procedures. h.6 revised health effect - impact code as an incorrect code was reported on the initial manufacturer device report number 1220648-2024-24062.
13721 | 1220648-2024-24063 | 20766903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766903&pc=OZD | 06/13/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through the protect iv study that the patient on impella cp support developed right groin hematoma and transfusion was given. additionally, there was a decrease in hemoglobin from baseline. the patient underwent complex percutaneous coronary intervention through right groin with impella placement for hemodynamic support. after the procedure it was noted a right groin hematoma. it was stable until impella removal on (b)(6) 2024 when it appeared to slightly increase in size; peripheral pulses remained palpable although some ecchymosis was noted. this event likely contributed to the decrease hemoglobin from baseline for which two transfusions were performed. hematoma was deemed to be resolved prior to discharge on (b)(6) 2024. patient baseline hemoglobin was already low. however, after index procedure it started to decrease reaching 7.6 g/dl on (b)(6) 2024 (requiring transfusion 1 unit on (b)(6) 2024 accomplished a hemoglobin of 9.2 g/dl) and 7.7 g/dl on (b)(6) 2024 also requiring transfusion (1 unit and reaching 8.8 g/dl). patient hemoglobin was noted to be 9.7 g/dl at discharge. on, (b)(6) 2024 hemoglobin was found to be 12.9 g/dl which is closer to the baseline value. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13722 | 1220648-2024-24064 | 20766907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766907&pc=OZD | 11/04/2024 | 12/27/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562265 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following the use of excessive force to disconnect the yellow luer from the purge sidearm, leaking was noted from the yellow check valve. both the purge cassette and purge solution were replaced, but the issue persisted. the impella 5.5 pump was subsequently removed in response to the purge leak. there was no patient harm. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. data logs were reviewed and the purge leak was confirmed on 11/4 by "purge system open" alarms. additionally, it can be seen that the team attempts to replace the purge cassette several times, but the low purge pressure persists. product was not returned for investigation, but clinical details report that excessive force was used when trying to disconnect the yellow luers during a purge cassette change. this caused a leak from the sidearm check valve. based on clinical details and data logs, the root cause of the purge leak was determined to most likely be a damaged sidearm due to excessive force.
13723 | 1220648-2024-24065 | 20766911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766911&pc=OZD | 11/04/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522103 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had bleeding from the access site. heparin was withheld and the bleeding slowed. the patient received three units of blood products. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, pt bleeding from all access sites as per the clinical description provided.
13724 | 1220648-2024-24066 | 20766912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766912&pc=OZD | 11/04/2024 | 12/20/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to acute-on-chronic congestive heart failure. while on support, the patient had no distal leg perfusion. the pump was explanted due to loss of pulses on lower extremity. the patient coded and expired. | [5 empty] | B2: the date of patient death is unknown. the impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13725 | 1220648-2024-24067 | 20766913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766913&pc=OZD | 11/04/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had dp pulse no longer able to be dopplered. the vascular study done revealed that there was no flow past the mid-calf. the impella cp was removed. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13726 | 1220648-2024-24068 | 20766915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766915&pc=OZD | 05/20/2024 | 11/23/2024 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492136 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported that after the impella cp explant, the patient was oozing blood from access site during routine check up. hemostasis was achieved, however, patient later got up to go to the bathroom and site began bleeding once again. femostop was applied and hemostasis was achieved. the patient did not receive any blood products. no serious patient harm or injury. on nov. 5th, 2024 the complaints team received a protect iv study that found a probable relationship between the patient's claudication to the impella procedure. there was claudication in right extremity at site of single sheath procedure, requiring stent placement for common femoral artery and superficial femoral artery lesion. impella procedure and closure device likely contributed. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13727 | 1220648-2024-24069 | 20766916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766916&pc=OZD | 10/31/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567099 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and it was aborted. the surgeon wanted to attempt 5.5 before using a cp. wires and pigtail passed easily but 5.5 would not advance. an impella cp was placed, and the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. b3 date of event incorrectly reported on manufacturer device report 1220648-2024-24069. g1 reporting contact email revised on manufacturer device report 1220648-2024-24069 in accordance with updated procedures.
13728 | 1220648-2024-24070 | 20766919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766919&pc=OZD | 11/05/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025536705 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed an infection of unknown origin during impella cp support. the infection resulted in circulatory failure and the patient ultimately passed away from multiple organ failure. it is unknown if the impella device caused or contributed to the infection that led to the patient's passing. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. added d4-model number updated corrections to the initial mfr report: g1 mdr reporting contact email updated.
13729 | 1220648-2024-24071 | 20766923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766923&pc=OZD | 11/01/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549021 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed an infection at the impella 5.5 artificial vascular graft during pump support. an intra-aortic balloon pump was introduced and the impella 5.5 pump was reintroduced from the contralateral side. the patient was given antibiotics to treat the condition. the patient survived the event. | It was further reported that due to the patient's cardiac function lack of recovery, the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. added: b2-selected death and added date of death. b5- mso review. h1-changed to death. h6 impact code added 1802 corrections to the initial mfr report: d6b has been updated. it is important to note the initial reported explant date was correct at the time of submission as the pump was explanted and reintroduced to the opposite side. the explant date has been updated now to reflect the final explant at the time of the patient expiration. g1 mdr reporting contact email
13730 | 1220648-2024-24072 | 20766924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766924&pc=OZD | 05/07/2024 | 12/24/2024 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  | 11/30/2025 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427737 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Unknown |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the device experienced flow and suction issues. the pump was repositioned, but the issue remained. the device console was then switched, which resolved the issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. no device was returned for evaluation. data logs were reviewed finding no motor current (mc) spike observed outside of p-levels change. drop in pump flow at p-9 observed on 5/7. many suctions occurred during the case and mostly at the date of the reported issue. no issue observed to confirm the rc of the reported low pump flow issue. the root cause of the low pump flow issue cannot be determined since the complaint device not returned for analysis and lack of details.
13731 | 1220648-2024-24073 | 20766926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766926&pc=OZD | 11/06/2024 | 11/23/2024 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313446 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a "system self-check failed" error on startup. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. data analysis: the console logs which showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service (fs). upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination. which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
13732 | 1220648-2024-24074 | 20766928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766928&pc=OZD | 11/07/2024 | 12/17/2024 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562046 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, a motor spike was noted, and flow equalization was noted at p7. there was a motor current upward trend and flow saturation. there was no reported resolution, and no reported patient harm. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | B2: the date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the saturated flow has been completed. there is no product returned. the data logs show that there are suction alarms and then there is a spike in motor current and placement signal, followed by saturated flows. during this time the purge pressure begins to rise. the root cause of the saturated flow is most likely due to biomaterial. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ corrections to the initial mfr report: b1-should have selected adverse event b2-should have selected death and added date of death b5 mso review d4-model number updated h1-changed to death h6 impact code 1802 added.
13733 | 1220648-2024-24075 | 20766930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766930&pc=OZD | 08/13/2024 | 01/09/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SMARTASSIST |  |  | OZD |  |  | IMPELLA CP SMARTASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2025482791 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was attempted to be implanted with an impella cp for mechanical circulatory support. during the implant attempt, the pump could not be advanced through the aortic valve, and the implant was aborted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the patient's aortic valve made insertion difficult as per the clinical description provided.
13734 | 1220648-2024-24076 | 20766939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766939&pc=OZD | 11/07/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562650 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown race/ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after multiple days of impella support, the patient continued to bleed at the access site. the patient was given packed red blood cells, and heparin levels were reduced. manual pressure and bandaging were utilized to achieve hemostasis. support continued with the implanted pump following the intervention. the patient ultimately expired, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13735 | 1220648-2024-24077 | 20766941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766941&pc=OZD | 11/07/2024 | 04/24/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  | 05/31/2025 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356748 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient had red urine. patient labs were unable to be obtained due to being hemolyzed. the physician ultimately determined that the patient was experiencing hemolysis due to impella. patients map maintained 115-130mmhg throughout the night. the device was explanted and patient is stable post explant. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data logs showed pump ran without issue for 8 hours then pump position was in aorta that triggered alarms impella position in ventricle/aorta. this event remained to the rest of the case. no other issues were found on the logs. pump was returned for analysis. visual inspection found no issues on the pump. the root cause of hemolysis was most likely due to pump was patient condition.
13736 | 1220648-2024-24078 | 20766942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766942&pc=OZD | 11/07/2024 | 08/25/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1383536 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 93 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Complainant in japan reported the automated impella controller (aic) was being prepped for an impella implant. when starting up the aic, there was a yellow alarm with controller error. the device was started three times, but the problem did not improve. the aic was replaced. no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13737 | 1220648-2024-24079 | 20766944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766944&pc=OZD | 11/07/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1209412 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us based console was having a preventive maintenance and found to have battery charge test failure. there is no patient involved | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. the console logs from the reported day of event did not contain any data relevant to this failure mode. during inspection of the console¿s components, a loose metal fragment (test point loop from one of pcbs) was found inside the housing, which most likely shorted either power supply or pbm, which in turn initiated failure of the remaining components. the cause of the aic issue was a defective power supply. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-24079. | The device was returned for evaluation, power battery manager board (pbm) was defective.
13738 | 1220648-2024-24080 | 20766947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766947&pc=OZD | 11/07/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the patient had an access site bleed. two units of blood products were given to the patient. the patient expired after 39.78 hours of impella support. the patient's cause of death was determined to be multiple organ failure, however the relationship between the access site bleeding and the cause of death is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13739 | 1220648-2024-24081 | 20766949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766949&pc=OZD | 11/07/2024 | 01/09/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544803 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient had an impella cp implant and during the implant, there was kinking in the guidewire. this caused a left ventricle perforation and a pericardial tamponade. pericardium was drained successfully; 750 mls were removed and 2 units of red blood cells were given. the coronary artery bypass grafting was completed and the patient was stable. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the ventricle perforation and product damage-guideline has been completed. abiomed 0.0180" guidewire was returned for evaluation. a kink in the guidewire near the distal tip of the wire with some coil unraveling present next to wire kink. clinical details noted that during wire insertion, a perforation of the ventricle occurred and resulted in pericardial tamponade. issue was resolved with draining of the pericardium and pump was able to be inserted and procedure completed without further issue. the root cause of the ventricle perforation was most likely due to a guidewire puncture since the perforation occurred during guidewire insertion. the root cause of the guidewire damage - great vessel cannot be definitively determined as limited clinical details were provided. added d9 device return date.
13740 | 1220648-2024-24082 | 20766951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766951&pc=OZD | 10/07/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1857507 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) was being assembled for service and it was noted that the rolling cart was broken. no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) cart issue has been completed. there are no logs necessary to analyze for review of a broken cart. the product did not need to be returned and a replacement was sent out to the customer.
13741 | 1220648-2024-24083 | 20766958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766958&pc=OZD | 10/14/2021 | 05/22/2025 |  | 11/23/2024 |  |  | 10/14/2021 | 11/23/2024 |  |  | 07/31/2023 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022108611 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella cp support, the patient experienced constant bleeding from the impella insertion site into the pectoralis major and upper arm. the patient was transfused with two units of packed red blood cells and one unit of plasma due to the blood loss. in was noted that additional surgery was performed to the site to resolve the bleed. it was additionally reported that the patient developed sepsis with an abdominal focus. the medical team decided the sepsis would be treated post impella explant. the patient remained in poor condition and was not a candidate for a left ventricular assist device or heart transplant. care was withdrawn, and the patient expired. the relationship between the sepsis and patient outcome to the impella is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13742 | 1220648-2024-24084 | 20766959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766959&pc=PYX | 09/07/2024 | 07/17/2025 |  | 11/23/2024 |  |  | 09/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED. INC. |  |  | US |  | 2025540864 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 2.5 for mechanical circulatory support. it was reported that the device exhibited low flow and was unable to go past p7 on support due to suction. it is unknown how this issue was resolved. additionally, the patient experienced ventricular tachycardia, the resolution for this issue is unknown. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Instructions for use: impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2024-24084 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2024-24084. g2 foreign was erroneously selected on the initial manufacturer device report number 1220648-2024-24084. h10 instructions for use were inadvertently omitted from manufacturer device report number 1220648-2024-24084. | Access site bleeding: the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. renal failure: the root cause of the renal failure was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. low pump flow: the clinical states that the rp was unable to go past p7 on support due to suction and due to the femoral placement, the inlet was located in the distal ivc and unable to offload the rv. the data logs show that there is suction throughout the case and a trend in the ps and flow. there are no positioning alarms seen. based on the clinical details and data logs, the root cause of the low pump flow is most likely due to the improper positioning as the inlet was located in the distal ivc and was unable to offload the rv. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
13743 | 1220648-2024-24085 | 20766962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766962&pc=OZD | 08/29/2021 | 11/23/2024 |  | 11/23/2024 |  |  | 08/29/2021 | 11/23/2024 |  |  | 01/31/2023 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED. INC. |  |  | US |  | 2022047424 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that the patient experienced a mild pericardial effusion, and the staff started the patient on nor-adrenaline which was tapered off in 12 hours. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13744 | 1220648-2024-24086 | 20766965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766965&pc=OZD | 10/29/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 10/29/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515221 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Wound Dehiscence | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with ischemic cardiomyopathy was implanted with and impella 5.5 device for mechanical circulatory support (mcs) and allowed for transfer to the tertiary center. a hematoma developed at the impella 5.5 access site; however, no intervention was required. approximately 30 days into support, the access site wound dehiscence occurred. the team reinforced blue suture hub with sutures to skin and a wound vacuum was possibly placed (unconfirmed). the following day it was noted the site covered with dressing, incision remained dehisced with no new plans related to both the impella 5.5 or the incision site. the patient was being planned for a mitral clip procedure the following week, and the impella 5.5 device would remain in place until then. approximately 10 days later, the patient was successfully weaned, and the device was explanted with vascular surgery in the operating room. the patient was noted to be stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of wound dehiscence has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d4 catalog number was missing on manufacturer device report 1220648-2024-24086. e4 should have been left blank on manufacturer device report 1220648-2024-24086.
13745 | 1220648-2024-24087 | 20766966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766966&pc=OZD | 11/07/2024 | 12/27/2024 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  | 01/31/2024 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486003 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp insertion, there was a perforation in the iliac artery. patient started to decline at this point. surgeon placed the wire in the true lumen and placed the peel away with dilator back into vessel. left ventricle was wired and impella placed without issue. patient started to improve after impella was on, peel away was left in per surgeon¿s recommendations. after the case the staff was unable to get doppler pulses in the right lower extremity. the reperfusion sheath was placed, and it was reported that the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of vascular damage - peripheral vessel and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13746 | 1220648-2024-24088 | 20766967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766967&pc=OZD | 11/07/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525263 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a large clot was attached to the catheter on the 9 fr portion in the descending aorta. this thrombus was noted during the removal of the pump. the pump was replaced with an impella 5.5. patient is stable post issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. e.1 revised us phone number as it was inadvertenly reported incorrectly on the initial manufacturer device report number 1220648-2024-24088.
13747 | 1220648-2024-24089 | 20766968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766968&pc=OZD | 11/07/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547052 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a cp pump placed to support a patient admitted in for a high-risk surgery. the pump was placed but was explanted after just 25 hours due to a leak at the red plug. the blood leakage occurred after there was low purge pressure during the support. the pump was explanted and replaced. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for product damage has been completed. the pump was returned for investigation. data log review was not required for this case run. the returned pump passed electrical testing and ran according to specifications during the test in a bucket of water. the red handle was open, and blood was observed inside. further investigation revealed that a cut was found on the catheter near the 20 cm mark, which allowed the blood to reach the red handle and leak from it. as per the clinical, blood leak was observed on the second day of use without any evidence of excessive force. the cause of product damage issue was most likely use related. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
13748 | 1220648-2024-24090 | 20766969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766969&pc=OZD | 11/10/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/10/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576531 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the right femoral artery was fully occluded by the impella repositioning sheath. a 6fr antegrade sheath was placed for distal limb perfusion and support with the implanted pump was maintained. additional information noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia could not be determined as the device was not returned nor sufficient clinical details.
13749 | 1220648-2024-24091 | 20766971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766971&pc=PYX | 11/11/2024 | 09/25/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502628 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient had access site bleeding. the patient was hemorrhaging at the access site. the patient required 4 units of packed red blood cells, 2 units of platelets, and an increase in vasopressors in order to maintain goal mean arterial blood pressures. a hemostatic suture was applied, and heparin was withheld. the staff stated that the bleeding was likely due to over-anticoagulation from systemic heparin. support was continued, no further information is available. | The complainant additionally reported that a placement signal not reliable alarm appeared during support two days later. there was no impact on support, and positioning was verified to be accurate. support continued. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major was unable to be definitively determined as several possible contributing factors were potential causes of bleeding. h6 health effect - impact code 4643 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-24091. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs confirm placement signal not reliable (psnr) alarm. there was a motor current spike, speed deviation, and immediate loss of central venous pressure. the dp sensor remained stable. the root cause of the optical issue was most likely a damaged sensor face. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number. b7 other relevant history, including preexisting medical conditions. d10 concomitant medical products and therapy dates.
13750 | 1220648-2024-24092 | 20766972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766972&pc=OZD | 11/08/2024 | 01/28/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being prepared for implant with an impella device for mechanical circulatory support. it was noted that the soft touch buttons on the automated impella controller (aic) were "stuck" and not responding. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. the first log entries indicating "imc-kbd error: 0xa1 0x00 0x" appeared, which suggests the issue occurs when the console screen is touched outside the soft touch button pads. after this, the case was paused, and the aic was not rebooted again. the console was tested upon arrival at field service. after performing 200 power cycles, the failure could not be reproduced. no damage was found to the keyboard or the keyboard- 4in1 cable. the root cause for the intermittent keyboard communication issue could not be determined due to the inability to reproduce. a.1 and a.3 information was added as it was inadvertently omitted from the initial manufacturer device report number. b.3 revised date of event as it was reported incorrectly on initial manufacturer device report number. d.2 revised common device name since the initial manufacturer device report number was submitted in accordance with updated procedures. d.4 revised catalog number as it was previously entered incorrectly on the initial manufacturer device report number. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number.
13751 | 1220648-2024-24093 | 20766975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766975&pc=OZD | 11/11/2024 | 12/24/2024 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530499 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed an infection at the impella access site. the patient was being treated with antibiotics. the patient was determined to not be a candidate for advanced therapies and was made dnr/ dni. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
13752 | 1220648-2024-24094 | 20766976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766976&pc=OZD | 11/11/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025549941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the patient was experiencing hemolysis. impella was too deep. impella was repositioned and hemolysis was resolved. suction alarms were present after pressures returned to baseline and was unable to get above p4. motor current was flat. impella was in good position. the decision was made to remove the pump and place a new. after removing pump, it appeared to be a piece of tissue was at the end which the physician believed was part of the papillary muscle. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿. section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of low pump flow and heart valve damage has been completed. the impella device was not returned for evaluation. the root cause of the heart valve damage issue was not determined since product was not returned and insufficient clinical information was provided. the root cause of the low pump flow issue was most likely biomaterial.
13753 | 1220648-2024-24095 | 20766977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766977&pc=OZD | 11/11/2024 | 01/15/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558095 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump had been supporting for 9 days when the pump stopped. the team attempted to troubleshoot and restart the pump; however, the pump could not be restarted and was explanted. no patient harm was reported. | [5 empty] | The investigation for the pump stop has been completed. the pump and data log were not provided for investigation. the pump stop occurred after 216 hours of case run. the cause of the pump stop issue was not determined since product was not returned and sufficient clinal information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13754 | 1220648-2024-24096 | 20766979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766979&pc=OZD | 11/11/2024 | 01/16/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517842 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to being implanted with an impella device for mechanical circulatory support. the device was not able to be inserted in the patient due to tortuosity in the vessel. placement was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues has been completed. no product/logs were returned. the impella implant was abandoned due to the patient's vessel tortuosity. the root cause of the delivery issues was patient condition related.
13755 | 1220648-2024-24097 | 20766980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766980&pc=OZD | 11/11/2024 | 12/23/2024 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123073 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella aic was not able to recognize the purge disc. the aic was exchanged. there was no patient involvement. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13756 | 1220648-2024-24098 | 20766982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766982&pc=OZD | 11/11/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563257 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient coded and flows were decreasing. the flows were as low as 0.1 lpm. there was a concern that the pump ingested a clot however, the clot was never confirmed. blood products were given to hopefully improve the flows. the patient expired after 1.00 hours of impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device nor the data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the low pump flow could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email.
13757 | 1220648-2024-24099 | 20766988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766988&pc=OZD | 11/10/2024 | 12/20/2024 |  | 11/23/2024 |  |  | 11/10/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552367 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, heavy bleeding was noted at the impella site. dressing was pulled down and manual pressure was being held. there was no hematoma. the decision was made to explant impella. manual pressure held until hemostasis obtained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13758 | 1220648-2024-24100 | 20766994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766994&pc=OZD | 11/05/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2025554604 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during insertion of the device, it was reported that the staff had too much wire in the left ventricle. the wire got wrapped around the inlet and kinked the catheter. the physician decided to remove the pump and replace it with another impella cp. no further information is available. | It was additionally reported that after the impella cp was replaced, the patient became hypokinetic with a mean arterial pressure in the 30's. fluids were given to the patient. the patient's pacer failed to capture, and the patient coded four times. staff then realized that 300cc of blood was in the patient's endotracheal tube and had caused the patient to not receive any ventilation for the previous twenty five minutes. support was discontinued and the patient expired. the relationship between the impella and the patient's outcome is unknown. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | Additional information is listed in b1, b2, g3, g6, h1, h6. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be use issue because the second pump was delivery successfully after the first pump was weaned due to kinking noticed on the catheter.
13759 | 1220648-2024-24101 | 20766995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766995&pc=OZD | 11/05/2024 | 08/14/2025 |  | 11/23/2024 |  |  | 11/05/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED. INC. |  |  | US |  | 1150343 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) exhibited a controller error. this occurred during preparation for a case with no reported harm to the patient. no further information is available. | [5 empty] | E1, e2, e3: the name and occupation of the initial reporter is unknown. the device was returned, and the evaluation is in process. upon completion of the investigation a supplemental mdr will be filed. | A.2 age should have been left blank on the initial manufacturer device report number 1220648-2024-24101 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2024-24101 as it is unknown. d.2 revised common device name as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-24101. d.4 revised serial number as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-24101. e.1 name prefix/title, first and last name are known and were added to the initial manufacturer device report number 1220648-2024-24101. phone number was entered incorrectly on the initial manufacturer device report number 1220648-2024-24101. e.2 and e.3 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24101. h.1 revised as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-24101. h.6 code 10 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24101.
13760 | 1220648-2024-24102 | 20766996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766996&pc=OZD | 11/11/2024 | 12/18/2024 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516567 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the while attempting to remove the device from the mitral structures and turn and advance it toward the apex, ventricle perforation occurred. according to abiomed support staff, the impella corkscrewed through the ventricular wall. a patch was sewn on to the left ventricle to repair the damage. the physician stated that area was as this as tissue paper and the sutures kept pulling through. the patch finally took. blood products were given. the impella was removed due to the issue. the patient outcome was noted to be critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warning; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings. ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the ventricular peforation has been completed. the impella was not returned for evaluation. based on the clinical details provided, the root cause of the ventricle perforation is most likely due to the patients condition at the ventricle tissue was thin.
13761 | 1220648-2024-24104 | 20766998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20766998&pc=OZD | 11/06/2024 | 01/16/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504536 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be implanted with an impella cp for mechanical circulatory support. the complainant reported that the impella cp was not able to pass through the open stent at the arch distal and insertion was aborted. no additional information is available, there was no reported harm to the patient. | [5 empty] | The investigation for the vascular damage and delivery issue has been completed. the impella was not returned for evaluation. however clinical details are sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the graft anastomosis. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. added b1 adverse event, b2 required intervention. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
13762 | 1220648-2024-24105 | 20767000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767000&pc=OZD | 11/06/2024 | 11/23/2024 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1177958 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation for a case the automated impella controller (aic) exhibited a purge disc failure alarm. the staff took the purge system out and put it back in, however the alarm persisted. the device was replaced with another aic controller, it was reported that the procedure was successful with no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of this issue was likely an electrical malfunction of either the purge pressure sensor or pud.
13763 | 1220648-2024-24106 | 20767001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767001&pc=OZD | 11/06/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/06/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567657 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. during delivery, the impella could not be advanced due to the patient's anatomy. the decision was made to abort insertion and implant an impella cp, it was reported that implant of the impella cp was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the cp was able to be inserted later. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings states the following: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
13764 | 1220648-2024-24107 | 20767005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767005&pc=OZD |  | 01/10/2025 |  | 11/23/2024 |  |  | 11/07/2024 | 11/23/2024 |  |  | 04/30/2024 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023251527 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported via literature that an impella 5.5 was placed in a patient with ventricular septal perforation (vsp) to provide circulatory support. during vsp surgery, it was confirmed that the impella had penetrated the free wall. the impella was pulled back and surgically repaired. additionally, at the time of removal, a blood clot was found attached to the outflow cage. | It was further reported that the patient was successfully weaned and explanted. |  |  |  |  | B5: date of awareness is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported thrombus and ventricle perforation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus and ventricle perforation could not be determined. b.5 patient outcome was added. e.1 revised address line 1 and city as they were reported incorrectly on the initial manufacturer device report number 1220648-2024-24107. g.2 additional selection was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-24107. h.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24107 and a new code was added.
13765 | 1220648-2024-24108 | 20767007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767007&pc=OZD | 11/02/2024 | 01/06/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533110 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a purge fluid leakage alarm occurred when the patient position was changed. the impella has been in use for beyond its service life, and a purge pressure increase alarm had also occurred just before. impella cp was removed urgently due to the pressure reservoir being damaged. patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of purge leak ¿ pump has been completed. the data logs were reviewed, and the purge issue was confirmed on the reported date. the high purge pressure failure mode was added because the voice of the customer expresses they noted a high purge pressure alarm shortly before the purge leak and it was confirmed in data log review. additionally, there were signs that blood/biomaterial had ingresses as high as the sidearm filter. the pump wasn't run to see if high purge pressure would reproduce because the sidearm had been cut off from the pump. based on data logs and returned product the root cause of the high purge pressure was determined to most likely be biomaterial. based on clinical details, data logs, and returned product, the root cause of the purge leak was determined to be a damaged sidearm reservoir. b3 date of event was reported incorrectly on manufacturer device report 1220648-2024-24108.
13766 | 1220648-2024-24109 | 20767008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767008&pc=OZD | 11/11/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  | 11/30/2025 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436210 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported upon delivery, the pump kinked at the arch. the pump was removed and replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | D.9 device returned date/information was added. h.6 medical device problem code 2682 was added due to investigation. h.10 insturctions for use for delivery issues was added. the investigation has been completed. the pump was returned for investigation. the pump had a cannula kink with no other abnormalities. the clinical information provided mentioned that the patient had a small aortic arch which led to a cannula kink during pump advancement. the pump implant was aborted due to that reason. therefore, the root cause of the delivery issue was patient condition related to patients small aortic arch. the cannula kink failure mode was changed to delivery issue since the cannula kink is a result of an issue faced during the pump delivery due to patient's vasculature. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13767 | 1220648-2024-24110 | 20767010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767010&pc=OZD | 11/04/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/04/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122397 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the automated impella controller (aic) had a controller error. this issue was discovered during prep. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller alarm-yellow alarm has been completed. the data logs were reviewed. upon boot up, the console registered no light since 0.47v is the approximate value of ambient light. then the optical bench reset itself but there continued to be no light seen from the lamp which triggered the controller error from the complaint. the console was also booted up 4 days later without triggering the low light voltage condition. the console was returned but the failure mode was not reproduced despite power cycling the console. the 5s was found to have marginally passing snr with barely failing contrast. the ccd array's fringe pattern was additionally found to have a small deformity. the cause of the bad lamp controller error was most likely due to a defective optical bench. corrections to the initial mfr report: g1 mdr reporting contact email
13768 | 1220648-2024-24111 | 20767012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767012&pc=OZD | 11/01/2024 | 12/27/2024 |  | 11/23/2024 |  |  | 11/01/2024 | 11/23/2024 |  |  | 06/30/2025 |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430984 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was being implanted with an impella cp device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. leading up to the event, the patient decompensated resulting in initiation of cardiopulmonary resuscitation (cpr) and advanced cardiac life support protocols. the decision was made to insert an impella cp pump, but the device was unable to pass the aorta-iliac bifurcation. multiple attempts were made, but the device could not advance through the calcified lesions. the family made the decision to stop resuscitation efforts and the patient passed away. the relationship between impella device and the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. corrections to the initial mfr report: g1 mdr reporting contact email.
13769 | 1220648-2024-24112 | 20767014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767014&pc=OZD | 11/01/2024 | 12/30/2024 |  | 11/23/2024 |  |  | 11/01/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547026 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the procedure, the iliac was meandering and was pressed too hard resulting in blood vessel damage. blood products were given as a result. the impella was removed and surgical repair was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage was most likely patient condition related.
13770 | 1220648-2024-24114 | 20767015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767015&pc=OZD | 11/08/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515306 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had no pedal pulse to the right leg. the patient's foot also started to look mottled. the impella was explanted and the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13771 | 1220648-2024-24113 | 20767016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767016&pc=OZD | 11/01/2024 | 12/31/2024 |  | 11/23/2024 |  |  | 11/01/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546552 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was access site oozing. a femostop was placed at the site and two units of packed red blood cells were given to treat the condition. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the pump was returned for investigation. upon investigation, catheter kink was noted near the blue suture hub and of the suture pad was broken. the observed product damage has no relation with the failure mode because the bleeding was controlled by using the femstop at the site. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13772 | 1220648-2024-24115 | 20767025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767025&pc=OZD | 11/11/2024 | 12/20/2024 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. on day 2 of the support, the limb became ischemic. the decision was made choice to explant and escalate the patient to the 5.5 pump for continued support. it was reported that the patient survived explant, no additional information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
13773 | 1220648-2024-24116 | 20767026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767026&pc=OZD | 11/11/2024 | 09/29/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | Device Sensing Problem | Positioning Problem |  |  |  |  | Arrhythmia | Ventricular Fibrillation | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | During repositioning of the pump the patient experienced ectopy which ultimately led to ventricular fibrillation and the patient was defibrillated once successfully. bedside echo done that showed tip of impella under pap. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump lost optical signal and later the pump stopped when the pump had been clamped. the pump was restarted successfully. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | Based on clinical details the root cause of the positioning issues was determined to most likely be a use issue. as the necessary information was not provided, the root cause of the vt/vf was not determined. without returned product, the root cause of the temporary pump stop could not be determined. b5 updated with information inadvertently not included in manufacturer device report 1220648-2024-24116. h6 health effect - clinical code 2130, 1721 health effect - impact code 4647 4628, medical device problem code 3009 were inadvertently not included in manufacturer device report 1220648-2024-24116. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the pump was not returned for investigation. data logs were reviewed, and placement signal not reliable (psnr) triggered just after starting up the pump. at this time, signal-to-noise ratio (snr), light, and contrast drop while gain and lamp increase. based on data logs, the root cause of the optical signal issues was determined to most likely be an air gap between the console and the pump plug. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-24116. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
13774 | 1220648-2024-24118 | 20767028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767028&pc=OZD | 11/08/2024 | 12/24/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560023 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following impella cp implant, the pump experienced lower than expected flows despite having optimized volume and positioning. placement signal alarms were continually alarming at the time of the issue. due to the ongoing issue, there were concerns about the catheter potentially having a kink due to the patient's tortuous anatomy. the pump was subsequently replaced and the issue was resolved. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation low pump flow and optical signal issue has been completed. the data logs showed that there was low pump flow at the beginning of the case. the product was returned and a kink in the cannula was found. based on the clinical details and product returned, the root cause of the low pump flow is due to a kink in the cannula. regarding the optical signal issue, the data logs showed that upon activation of the pump, the snr decreased, gain and light were stable and lamp and contrast were most likely stable. the returned pump's parameters were out of specification when returned. biomaterial or dried dextrose was found on the sensor face. the pump was soaked in warm water for 24 hours and it was removed. the parameters were taken again and they were in specification. no placement signal unreliable alarm was reproduced. the root cause of the optical signal issue is most likely due to biomaterial.
13775 | 1220648-2024-24117 | 20767030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767030&pc=OZD | 11/11/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance beyond the descending aorta. the implant was aborted until later that day when a second pump could be successfully placed. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, due to pad of patient doctor was unable to place impella past descending aorta. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ warnings & cautions: warnings. pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ warnings & cautions: warnings. pre-support evaluation. axillary insertion of the impella 5.5 catheter. use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ warnings & cautions: warnings. pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13776 | 1220648-2024-24119 | 20767031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767031&pc=OZD | 11/11/2024 | 01/15/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547865 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, a leak was confirmed near the pressure reservoir in the purge fluid line. fluid accumulated inside the side arm retainer and was also leaking to the outside. the amount of accumulated fluid increased and decreased, and new fluid had been visually confirmed to be coming out. there was no significant change in the purge pressure or purge flow rate, and no change in the motor current consumption. no visible damage was found. no patient harm was reported. | [5 empty] | The investigation for the purge leak has been completed. impella 5.5 was returned for evaluation. upon visual inspection, dried dextrose was present around purge sidearm retainer. no damage to sidearm retainer was noted. pump was returned cut along the catheter but still able to purge. pump was connected to the console and purged for a few hours in attempt to reproduce reported purge leak in sidearm. the purge leak in the sidearm reproduced with purge fluid pooling in the sidearm retainer and leaking out of retainer seams. the sidearm retainer was removed and the purge sidearm was wiped down to remove any excess purge fluid. upon further examination of the purge reservoir, purge fluid was present within the clear plastic retainer ring of the reservoir. no data logs were returned for evaluation. clinical details noted that the a purge leak was seen in the field at the sidearm and fluid was pooling in the sidearm retainer. no change to purge pressure or flow rate were noted and motor current remained stable. the amount of fluid pooling in and around the purge sidearm was noted to be fluctuated throughout support. the root cause of the purge leak was most likely due to a damaged sidearm reservoir. added d9 device return date. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
13777 | 1220648-2024-24120 | 20767032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767032&pc=OZD | 11/11/2024 | 12/23/2024 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546504 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the impella access site had continuous bleeding. forward tension was attempted with angle matching and manual pressure. the physician elected to explant the device due to bleeding complication. two units of blood products were administered. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13778 | 1220648-2024-24121 | 20767033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767033&pc=OZD | 11/11/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564933 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a leak was noted from the purge sidearm. no alarms were occurring. the pump was replaced as precautionary measure. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi: section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13779 | 1220648-2024-24123 | 20767034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767034&pc=OZD | 11/06/2024 | 01/23/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561269 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Renal Failure | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the patient¿s liver and kidneys were shutting down. additionally, suction alarm occurred and reduced flow rate. motor current values suddenly significantly increased, then a pump stop event with alarm. physicians tried to start the device a few times again but with no success. after removing the device, a blood clot was visible in the cannula. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the immpella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop, renal failure, and thrombus has been completed. data logs were reviewed and the pump stop was confirmed. there is a distinct step up in motor current that is sustained and speed deviation without a change in p-level. this is followed by the first trigger of alarm 13 (motor current mc high) and purge pressure spikes to roughly 900 mmhg without a trigger of the high purge pressure alarm over a 50 second period. these signatures are indications of biomaterial getting sucked into the purge gap. there are no suction alarms prior to the pump stop as was reported. additionally, data logs show that the pump was able to restart on its own twice after the initial stop. the first time, it runs for 8 minutes until a mc spike to 1245 and alarm 13 trigger. the second time, it runs for a 2 minute period with two distinct motor current spikes to 1455 and 1446. at the time of the second spike, the pump stops and is unable to be restarted. clinical details also report that a clot was found to be in the cannula during explant. product was not returned for investigation. based on clinical details and data log review, the root cause of the pump stop was determined to most likely be biomaterial. as the root cause of the pump stop was determined to most likely be biomaterial, the investigation into the pump stop will be representative of the investigation into thrombus. data logs were reviewed there were several factors that could have potentially played a role in the renal failure, however, none that are conclusive. on the first day of support, there were two triggers of impella in aorta alarms, however, they each resolved within a minute of the trigger. additionally, there were 11 triggers of suction alarms throughout the case, but they all resolved within a couple minutes of triggering as well. the placement signal had a slight trend downward throughout the case, which aligns with clinical details provided that the patient's condition was deteriorating. finally, between (b)(6) 2024, pump flows were saturated while running at p-8, but there were no motor current spikes that didn't correlate with a change in p-level leading up this. it was found that the pump stop was likely connected to biomaterial ingestion, but this ingestion event was 2 days after the reported renal failure, so it is unlikely to have been the cause for both. product was not returned for investigation. based on limited clinical details and inconclusive evidence from data logs, the root cause of the renal failure could not be determined. corrections to the initial mfr report: d4 updated udi number. | B3 date of the event was inadvertently reported incorrect on manufacturer device report 1220648-2024-24123.
13780 | 1220648-2024-24122 | 20767035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767035&pc=OZD | 11/08/2024 | 12/20/2024 |  | 11/23/2024 |  |  | 11/08/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, an aortic arch clot was noted unrelated to the impella. the family elected to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13781 | 1220648-2024-24124 | 20767036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767036&pc=OZD | 11/11/2024 | 01/08/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532954 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced limb ischemia. a leg bypass cannula was implemented and the ischemia resolved. support continued and it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13782 | 1220648-2024-24125 | 20767039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767039&pc=OZD | 11/11/2024 | 01/03/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560043 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support ongoing for a patient admitted in for post-cardiotomy cardiogenic shock from a multivessel coronary artery bypass graft x4. the pump was placed and supported for 6 days with persistent hemolysis signs. the hemolysis was seen with urine color change and elevated labs of plasma free hemoglobin and lactate dehydrogenase. the harm was treated by blood product infusions. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
13783 | 1220648-2024-24126 | 20767042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767042&pc=OZD | 11/11/2024 | 01/15/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533845 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump. it was documented that following insertion; the pump was unable to advance to the aortic arch due to the patient's tortuous anatomy. the pump kinked during the attempt and the decision was made to abort the implant. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the delivery issues has been completed. the pump was returned and upon investigation it was noted that the cannula was kinked. the root cause of delivery issue is determined to patient condition because of the tortuosity condition of the vessels of the patient. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. dance.¿
13784 | 1220648-2024-24127 | 20767044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767044&pc=OZD | 11/11/2024 | 01/02/2025 |  | 11/23/2024 |  |  | 11/11/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561035 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the flows went down suddenly and suction alarms were present. echocardiography imaging by the staff said the position was adequate and filling pressures were within normal limits. partial thromboplastin time (ptt) was therapeutic for last several days with last ptt 79.1 (60-80). there was a suspected clot in the cannula. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started and run for about 114 hours without issue, motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of low flow was biomaterial ingestion.
13785 | 1220648-2024-24128 | 20767045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767045&pc=OZD | 11/12/2024 | 01/07/2025 |  | 11/23/2024 |  |  | 11/12/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient moved, and the pump got pulled and moved causing a hematoma. the impella was removed due to massive hematoma. a new impella was inserted on right groin. the team could not get control of bleed. patient received 3 units of blood, was coded for 30 plus minutes. the family was consulted and agreed to stop all therapies and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement. while moving the patient over to catheterization lab table, the pump got pulled and moved causing a hematoma as per the clinical description provided. h.6 health effect - impact code 4629 was reported incorrectly on the initial manufacturer device report number 1220648-2024-24128. the correct code was added.
13786 | 1220648-2024-24129 | 20767046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20767046&pc=OZD | 11/13/2024 | 12/04/2025 |  | 11/23/2024 |  |  | 11/13/2024 | 11/23/2024 |  |  |  |  | 11/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1870819 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported for an automated impella controller (aic) that the field service team was performing the configuration and discovered that the cloud light emitting diode (led) and airplane led were not illuminating when pressed. troubleshooting of the console was unsuccessful, which consisted of restarting the console, a 15-second airplane press with no illumination when the button was pressed, a 60-second press with no illumination when the button was pressed and confirming that the universal serial bus (usb) was plugged in properly. there was no patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections to the initial mfr report: d2 device name has been updated. d9 return date added. updated h3 to device evaluation anticipated. h6 type of investigation code added. | The product was returned for analysis. after turning on the aic, rlm was confirmed to not light up leds. 5v was confirmed to come out of the 4-in-1 and received by the usb-a at the i/o port. the uusb was then pushed further into the rlm's uusb connector and the module powered on. the rlm was found to have symptoms of usd card corruption. there was no evidence of "loss of opm data" either in device logs or in complaint description.the rlm was not powering up because the uusb out of the backstrap cable was not properly connected to the rlm. g1 reporting contact email updated.
13787 | 1220648-2024-23881 | 20755959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20755959&pc=OZD | 11/16/2022 | 11/22/2024 |  | 11/22/2024 |  |  | 11/16/2022 | 11/22/2024 |  |  | 09/30/2024 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023258435 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient was in a computed tomography (ct) scan when it was noted that p level decreased to p2 and then impella position in aorta alarms occurred. there was also a sudden increase in purge pressure (pp) and drop in purge flow (pf) leading to purge system blocked alarm. when the patient was placed back on p6 motor current again, pulsatile and alarm resolved. left ventricle (lv) signal was noted to be high and flows saturated. pf/pp improved. ct revealed a large ischemic stroke. actions taken are unknown. additional information noted that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
13788 | 1220648-2024-23882 | 20756016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756016&pc=OZD | 10/30/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/30/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563135 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella 5.5. after multiple attempts, the impella 5.5 did not cross at anastomosis. after a final attempt, the pump and wire were removed and the pump drive shaft was found to have a large kink. a new access site was used and a new pump was opened and successfully implanted. the patient remained stable and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition. the kink of cannula was the consequence of delivery issue.
13789 | 1220648-2024-23883 | 20756092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756092&pc=OZD | 07/16/2024 | 12/16/2024 |  | 11/22/2024 |  |  | 10/30/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received multiple notifications via abiomed¿s long-term outcome and quality indicator (loqi) database reporting adverse events with "possible" relationship to the impella cp pump. the adverse events were ventricular tachycardia that required cardioversion, acute kidney injury which required continuous renal replacement therapy and progressed to renal failure that required hemodialysis, right lower extremity thrombus that required four compartment fasciotomy and thrombectomy of the right femoral artery and right external iliace artery, and right lower extremity ischemia that required above the knee amputation. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into vf/vt/af/at, renal failure, thrombus, and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13790 | 1220648-2024-23884 | 20756144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756144&pc=OZD | 10/31/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492533 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and after the implant, low pump flows were noted. on fluoroscopy image, kinking of the cannula between the radio pac marker and the outflow was noted. the staff tried to straighten the kink and switched to manual control to improve the flow to no avail. the pump was explanted and replaced. the second procedure was performed without any issue and the procedural outcome was successful and the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. there were no data logs returned, but the returned product had a kink in the cannula. based on the clinical details and product returned, the root cause of the low pump flow is most likely due to a kink in the cannula which most likely occurred during insertion. d.9 device return date/information was added.
13791 | 1220648-2024-23885 | 20756208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756208&pc=OZD | 10/31/2024 | 03/27/2025 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555645 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had impella 5.5 support and during support, the patient had ventricular tachycardia (vt). the impella was deep and the staff pulled the pump back. the vt and alarms resolved. the patient remains on support. | Additional information was received. it was noted that the pump had suction alarms throughout support. one unit of red blood cells and albumin were given to the patient. furthermore, the patient had hemolysis while on support. pump repositioning did not resolve the issue. the impella was removed and the hemolysis subsided. upon removal of the pump, large clots were visualized on the inflow and outflow of the impella. the patient survived the explant. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B.1 revised due to additional information received. b.5 additional information was received. d.6b explant date was added. d.9 date device returned/information was added. h.6 codes were added due to additional information received. the investigation for the ventricular tachycardia (vt), thrombus, low pump flow, and hemolysis has been completed. as the necessary information was not provided, the root cause of the vt was not determined. as the necessary information was not provided, the root cause of the thrombus was not determined. data logs were reviewed and the reported intermittent suction alarms throughout the case were confirmed. there were a total of 169 alarms throughout the case and the concentrations do align around the dates they were reported. there are no motor current spikes throughout the case, but the purge pressure is trending up slightly, though never going above specification. additionally, pump flow rates are unstable and sometimes below specification for the p-level. particularly during times of many suction alarms, it can be seen that the alarms will abruptly resolve and pump flows will improve - at least for a short period - indicating that the reported blood volume given to the patient improved pump performance. the returned product was investigated and biomaterial was found to be in the pump outlet. the biomaterial was removed and the pump ran in a bucket of water without low flows reproducing. since there were no motor current or purge pressure spikes during the case, and the biomaterial was clearly contacting the impeller on returned product, the biomaterial likely was not the cause of the low flows. based on clinical details and data log review, the root cause of the low pump flows was determined to most likely be patient condition. data log review did not show any abnormalities with motor current or purge pressure during the case, particularly the last few days of support leading up to hemolysis. there are some triggers of position unknown alarms, but none suggesting that the pump is in the aorta or ventricle. there are many triggers of suction alarms, the placement signal is relatively noisy, and pump flows are unstable and sometimes below specification. there was one particular point on (b)(6) 2024 where pump flows significantly dropped and ps spiked over the course of 20 minutes but there was no change in the motor current, p-level, or purge pressure. finally, as mentioned above, there were several dates that it was reported blood volume was administered and the logs showed suction alarms and pump flows suddenly improve. the returned product did have biomaterial wrapped around the impeller; however, data log review does not suggest that this biomaterial was worsening pump performance during the case based on the motor current and purge pressure stability. based on clinical details and data log review, the root cause of the hemolysis was determined to most likely be patient condition. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ e.1 added fax number as it was inadvertently omitted from the initial manufacture device report number 1220648-2024-23885.
13792 | 1220648-2024-23886 | 20756238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756238&pc=OZD | 10/31/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557336 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, bleeding was noted from the impella access site but also from other puncture sites. the patient was given one unit of red blood cells. the patient remained stable and the procedural outcome was successful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient condition since the patient was bleeding from multiple sites.
13793 | 1220648-2024-23887 | 20756305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756305&pc=OZD | 10/31/2024 | 12/24/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563804 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was placed in the catheterization lab and when getting ready to transfer the patient to the intensive care unit, a purge pressure high alarm went off followed by purge system blocked alarm. the left ventricle waveform above aortic waveform and high flows on associated p-level indicated flow saturation. the impella cp was removed and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the high purge pressure and saturated flow has been completed. data logs were reviewed finding there were no distinct motor current spikes seen but spikes in placement signal followed by a slight decrease in purge flow. this was followed by an increase in purge pressure and decrease in purge flow over the course of 25 minutes. after 25 minutes, high purge pressure alarms were triggered and there was saturated flow. the product was returned and there was no biomaterial in the purge gap or wrapped around the impeller. however, there was blood found in the purge line. there were no kinks found in the purge tubing. the root cause of the high purge pressure is most likely due to biomaterial. the root cause of the saturated flow is most likely due to biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ corrections to the initial mfr report: g1 mdr reporting contact email
13794 | 1220648-2024-23888 | 20756362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756362&pc=OZD | 11/02/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/02/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563758 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. it was reported when attempting to insert a long 25 centimeter (cm) peel away sheath, the sheath could not pass farther past arteriotomy. it was then noted that the stiff wire had a significant kink. it was then decided to utilize the right femoral artery for access then intervention via radial artery. a 13 cm impella sheath was inserted and positioned after 8, 10, & 12 french dilators. left ventricle (lv) access was attained with a pigtail catheter and 0.018 wire was positioned in the lv. the catheter was removed and the impella cp was then inserted over the wire but could not pass the iliac/aorta bifurcation. the impella cp was removed. the process was repeated with a 25 cm impella sheath but again was unable to pass iliac-aorta threshold. it was then decided to proceed with intervention without impella cp support. no patient injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the access site bleed, introducer damage, and delivery issues has been completed. the product was not returned for investigation. the data log review was not required for this case run. moderate plaque and tortuosity were reported while accessing the left femoral artery. the sheath couldn¿t pass through the arteriotomy, and it was found that the stiff guidewire was kinked. a small hematoma was controlled by applying manual pressure during the guidewire kink issue. the doctor changed the access site to the right femoral artery (rfa). the pump was not able to pass through the iliac/aorta bifurcation. the doctor tried to save time, but the attempt was unsuccessful. the cause of the delivery issue and product damage was most likely patient condition related to diseased/small vessels, which could also contribute to the access site bleeding and product kinking.
13795 | 1220648-2024-23889 | 20756467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756467&pc=OZD | 11/02/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/02/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025514922 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device as a bridge to a left ventricular assist device. it was reported that the impella 5.5 pump began to experience high purge pressure with low purge flows. a tissue plasminogen activator lysis protocol was sent to the staff to troubleshoot the issue, but the pump was ultimately replaced a few hours later. there was no patient harm resulting from the failure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. pump was not returned for investigation. data logs were reviewed and purge pressure was seen rising over the time of 2 hours at the same p-level. the last 15 minutes had a huge purge rise from 700mmhg to 1050mmhg followed by a motor current spike. this was followed purge pressure high and purge blocked alarms. these are characteristics of biomaterial ingestion. therefore as per the data log review, the root cause of the high purge pressure issue was biomaterial.
13796 | 1220648-2024-23890 | 20756501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756501&pc=OZD | 10/31/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511484 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the automated impella controller was replaced and when swapping to the replacement console, there was an ¿impella defective alarm¿. the pump was switched back to the original console and the same error occurred. the pump was then replaced. there was no reported patient harm. this is one of three related reports. the other two related reports are complaint number (b)(4). | [5 empty] | The investigation for the pump not detected has been completed. impella 5.5 pump was returned. visually inspected and no pump plug pins damage, biomaterial observed. electrical resistance test was done and all phase values were in normal range. pump was connected to console and impella defective alarm reproduced. red handle was opened and dried blood observed inside the circuit board of red handle. cather jacket from kink protector to motor housing side was cut open and blood was ingressed inside the catheter jacket completely. the external catheter jacket was inspected and hole in catheter, tearing evidence was found. no other abnormalities were seen. data logs were verified and impella defective alarm observed and after swapping to loaner console. no other issue seen. the cause of the pump not detected (impella defective) issue was blood ingress due to a hole in the catheter resulting from improper use. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
13797 | 1220648-2024-23891 | 20756561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756561&pc=OZD | 10/31/2024 | 03/23/2025 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1751872 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while trying to swap to a loaner automated impella controller (aic) an ¿impella defective alarm¿ appeared so they then tried to go back to the original console and received the same message. this issue is for the original aic. there was no reported patient harm. this complaint number two of three related reports. the other two related reports are complaint number (b)(4). | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of pump not detected has been completed. the console logs confirm that impella defective alarm was sent multiple times. the console's internal circuitry was inspected and no observable anomaly observed. the failure could not be reproduced when a similar known good pump was plugged into the console multiple times. the root cause of the impella defective alarm could not be determined as a console failure since it could not be reproduced. no problem found with this console during testing. d2 common device name revised on manufacturer device report in accordance with updated procedures.
13798 | 1220648-2024-23892 | 20756622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756622&pc=OZD | 10/31/2024 | 03/20/2025 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1653274 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while trying to swap to a loaner automated impella controller (aic) an ¿impella defective alarm¿ appeared so they then tried to go back to the original console and received the same message. this issue is for the loaner aic. there was no reported patient harm. this is complaint number three of three related reports. the other two related reports are complaint number (b)(4). | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the pump not detected has been completed. console logs confirm that impella defective alarm was issued multiple times. the console was returned for evaluation. the console's internal circuitry was inspected and no observable anomaly observed. the failure could not be reproduced when a similar known good pump was plugged into the console multiple times. the root cause of the impella defective alarm could not be determined as a console failure since it could not be reproduced. no problem found with this console during testing.
13799 | 1220648-2024-23893 | 20756810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756810&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531652 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient brought to the electrophysiology lab, prepped and draped in usual fashion. during insertion of an impella cp, calcium was noted at the insertion site of the sheath (which was not visualized well on initial angiogram). the left femoral artery was accessed using micropuncture and ultrasound and angiogram again showed an adequate size vessel and did not appear calcified. however, when the wire was inserted for sheath exchange, the surgeon felt the iliac. the aorta had diffuse calcifications (no focal lesions that could be easily treated). during this time, the impella cp had been opened and prepped but the surgical team agreed to abort impella cp placement due to high risk for vascular complications. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since iliac and aorta had diffuse calcifications.
13800 | 1220648-2024-23894 | 20756958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756958&pc=OZD | 10/30/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/30/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537910 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp for mechanical circulatory support. during support the patient experienced hemolysis. two units of packed red blood cells were provided to the patient. the patient was escalated to impella 5.5 device for support. no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed pump ran without issue during support. there were suction alarms triggered during the case but resolved quickly. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email.
13801 | 1220648-2024-23895 | 20756992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20756992&pc=OZD | 01/15/2022 | 11/22/2024 |  | 11/22/2024 |  |  | 01/15/2022 | 11/22/2024 |  |  | 06/30/2023 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098185 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.5 for mechanical circulatory support. during support, the patient experienced bleeding issues and heparin was withheld. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13802 | 1220648-2024-23896 | 20757104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757104&pc=OZD | 10/30/2024 | 03/07/2025 |  | 11/22/2024 |  |  | 10/30/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1490418 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the console was observed to have damage at the plug. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the supplemental mdr mfr report #1: h1-should have been malfunction. | The investigation for the console (aic) damage before therapy has been completed. console logs did not contain any information relevant to this failure mode. the aic was returned for evaluation. upon functional testing the issue was confirmed. judging by weak imprint of the strain relief bracket on the power cord shielding/insulation, the power cord set was most likely inadequately assembled by external vendor, so that after operator pulled by the cord, it became loose. the cause of the issue was a defective cable.
13803 | 1220648-2024-23897 | 20757139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757139&pc=OZD | 10/31/2024 | 12/23/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  | 01/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305672 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella migrated into the aorta and the position was adjusted in the cathlab. the decision was made to explant the pump and support with inotropes if needed. the impella successfully explanted and just on dobutamine, tolerating well. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique and the impella migrated into aorta. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
13804 | 1220648-2024-23898 | 20757263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757263&pc=OZD | 11/02/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/02/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a large hematoma at the impella access site. the site was redressed with angle match dressing and forward tension sutures were re-applied and hemostasis was achieved. additionally, the patient had hemolysis while on support that was noted by the patient's pink-red urine. one unit of packed red blood cells was provided to the patient for both the access site issue and the hemolysis. support was continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and access site bleed could not be determined.
13805 | 1220648-2024-23900 | 20757438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757438&pc=OZD | 11/04/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1560022 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the automated impella controller (aic) shut off unexpectedly during case simulation of functional check. the console was turned on, and did not connect to ac power. there was no reported patient involvement. | [5 empty] | The impella device was returned and evaluated. the root cause of the unexpected shutdown and ac power detection issues were determined to be cause by a defective power supply.
13806 | 1220648-2024-23901 | 20757459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757459&pc=OZD | 11/03/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534670 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella 5.5 for mechanical circulatory support. it was reported that the medical team was unable to pass a stenosed innominate junction. outer buddy wire, intravascular lubricants, and .027 inner rail all were utilized to attempt to cross the innominate junction without success. due to multiple advancement attempts, the cannula was kinked and scored in multiple locations and the catheter collapse at the catheter motor junction. the implant was aborted. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
13807 | 1220648-2024-23902 | 20757469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757469&pc=OZD | 11/01/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486749 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient exhibited mottling of the foot overnight. decision was made to explant the impella to help with perfusion to the right lower extremity (rle). the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
13808 | 1220648-2024-23899 | 20757503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757503&pc=OZD | 11/01/2024 | 03/24/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552614 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. while on support there was low purge flow alarms. the purge flow alarm was troubleshot by pigtail catheter and aic (automated impella controller) replacement. as the troubleshooting was completed without success, the was pump removed and replaced with a new pump. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of high purge pressure has been completed. the product was returned and evaluated. there were no abnormalities found upon visual inspection. upon running the pump with a test purge cassette using dextrose five percent (5%) in water, the low purge pressure and high purge flow reproduced with purge pressure at approximately 150 mmhg and purge flow at approximately 30 milliliters per hour (ml/hr). there were no leaks observed. the data logs also showed low purge pressure from the beginning with purge flow at approximately 30 ml/hr. the glucose concentration was 5% in field. no other abnormalities were seen with the data logs. although the pump passed all checks in manufacturing and was built to specification, it is likely that the tolerance stack up of the specific components of this build lead to less resistance and higher purge flow. the root cause of the low purge pressure was most likely a design limitation of the pump. d.4 revised lot number was inadvertently omitted from the initial manufacturer device report number. g.1 revised reporting contact email since the initial manufacturer device report was submitted in accordance with updated procedures. h.3 if the device was not evaluated selection was incorrect on the initial manufacturer device report. device evaluation anticipated, but not yet begun should of been selected. h.6 code 4112 should not of been reported on initial manufacturer device report as it belongs on this follow up report.
13809 | 1220648-2024-23903 | 20757587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757587&pc=OZD | 11/01/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1223935 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the grey selector knob on an impella aic (automated impella controller) was not working. the aic was exchanged, and problem was resolved. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) rotary knob issues has been completed. data logs were reviewed which showed that the day before the complaint date, there is a single missed kbd communication in the logs for ic4168 on boot up but this would have not been visible to the user. no button presses were attempted and the console was immediately turned off. the next day, the console is booted up without the kbd error and the console is immediately turned off by the user. the console is rebooted and pump 508056 is connected and run for an hour. near the end of the pump's connection to ic4168, the user opens numerous menus opened but the logs do not show any evidence of the gray knob being used. the console was returned for testing and after running a pump for an hour, the failure mode reproduced. despite rotating the gray knob, only a small subset all of the menu's options were selectable. the impellatronic board was replaced with a known-good board and the failure did not reproduce. the cause of the gray knob's inability to properly scroll was due to a defective impellatronic board. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email. h6 impact code 4629.
13810 | 1220648-2024-23904 | 20757649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757649&pc=OZD | 11/04/2024 | 12/26/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  | 08/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump. it was documented that following insertion, the impella pump was unable to advance into the abdominal aorta due to a fractured abdominal aortic stent. after multiple failed attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since unable to advance into abdominal aorta due to fractured abdominal aortic stent.
13811 | 1220648-2024-23905 | 20757777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757777&pc=OZD | 11/01/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562080 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.5 device for mechanical circulatory support. during implantation the pump was not able to cross into the aorta due to the patient's tortuous anatomy. after multiple attempts, a kink was noted to the impella 5.5 cannula and the device was removed over the wire. | [5 empty] | The investigation of delivery issues has been completed. the pump was returned and investigated. upon inspection kink was noted on the catheter near the cannula and the cannula of the pump. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through due to the tortuous condition of the vessels and the cannula was kinked during the insertion. g1 reporting contact email revised on manufacturer device report 1220648-2024-23905 in accordance with updated procedures. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
13812 | 1220648-2024-23906 | 20757778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757778&pc=OZD | 10/31/2024 | 12/23/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551820 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was inserted but there was pump saturation, decreasing purge flows and increasing motor currents that occurred. additionally, the impella cp failed initially during pump start but was restarted on first attempt. the impella cp was exchanged. it was further reported that the .018¿ guidewire was unable to be removed without dislodging the impella cp cannula. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use:impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention.section: potential adverse events (united states).¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿section: warnings & cautions: warnings¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿section: warnings & cautions: warningssection: pre-support evaluationsection: axillary insertion of the impella 5.5 cathetersection: direct aortic insertionsection: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿section: warnings & cautions: warnings¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the pump did not start and removal issues has been completed. data logs were reviewed, and the failed pump start was confirmed. during the initial attempt, alarm 13 triggered. the second attempt to start the pump 20 seconds later did work, but pump flows were immediately saturated. the pump was turned to p-9 and mean flows were 4.3 l/min (specification: 3.3-3.7 l/min). returned product was investigated and biomaterial was found to be wrapped around the impeller. the pump was attempted to be run in a bucket, but alarm 13 reproduced twice. the pump was unplugged, and once the biomaterial was removed from around the impeller, it was able to start up without problem. based on data logs and returned product, the root cause of the pump did not start was determined to most likely be biomaterial. clinical details further state that the customer states that the complication was guidewire issues during delivery. data logs are not relevant to the complication and the guidewire was not returned for investigation. clinical details do not report that excessive force was used or any patient conditions that would cause the issue. due to limited clinical details and product not being returned, the root cause of the delivery issues could not be determined.
13813 | 1220648-2024-23908 | 20757846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757846&pc=OZD | 11/01/2024 | 12/26/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479720 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.5 device for mechanical circulatory support due to cardiomyopathy. it was reported that the impella 5.5 was explanted due to the inability to be repositioned after an impella in aorta alarm occurred. patient was stable after removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was patient movement related while he stood up, coughed aggressively and got ¿impella in aorta¿ alarm which was unable to reposition. added h6 impact code 4628 corrections to the initial mfr report: g1 mdr reporting contact email.
13814 | 1220648-2024-23911 | 20757998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20757998&pc=OZD | 11/01/2024 | 12/20/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support due to cardiomyopathy. while on support, slight oozing was noted at the access site. there was no additional information regarding intervention. additionally, the patient went in supraventricular tachycardia (svt) overnight requiring adenosine x2. the patient was escalated to an impella 5.5 with continued venoarterial extracorporeal membrane oxygenation (va-ecmo) support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): "arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia." | The investigation for the ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
13815 | 1220648-2024-23912 | 20758032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758032&pc=OZD | 11/01/2024 | 01/21/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567162 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Sense | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Other | Life Threatening | Required Intervention | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient needing an impella 5.5 device for mechanical circulatory support due to cardiomyopathy. while on support the patient experienced acute swelling and a large hematoma formed under the skin at the graft attachment site. the staff held pressure to the area. the impella waveforms became damped and a vascular surgeon was called. the physician began evacuating and suctioning the hematoma. during the process the patient went into cardiac arrest. 4 units of packed red blood cell , 1 plasma, epi, and bicarb push were administered. return of spontaneous circulation was achieved. the patient was taken to the operation room and the surgical site was opened and cleaned. the physician was unable to find the source of bleed that caused the hematoma. no further information was provided as to the outcome for the patient. | It was also reported that one month after the patient was taken for evacuation of the hematoma, the patient started experiencing gastrointestinal bleed. an attempt was made to gradually wean off the impella as there were no long-term options for this patient. intravenous heparin was stopped and the purge solution was changed to dextrose fiver percent in water with heparin. additionally, the patient was started on protonix and given one unit of packed red blood cells. a failed optical sensor was noted. discussions were made with the family to consider palliative care as all other options had been exhausted. the pump was weaned unsuccessfully and a decision was made by the family to withdraw care and move to comfort measures. the patient eventually expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding, placement signal issue, and vascular damage - peripheral vessel has been completed. the failure mode (fm) placement signal issue was changed to optical signal issue because the 5.5 has an optical sensor. the fm access site bleeding is captured under vascular damage as vascular surgery was called due to the hematoma present. due to lack of product returned and limited clinical details, the root cause of what caused the vascular damage - peripheral vessel cannot be determined. the data logs show that the placement signal goes out on the 27th at the same time as shockwave was used for pci. based on the clinical details and the 5s parameters in the data logs, the root cause of the optical signal issue is most likely due to a damaged sensor from shockwave. | The investigation is still on-going. should the device or any new information is received, a supplemental manufacturer device report number will be filed. b.1 added product problem due to new information received. b.3 added death and date of death due to new information received. b.5 new information was received and added. d.6b explant date was added. h.1 selection was changed to death due to new information received. h.6 codes were added due to new information received. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23912. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23912. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23912 was submitted in accordance with updated procedures.
13816 | 1220648-2024-23913 | 20758123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758123&pc=OZD | 11/01/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512468 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the flows were down on the impella despite additional volume. additionally, the placement signal and flow metrics on the automated impella controller (aic) would not populate during and after insertion of the impella cp. the staff disconnected and reconnected the white power plug but noticed no change. the decision was made to exchange the pump for another impella cp. the following pump worked as expected. the patient coded, the family made the patient do not resuscitate, care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the optical signal issue has been completed. product was returned and evaluated. the pump's 5s parameters were in spec. the pump passed the laser continuity test and there was no significant biomaterial found on the pump. no other abnormalities were found. the pump ran in a bucket in the lab for 15 minutes. optical parameters were in spec and placement signal and pump flow displayed on the aic screen. field log review showed no issues with the optical parameters, but the placement signal was leveled out at zero for the majority of the case and the pump flow also dropped to zero. imc logs showed "err opsens sensor failure" events, which pointed to the console for investigation. no problem was found with the pump and it is apparent the console was the potential cause of the issue.
13817 | 1220648-2024-23915 | 20758251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758251&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  | 10/31/2024 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023284182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Ectopic Heartbeat | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella cp implanted. during the implant, the patient had ectopy. the implant continued and during the implant, a perforation to the posterior left ventricle occurred which caused the patient to decompensate requiring massive transfusion protocol, intubation and exploration by a trauma surgeon. it was further reported that the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of ventricle perforation and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13818 | 1220648-2024-23914 | 20758252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758252&pc=OZD | 11/05/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/05/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC |  |  | US |  | 1691800 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller experienced a purge disc failure during product training. there was no patient involvement or harm associated with the failure. | [5 empty] | The console was returned and an evaluation was performed to investigate the reported issue. during the evaluation it was discovered that the purge disc retainer was no longer attached to the purge assembly. upon conclusion of the investigation, the reported issue was confirmed and the cause was traced to a broken purge disc retainer. should additional relevant information become available, a supplemental report will be submitted.
13819 | 1220648-2024-23917 | 20758382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758382&pc=OZD | 07/03/2023 | 11/22/2024 |  | 11/22/2024 |  |  | 07/03/2023 | 11/22/2024 |  |  | 03/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Fragmentation | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella cp support and while on support the 0.018" wire tip came unraveled but was attached to the main body of wire. the wire was completely removed and inspected. flouro and echocardiogram was used to verify that nothing remained in the body. it was further reported that the patient expired on impella cp support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into the impella cp has been completed. no product was returned. per the clinical investigation, the root cause of the guidewire damage issue was not determined since product was not returned and insufficient clinical information provided. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13820 | 1220648-2024-23921 | 20758523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758523&pc=OZD | 11/05/2024 | 01/08/2025 |  | 11/22/2024 |  |  | 11/05/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2025556104 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed limb ischemia. the physician decided to explant and replaced with another impella. it was reported that the patient survived explant, no further information is available. | [5 empty] | E1, e2, e3: the name and occupation of the initial reporter is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 revised model number andunique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23921. e.1 revised name prefix/title, first name, last name, facility name, address line 1 and city as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23921. e.2 revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23921. g.1 revised manufacturing site fax number as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23921. h.6 code 3221 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23921 and was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-23921 in accordance with updated procedures.
13821 | 1220648-2024-23920 | 20758587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758587&pc=OZD | 06/28/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 06/28/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025483544 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Other | Death | Life Threatening | [22 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported an unsuccessful attempt to implant an impella cp device for mechanical circulatory support for to patient experiencing acute myocardial infarction/cardiogenic shock. shortly after arrival to the cardiac catheterization lab, the patient ventricular fibrillation arrested. angiographic imaging revealed patent coronaries on the left system. advanced cardiac life support was initiated, and the impella cp device was opened and prepped for insertion. the impella sheath was inserted in right femoral artery. resuscitative efforts were made and, however, stopped prior to the impella cp device being implanted. the patient expired in the procedural area. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of delivery issue is not determined because necessary clinical information was not provided. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions. never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
13822 | 1220648-2024-23922 | 20758606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758606&pc=OZD | 11/05/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/05/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025562925 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2024 |  |  |  |  | The user facility reported a patient presenting for aortic valve and mitral valve replacement was implanted with an impella cp pump for mechanical circulatory support. following pump and ecmo placement, the patient lost pulsatility down their impella leg. attempts at reperfusion were unsuccessful and diffuse/extensive thrombus formation was discovered down the leg. a thrombectomy was recommended, however the patient's family declined due to the patient's overall poor prognosis. the decision was eventually made by the patient's family to withdraw care and permission was given to harvest organs. it can not be definitively denied that the conditions may have contributed to the patient's passing. | [5 empty] | The impella device was not returned for evaluation. upon conclusion of the investigation, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: d4 updated model number f6/f8 should have been left blank.
13823 | 1220648-2024-23923 | 20758628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758628&pc=PYX | 11/12/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  | 11/30/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024431862 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella rp device for mechanical circulatory support. during impella support, the impella rp was suspected of clot ingestion. the impella flows decreased while motor current increased. pump stop alarm became active. the impella rp was removed. the patient is in stable condition after removal. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system. section: troubleshooting the purge system, ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop and thrombus has been completed. pump was not returned for investigation. data logs were returned and there was a motor current spike observed along with flows dropping at the steady p-level. these characteristics are similar to that of biomaterial ingestion. clinical also mentions that at explant biomaterial was observed. therefore, as per data log review, the root cause of the pump stop issue was most likely biomaterial. no further information regarding the patients anticoagulation status, infection status, indwelling lines, etc. was provided. corrections to the initial mfr report: g1 mdr reporting contact email.
13824 | 1220648-2024-23928 | 20758641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758641&pc=OZD | 11/07/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545155 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had thrombosis of both left anterior descending artery (lad) and circ stents. a thrombectomy was performed. additionally, when removing peel away sheath and inserting repo sheath there was a formation of large hematoma and pulsating arterial bleeding coming from access site that would not resolve with angle matching or manual pressure. the decision was made to explant the device. patient survived the procedure. | [5 empty] | Explant date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system (access site bleed) section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system (thrombosis) section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombus could not be determined.
13825 | 1220648-2024-23924 | 20758666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758666&pc=OZD | 11/07/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563043 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during the implant procedure, it was noted that the patient had unknown stenosis above aorta iliac junction. the implant was aborted. no patient harm or injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
13826 | 1220648-2024-23927 | 20758708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758708&pc=OZD | 08/31/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 08/31/2021 | 11/22/2024 |  |  | 04/30/2023 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022083141 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported that during impella 5.0 support, the patient experienced bleeding issues from the axillary impella access site. three units of packed red blood cells were transfused to the patient and heparin was removed from the purge system to manage the complication. support continued with the implanted pump following the intervention, however the decision was made to withdraw patient care. after seven days of support, care was withdrawn and the patient expired. the relationship between the impella and the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13827 | 1220648-2024-23931 | 20758821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758821&pc=OZD | 11/07/2024 | 08/25/2025 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147988 |  | 0042-0000-U |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the impella cp was plugged into the automated impella controller (aic), the aic was alerting that the impella was not plugged in all the way and then gave an optical sensor error. the aic was exchanged without any patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: h6 type of investigation code added.
13828 | 1220648-2024-23932 | 20758837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758837&pc=OZD | 11/07/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520561 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had a cp pump placed to support a patient admitted in with vfib and cpr with shock x8. the pump was placed but the poor positioning of it prompted pump removal and replacement after 2 days. the new 2ndpump was successfully placed for support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13829 | 1220648-2024-23930 | 20758889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758889&pc=OZD | 09/08/2024 | 10/06/2025 |  | 11/22/2024 |  |  | 09/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479675 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced hemolysis while on impella support. plasma free was greater than 300 at the time and increased from 252. it was later noted that the patient received total of five units of packed red blood cells due to a gastrointestinal (gi) bleed. the following day, the patient was still experiencing large gi bleed and received four units packed red blood cells overnight. hemoglobin was reported to be at six the following morning. subsequently, the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of vascular damage - peripheral vessel and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13830 | 1220648-2024-23934 | 20758895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758895&pc=OZD | 11/08/2024 | 01/09/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | GM |  | 2025510079 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 insertion the guidewire kinked due to the patient tortuous vessel. the guidewire was exchanged and impella was able to be inserted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, when trying to place impella 5.5 gw kinked due to tortuous anatomy of the vessel as per the clinical description provided. corrections to the initial mfr report: d4 updated model number
13831 | 1220648-2024-23933 | 20758913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758913&pc=OZD | 09/10/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 09/10/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466043 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported continuous bleeding issue occurred. the patient received one unit packed red blood cells and one of platelet count. the patient had surgical repair completed at the left groin incision due to continuous bleeding. additional information noted that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation has been completed. the device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13832 | 1220648-2024-23937 | 20758970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758970&pc=OZD | 11/08/2024 | 01/13/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629292 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported by the biomedical engineer that the automated impella controller (aic) had a controller error. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis confirmed the reported failure mode. the device analysis was able to reproduce the reported failure. root cause: the root cause for the controller error (defective optical sensor lamp) was most likely due to the defective optical bench or the lamp assembly e2 should have been selected as no on manufacturer device report 1220648-2024-23937. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
13833 | 1220648-2024-23935 | 20758976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20758976&pc=OZD | 11/08/2024 | 09/24/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5-EU |  |  | OZD |  |  | IMPELLA 5.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025515657 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient undergoing coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced arrhythmia. reportedly upon a recommended inspection of the aortic valve (av) it was noted that the av did not have good closure, the valve was adjusted. the patient remained on impella support and was successfully weaned. patient survived the procedure. | The complainant also reported that there were placement signal issues. the pump was confirmed to be in the proper position. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. as the necessary clinical information was not provided, the root cause of the vt/vf was not determined." | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed of optical signal issue: data logs were reviewed and there were several distinct drops in the pump's signal-to-noise ratio (snr)/contrast through the first couple days of support. additionally, the placement signal is trending downward throughout the case to values lower than expected. returned product was investigated and there were no abnormalities noted. the pump passed the laser continuity test. it was then connected to a console and the 5s parameters were all within specification. finally, it was run in a bucket for 2 hours and all 5s parameters continued to be within specification. based on the reported clinical details, data logs, and returned product, the root cause of the optical signal issues was determined to be patient condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1. product problem updated. b5. updated to include optical signal issue description. h6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-23935. b7. other relevant history, including preexisting medical conditions updated. d10. concomitant medical products and therapy dates updated.
13834 | 1220648-2024-23938 | 20759011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759011&pc=OZD | 11/08/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550469 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that once the peel away sheath was done, a hematoma was noted. reportedly to address the hematoma and oozing, the perclose sutures were tightened and mattress suture was done. the site was still oozing, and the decision was made to remove the pump and place a 16 french dry seal sheath. the site oozing was resolved and another impella was successfully placed in the same area with no further oozing reported. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
13835 | 1220648-2024-23939 | 20759076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759076&pc=OZD | 11/08/2024 | 12/01/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147986 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported by the field service engineer that upon startup of the automated impella controller, the screen showed an alarm failure of "system self-check failed/change console immediately". there was no patient involvement since the reported observation was found during preventive maintenance. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
13836 | 1220648-2024-23940 | 20759121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759121&pc=OZD | 10/31/2024 | 12/19/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526482 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. through the axillary artery. however, it was not successful. the pump was explanted and placed with the femoral access successfully. during support, the patient had a hematoma and significant output from the jp drain. the patient's right groin was also noted to be oozing with a hematoma. protamine was given, pressure was applied and two units of packed red blood cells were provided and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had tortuosity. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
13837 | 1220648-2024-23943 | 20759283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759283&pc=OZD | 10/31/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546116 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in heart failure/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient was having a large amount of ectopy. the impella was repositioned under echocardigram from 6.5 centimeters to 5.3 centimeters. there was an increase of ventricular ectopy when the patient was lying flat. the impella kept being reposition. additionally, the patient¿s dressing was saturated and leaking a little blood. the site was removed and redress. heparin was stopped due to bleeding from the jp site. it was further reported that the impella sterile sleeve separated from the locking hub. nursing stated that there was no force applied to cause this separation. the patient did ambulate the halls regularly. after ten days of impella support, the patient had a heart transplant and the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the access site bleed, ventricular ectopy, and product damage has been completed. there was no product returned. due to lack of evidence, the root cause of the access site bleed, ventricular ectopy, and product damage cannot be determined as the product was not returned and sufficient clinical details were not provided. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿.
13838 | 1220648-2024-23944 | 20759309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759309&pc=PYX | 11/15/2020 | 02/18/2025 |  | 11/22/2024 |  |  | 11/15/2020 | 11/22/2024 |  |  | 11/30/2020 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1387832 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella rp for mechanical circulatory support. it was reported that the pump had stopped during treatment. it was noted that the staff believed the pump may have been misplaced inside the right ventricle before the pump stop occurred. it was also noted that the patient received support from the device for 32 days which is beyond the expected life of the device. additionally, a literature review reported signs of hemolysis on the 15 day of support which resolved with blood volume. it also mentions the pump was displaced into the right ventricle, leading to explant. once care was withdrawn the patient expired. the relationship between the impella and cause of death is unknown. | Additional information has been received that information about care withdrawn/patient expired were incorrect. the patient has been successfully weaned and survived the incident. |  |  |  |  | The investigation into the impella rp has been completed. the impella pump was returned for investigation. product review reveals broken lemo connector pins. no biomaterial or abnormalities were found on the pump, although the pump appeared to have been cleaned prior to return. the data logs reveal a sudden motor current spike and then a high motor current pump stop without any electrical alarms. there are no motor current abnormalities prior to this. the cause of the pump stop was a bearing retainer wear. the cause of positioning issue was not determined as the clinical account does not dictate what caused the pump to move into the right ventricle. the cause hemolysis was patient condition related as the addition of blood volume resolved the hemolysis. d.4 expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Literature citation: gómez, m. g., uribarri, a., calvar, j. a. s. r., & stepanenko, a. (2021). clinical use of percutaneous mechanical circulatory assistance in a patient with end-stage right-sided heart failure and massive tricuspid insufficiency due to congenital heart disease: first-in-the-world case report. european heart journal - case reports, 5(8). https://doi.org/10.1093/ehjcr/ytab269 a.2 revised age as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23944. b.2 updated due to new information received since initial manufacturer device report number 1220648-2024-23944 was submitted. b.5 updated with clarifying information received since initial manufacturer device report number 1220648-2024-23944. d.4 revised expiration date and unique identifier (udi) as they were inadvertently omitted on initial manufacturer device report number 1220648-2024-23944. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-23944 as it was previously entered incorrectly. g.2 added selection as it was inadvertently omitted on initial manufacturer device report number 1220648-2024-23944. h.1 revised type of report due to new information received since initial manufacturer device report number 1220648-2024-23944 was submitted. h.3 revised as evaluation summary was inadvertently omitted on initial manufacturer device report number 1220648-2024-23944. h.6 due to new information received, health effect - impact code 1802 should not have been reported the initial manufacturer device report number 1220648-2024-23944. h.10 added literature citation as it was missing from the initial manufacturer device report number 1220648-2024-23944. attachments added literature report as it was missing from the initial manufacturer device report number 1220648-2024-23944.
13839 | 1220648-2024-23946 | 20759470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759470&pc=OZD | 04/07/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 04/07/2021 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED. INC. |  |  | US |  | 1470039 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the patient experienced bleeding and heparin was discontinued. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13840 | 1220648-2024-23947 | 20759530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759530&pc=OZD | 11/03/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551939 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding at the impella access site. as a result of the bleed, the impella cp was weaned and explanted and the patient was taken for surgical repair for a tear in the arteriotomy. the patient was intubated for the vascular surgery and was stable after the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13841 | 1220648-2024-23948 | 20759554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759554&pc=OZD | 06/04/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 06/04/2021 | 11/22/2024 |  |  | 12/31/2022 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022027308 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was being escalated to an impella 5.5. after the impella 5.5 was in the right axillary artery, the impella cp was removed from the right femoral artery. upon an attempt to close the impella cp access site, the hydrophilic tip of the perclose came off in the arterial track. the physician applied pressure to the site while they waited on a vascular surgeon. the femoral artery was cut down, repaired, and closed by the vascular surgeon. the patient received a massive blood transfusion protocol due to the blood loss caused by the perclose failure. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
13842 | 1220648-2024-23950 | 20759622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759622&pc=OZD | 11/01/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534720 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old female patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient began to experience hemolysis. volume was added and positioning was verified, however labs continued to come back hemolyzed. the decision was eventually made to remove the pump in response to the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email.
13843 | 1220648-2024-23951 | 20759656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759656&pc=OZD | 11/03/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492753 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella was found to be under papillary muscle on the echocardiogram. the staff was contemplating exchanging the pump versus repositioning. additionally, the patient had severe hemolysis so the decision was made to escalate patient to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and positioning issues has been completed. the impella has not been returned for evaluation. data logs were reviewed which showed that the pump was in the improper position corresponded with the time hemolysis was reported per clinical description. no other issues were found on the logs. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was most likely use related.
13844 | 1220648-2024-23942 | 20759990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20759990&pc=PYX | 05/10/2021 | 05/27/2025 |  | 11/22/2024 |  |  | 05/10/2021 | 11/22/2024 |  |  | 02/28/2023 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022056133 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. it was reported that the patient was rolled while attempting to place hypothermia cooling pads on the patient early this morning. the patient was rolled and the impella rp was subsequently pulled from the (pa) pulmonary artery. a chest x-ray showed the outlet of the impella rp in the ra (right atrium). the physician decided that with the patient's condition, they would rather not move them to replace the pump. the rp was pulled and a mattress suture was placed on the access site. it was noted that the patient was withdrawn from care and expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with automated impella controller instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance." | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13845 | 1220648-2024-23955 | 20760464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20760464&pc=PYX | 01/12/2022 | 11/22/2024 |  | 11/22/2024 |  |  | 01/12/2022 | 11/22/2024 |  |  | 06/30/2022 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1485188 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Us complainant reported the patient was on impella rp for hemodynamic support. while on support it was noted the patient was given lots of volume due to bleeding copious amounts from chest tubes. the patient received over 63 blood products. hemodialysis was also successfully initiated. however, post dialysis the patient was much worse. spouse decided to withdraw care, patient expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella rp with smart assist system with automated impella controller. section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
13846 | 1220648-2024-23956 | 20760624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20760624&pc=OZD | 10/30/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/30/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507344 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed limb ischemia and vascular surgery was consulted. the patient was taken to the operating room and had bilateral above the knee amputation performed. it was also reported that the impella was malpositioned with persistent low pump flow of 0.0 to 0.1 liters per minute. the physician reportedly attempted to unsuccessfully reposition the pump and suspected that the inlet was caught in the mitral valve structure. the physician made the decision to remove this impella and place a new impella, which was successfully placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. the root cause of low pump flow issue most likely was improper positioning issue due impella inlet was caught in the mitral valve structure, unable to reposition. the failure mode positioning issues removed because failure mode, low pump flow root cause is supporting this issue. after further case review limb ischemia was removed and should have been against the 2nd cp pump and will be reported in manufacturer device report 1220648-2024-23958. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23956. b1 changed to product problem. b2 should have been left blank. h1 type of reportable event. h6 codes 1942 and 4626 were reported incorrectly. new codes were added to health effect - clinical code and health effect - impact code. h10 instructions for use.
13847 | 1220648-2024-23957 | 20760675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20760675&pc=OZD | 06/22/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 06/22/2021 | 11/22/2024 |  |  | 06/22/2022 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1492674 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Hematuria | Thrombosis/Thrombus | [10 empty] | Life Threatening | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced new onset hematuria due to a clot in the bladder and dialysis was continued. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
13848 | 1220648-2024-23949 | 20760759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20760759&pc=OZD | 11/03/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537919 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp was placed to support a postpartum cardiomyopathy patient who was coagulopathic post cesarean section and hemorrhaging from the uterus post delivery. it was reported while on impella cp support, the patient experienced oozing at the impella access site. as a result of the bleeding, the patient was transfused with multiple units of blood, cryoprecipitate, platelets, fresh frozen plasma, and albumin. the patient's heparin was also withheld. support was continued with the implanted pump following the intervention. the patient was successfully weaned and explanted after 30 hours of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding additional sutures.
13849 | 1220648-2024-23381 | 20761069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761069&pc=OZD | 08/11/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 08/11/2021 | 11/22/2024 |  |  | 10/31/2022 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022026610 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella ld device for mechanical circulatory support. while on support, patient was bleeding. as a result, patient was given blood and platelets. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13850 | 1220648-2024-23958 | 20761349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761349&pc=OZD | 10/31/2024 | 12/20/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia of the lower left extremity. the pulses were lost at the limb and a vascular surgery was consulted. additionally, the patient had a new rectal bleed, and multiple blood transfusions were ordered. the physician reported the ischemia was due to venoarterial extracorporeal membrane oxygenation and the patient¿s condition not the impella. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the vascular damage and the limb ischemia has been completed. no product was returned for analysis. the root cause of vascular damage - peripheral vessel and the limb ischemia was not determined as insufficient clinical information was provided.
13851 | 1220648-2024-23962 | 20761586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761586&pc=OZD | 10/06/2024 | 11/03/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395260 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported ongoing support for a patient admitted in acute myocardial infarction/cardiogenic shock (ami/cgs). the pump lost the placement single but would return at certain p levels, but the p level was not appropriate support. the pump was explanted and extracorporeal membrane oxygenation replaced. in the investigation of the pump the automated impella controller (aic) was found to be having optical signal issues. the patient survived the event. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show multiple occurrences of placement signal not reliable (psnr) alarms (#1036) and multiple psnr events. there were also multiple suction alarms (#73) triggered by motor current algorithm. ltlog data indicated correlation between motor current and optical signal signal-to-noise ratio (snr) - for the first 2.5 hours of the case snr was low and placement signal was most of the time unreliable but improved after p-level was reduced. rtlog data shows in detail that despite sufficiently high snr at the very beginning of the case, as the p-level increased, the snr decreased, indicating pool quality of the signal resulting in psnr events and alarms. further detailed analysis of placement signal and its parameters presented in rtlog shows that snr and optical signal become less noisy for higher average values of placement, even at constant p-level. device was evaluated at aachen fs regarding issues reported in (b)(4) and (b)(4). it¿s been determined that analog output of the ccd detector in optical bench was distorted, which negatively affected optical placement signal measurement. placement signal issues investigated in (b)(4) were caused by failure of the optical bench. investigation (b)(4) describes findings related to rlm malfunction. there were no losses of opm data during the case. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h3 updated. h6 updated. d10 concomitant medical products and therapy dates updated. | D9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-23962 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13852 | 1220648-2024-23965 | 20761615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761615&pc=OZD | 11/01/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517036 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 5.5 pump support ongoing for a patient admitted in for a high-risk surgery. the pump stopped after 38 hours of use and was replaced. no lasting harm or injury from the interruption in support. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the pump stop has been completed. the pump was returned for investigation. data logs were not provided for this case. the returned pump passed the electrical tests, and significant biomaterial was observed on the impeller outflow cage. the pump was able to start after the biomaterial was removed. during testing, the pump flow appeared to be saturated, and further inspection suggested some biomaterial in the purge gap. as per the given clinical report, the pump stopped after 39 hours of the case run due to a purge-related alarm. the cause of the pump stop issue was biomaterial, based on returned product investigation.
13853 | 1220648-2024-23966 | 20761624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761624&pc=OZD | 11/01/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522371 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp was placed, and impella sensor failure alarm occurred. the impella was exchanged. the patient survived the event. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the optical signal issue was not determined as no product or data logs returned for analysis.
13854 | 1220648-2024-23968 | 20761676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761676&pc=OZD | 10/31/2024 | 12/26/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451622 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Life Threatening | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | PROTEK DUO | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 5.5 pump placed as care escalation from a cp and extracorporeal membrane oxygenation for the patient admitted in acute myocardial infarction/cardiogenic shock. the pump had been supporting for five days when there was ventricular tachycardia storm that prompted protek placement. the pump remains on for support and had not been explanted. approximately, five days later there was a complaint of hemolysis during the support. the patient had urine color changes. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation vf/vt/af/at and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined b3 date of event of reported incorrectly on manufacturer device report 1220648-2024-23968.
13855 | 1220648-2024-23970 | 20761680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761680&pc=OZD | 11/03/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561589 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had a head computed tomography which revealed subacute long posterior ciliary artery and middle cerebral artery lesions. there was no evidence of a thrombus when reviewed with echocardiography. the patient was noted to have the ability to move all extremities and follow commands. the patient was concurrently on venoarterial extracorporeal membrane oxygenation support and prior to impella was on an intra-aortic balloon pump. the patient remained on impella cp support until they were stable enough to determined if they no longer required mechanical circulatory support. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cerebrovascular accident has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the stroke/cerebrovascular accident could not be determined.
13856 | 1220648-2024-23972 | 20761931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761931&pc=OZD | 11/03/2024 | 09/29/2025 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474248 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Asian | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in non-st elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support as an escalation of care from an intra-aortic balloon pump. it was reported while on impella 5.5 support, the patient experienced an ischemic stroke and was treated with a thrombectomy. the medical team was unable to determine the origin of the stroke and it unknown if it was related to the impella device. the patient was unable to follow commands or move on the left side. the medical team made a care plan to wean the impella as the patient was no longer a candidate for advanced therapies. later, the patient had computed tomography which revealed a new stable right frontal subarachnoid hemorrhage. heparin was withheld as a result of a brain bleed. the patient's prognosis was poor and a plan was made to withdraw care. the patient was withdrawn from care and expired. the relationship between patient's outcome and the impella is unknown. | The complainant also reported that, after four days of support, the impella 5.5 lost its placement signal. the signal loss prompted no explant or intervention. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal review and stroke. the pump was not returned for investigation. the data log shows several temporary drops in signal-to-noise ratio (snr) to 2500 and contrast below 15 during the reported placement signal discrepancy issue, with no correlation to motor current. there was no major alarm reported on suction, not the positioning issue. placement signal correlation with patient condition or device performance was not confirmed from the data log. the cause of the optical signal issue was not determined since product was not returned and sufficient clinical information was not provided. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-23972. d10 concomitant medical products and therapy dates updated.
13857 | 1220648-2024-23975 | 20761950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761950&pc=PYX | 11/04/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025536963 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp flex for mechanical circulatory support. it was reported that after the impella rp flex was inserted and started, a motor spike and pump stopped occurred within seconds of being started. a restart was attempted with momentary success, but the pump again stopped with motor spike within seconds. the pump was explanted through the peel away sheath and a new pump was successfully placed. there was no harm to the patient reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation into pump did not start has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13858 | 1220648-2024-23974 | 20761959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20761959&pc=OZD | 11/06/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025473872 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient had an episode of ventricular tachycardia during impella 5.5 support. the pump was deemed to be too deep and was pulled back in response to the condition. as a result of the arrhythmia, the pump was scheduled for explantation the following morning. patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13859 | 1220648-2024-23973 | 20762041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762041&pc=OZD | 11/04/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550499 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding at the impella access site. manual pressure was held, but the bleed did not resolve. an additional suture was deployed and hemostasis was achieved. support continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, intervention with cut down and suturing to gain hemostasis as per the clinical description provided.
13860 | 1220648-2024-23978 | 20762064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762064&pc=OZD | 11/06/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2025563607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced ventricular tachycardia (vt) and ventricular fibrillation (vf). the patient was defibrillated multiple times to resolve the arrhythmias. the impella position was confirmed after the arrhythmia event. the relationship to the impella and the vt/vf is unknown. support continued with the implanted pump following the intervention. the patient remained on support for 246 hours, and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined.
13861 | 1220648-2024-23976 | 20762068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762068&pc=OZD | 11/06/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484451 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the right groin was oozing slightly. a stitch was placed at both sides and the oozing completely resolved. the site was angle matched and dressing was applied. support continued with the implanted pump following the intervention. after 5 days of support, the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding additional sutures.
13862 | 1220648-2024-23979 | 20762094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762094&pc=OZD | 11/04/2024 | 12/27/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542888 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported while on impella 5.5 support, the patient was heparin-induced thrombocytopenia positive. heparin was stopped. additionally, the patient developed a gastrointestinal (gi) bleed. the patient received a blood transfusion due to the gi bleed. support was continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombocytopenia and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13863 | 1220648-2024-23980 | 20762118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762118&pc=OZD | 11/06/2024 | 01/08/2025 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025561523 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella cp support developed bleeding from the impella access site. the 14f sheath introducer was left in to manage the bleeding and blood products were administered. hemostasis was achieved. the patient remained on support following the intervention, until they were successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ impella cp with smart assist system. section: general patient care considerations. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | Updated with missing information on manufacturer device report 1220648-2024-23980-01.
13864 | 1220648-2024-23910 | 20762128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762128&pc=OZD | 11/01/2024 | 01/07/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  | 01/31/2024 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181114 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. after removal, a mobile thrombus was observed in the ascending aorta. furthermore, thrombus in both common iliac arteries resolved with conservative treatment and continued anticoagulant therapy. the patient was bridged to a left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)". corrections to the initial mfr report: g1 mdr reporting contact email.
13865 | 1220648-2024-23919 | 20762157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762157&pc=OZD | 04/19/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 04/19/2021 | 11/22/2024 |  |  | 04/30/2021 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1426710 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered delivery issues. they were unable to advance the impella 5.0 due to the patient's anatomy and so the implant was aborted and the decision was made to implant an impella cp instead. there was no reported patient harm in regards to this event. | [5 empty] | The impella device was not received form the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as a part of the retrospective review of historical records.
13866 | 1220648-2024-23982 | 20762159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762159&pc=OZD | 11/06/2024 | 01/20/2025 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025557238 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient of unknown race with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during the support patient developed oozing from the impella access site. due to the bleeding bivalirudin was withheld. support continued with the implanted pump following the intervention. the device support is still ongoing. | Additional information received that the patient developed hemolysis. the impella was verified to be in the correct position and there was no evidence on imaging or a change in purge flows. the decision was made to exchange the pump and there was thrombus visualized on the outlet. a new impella 5.5 was successfully placed and the patient remains on support. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and thrombus is still pending, should any new information become available and supplemental report will be submitted. the returned product showed biomaterial wrapped around the impeller. the root cause of the hemolysis is due to biomaterial based on the data log analysis and returned product. instructions for use as follows impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock . section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ a5 ethnicity was reported incorrectly on manufacturer device report 1220648-2024-23982 | The investigation of access site bleeding and thrombus has been completed. due to lack of clinical details and insufficient product returned, the root cause of the access site bleeding and thrombus cannot be determined.
13867 | 1220648-2024-23981 | 20762296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762296&pc=OZD | 11/03/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565080 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced limb ischemia. as a result of the limb ischemia, the impella was explanted. the patient was successfully weaned and explanted with no further impella related issues. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13868 | 1220648-2024-23926 | 20762308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762308&pc=OZD | 11/07/2024 | 12/19/2024 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535318 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the single access technique was used. the pump access site had a bleed and hematoma when the peel away was removed and reposheath placed. the medical team could not control the bleed with pressure or angle matching. the pump was explanted, and 2 units of right blood cells (rbc) were given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13869 | 1220648-2024-23941 | 20762316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762316&pc=OZD | 04/24/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 04/24/2021 | 11/22/2024 |  |  | 02/28/2023 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2022052262 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | Black Or African American | Death | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.5 for mechanical circulatory support. during support, it was reported that the device was being ran at subtherapeutic levels due to bleeding concerns. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13870 | 1220648-2024-23961 | 20762320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762320&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484442 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Separation Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, after repositioning, the patient's urine output became pink/red. the patient remained on low dose levo overnight with low impella support. discussed leaving impella in longer due to continued levo but decided on removal based on oozing concerns and hemolysis concerns. the patient was successfully weaned and device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the product and data were not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the clinical details outline that the clear ring on the introducer was unable to break off of the catheter. a failure mode for introducer damage was added to capture this. the root cause of the introducer damage was unable to be determined as no product was returned for analysis. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23961. b1 adverse event and product problem. h6 code 4641 and 4644 were reported incorrectly. new codes was added to health effect - impact code.
13871 | 1220648-2024-23918 | 20762420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762420&pc=OZD | 11/05/2024 | 01/09/2025 |  | 11/22/2024 |  |  | 11/05/2024 | 11/22/2024 |  |  | 11/30/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET |  |  | ABIOMED, INC |  |  | US |  | 2025495564 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed signs of hemolysis during impella cp support. the patient's urine color had changed and their pfhg levels had risen. attempts to troubleshoot the issue included echo positioning checks and fluid boluses, but the hemolysis persisted. the patient was put on dialysis and remained on pump support for one more day despite the noted condition. | [5 empty] | The impella device was discarded by the customer and therefore, evaluation of the device is not possible. upon conclusion of the investigation, a supplemental report will be submitted. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. the pump and data log were not provided for investigation. according to the clinical team, there was concern over hemolysis on the third day of pump use. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided.
13872 | 1220648-2024-23985 | 20762466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762466&pc=OZD | 11/09/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/09/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 81-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed hemolysis. the pump was removed in response to the noted condition and an impella cp pump was placed for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. data logs confirmed the failure mode and clinical narrative. logs showed placement signal was observed to be trending downward approximately nine hours into support. weaning was begun due to the hemolysis 13 hours into support. the product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely patient condition related (due to trending placement signal and since the spike in motor current was at p1 after the weaning process was already initiated). b.7 information was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23985. d.4 revised serial number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-23985. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23985 was submitted in accordance with updated procedures.
13873 | 1220648-2024-23925 | 20762561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762561&pc=OZD | 09/11/2024 | 02/10/2025 |  | 11/22/2024 |  |  | 09/11/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525457 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 placed via axillary insertion for a patient admitted in acute myocardial infarction/cardiogenic shock. the pump was placed via the axillary. however, there was an access site bleed at the site. the team achieved hemostasis at the site by placement of a hemoshield graft, side suturing and the delivery of blood products. support proceeded for six days where the patient remained stable and the procedural outcome was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was inadvertently reported incorrectly on manufacturer device report 1220648-2024-23925.
13874 | 1220648-2024-23983 | 20762584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762584&pc=OZD | 11/08/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/08/2024 | 11/22/2024 |  |  | 07/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441407 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | Yes | Incomplete or Inadequate Priming | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was in implanted, however the staff was unsure if the device crossed the valve. device was not primed or setup completed appropriately prior to insertion. the device was explanted, it was noted that proper steps were not completed prior to insertion. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of priming issue has been completed. data logs were reviewed and during case start, the purge cassette and pressure disc were recognized without a problem. during the first attempt to prime the cassette, the procedure got to 90% and then appeared to stall out as it did not progress any further for over two minutes before the user canceled the case start. they then restarted the procedure and got to 50% before the console triggered the "purge fluid not detected error." at this time, the purge ticked in real time and logs showed the piston stalling. there are signs that there was no fluid for priming the cassette, but it is unclear why. the purge cassette was then unplugged, plugged back in, and the case started to get up to step 7, but they canceled and aborted. clinical details align with what was seen in the logs. the priming procedure was not correctly done at first, but when the representative arrived and helped to correctly complete it, the decision was made to abort the case. the returned product was investigated and there were no visual abnormalities. the purge cassette was plugged into a console and the priming procedure was completed successfully and then it was connected to the returned pump and was run for five minutes without issue. based on all the above information, the root cause of the priming issue was determined to be a use issue. d.9 device return date/information was added. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23983 was submitted in accordance with updated procedures. h.6 investigation findings code 3221 was inadvertently missed on the initial manufacturer device report number 1220648-2024-23983 and was added.
13875 | 1220648-2024-23916 | 20762658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762658&pc=OZD | 11/05/2024 | 12/18/2024 |  | 11/22/2024 |  |  | 11/05/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025552979 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a hematoma around the sheath. heparin was stopped due to the hematoma and the sheath clotted off. two units of blood were given in response to the condition. plans were made to stop using the sheath and to continue to wean down impella cp support. | [5 empty] | The impella device was not returned by the customer for evaluation. upon conclusion of the investigation, a supplemental report will be submitted. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d4 model number was reported incorrectly on manufacturer device report.
13876 | 1220648-2024-23971 | 20762768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762768&pc=OZD | 11/04/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1717562 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that during preparation of an automated impella controller, errors controller error and de-air occurred. this complaint was not related to a clinical procedure. | [5 empty] | The investigation for the reported alarm issue has been completed. console logs from the reported day of event were reviewed and found to be mostly consistent with the complaint and showed alarms related to de-airing (¿air in purge system¿ and ¿purge pressure low¿) occurred while a demo pump and purge cassette were utilized, but there were no ¿controller error¿ alarms recorded. console was tested with a known-good purge cassette, but the issues were not reproduced, and the console operated within specification. the ultimate root cause of reported alarms could not be determined, but likely issues were use-related, and caused by a non-clinical use scenario. we could not exclude either defective demo purge cassette or defective demo pump - neither of which was returned for analysis.
13877 | 1220648-2024-23989 | 20762829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762829&pc=OZD | 11/12/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in for a high risk percutaneous coronary intervention (hrpci). the pump was placed but there was low flows from suction. the team explanted and replaced the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23989 in accordance with updated procedures.
13878 | 1220648-2024-23988 | 20762945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20762945&pc=OZD | 11/10/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/10/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella cp device for mechanical circulatory support. it was reported that the nurse turned on the impella before placement in the patient, the easy guide lumen was still present in cannula. the staff noted that product mishandling by the hospital staff led to cannula damage. it was reported that the procedure was successful, another device was utilized. no further information is available. | [5 empty] | The investigation for the pump did not start has been completed. data logs were not returned for review. returned product confirmed the clinical story: the full red lumen was still in the pump. there was a kink/damage in the lumen, likely from when the team attempted to start the pump while it was still in the cannula. based on clinical details and returned product, the root cause of the pump did not start was determined to be a red lumen issue. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states the following: section: using the automated impella controller with the impella catheter "achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup." section: using the automated impella controller with the impella rp with smartassist system catheter. "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the "purge pressure high" alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen." | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
13879 | 1220648-2024-23977 | 20763028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763028&pc=OZD | 11/04/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/04/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529550 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at impella access site. systemic heparin was withheld and the patient was transfused with one unit of red blood cells and one unit of platelets. support was continued with the implanted pump following the intervention. the patient was successfully weaned and explanted after completion of a high risk percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding ¿ major has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient condition because of the bleeding from multiple sites along with the impella site.
13880 | 1220648-2024-23952 | 20763096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763096&pc=OZD | 11/01/2024 | 12/20/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545549 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 60-year-old female patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, hematoma developed after peel-away sheath was exchanged for repo sheath. manual pressure was held for two hours. bruising was noted. pressure was relieved and no oozing was present. the physician opted to inject lidocaine/epinephrine intradermal to help. complications were successfully managed. additionally, the impella had to be kept shallow otherwise the pigtail tickles patient¿s ventricle resulting in a significant amount of arrhythmias. impella had intermittent ¿impella in aorta alarm¿ that is triggered when patient¿s systolic blood pressure (sbp) was below 90 mmhg. echo ordered to verify position. the physician ordered to keep impella shallow, which helps with arrhythmias but risks impella inlet from coming out of the ventricle and causing suction on atrioventricular. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the positioning issue and the arrhythmias has been completed. no product was returned for analysis. clinical details mention patient had small aortic arch/ventricle which contributed to positioning issues. the root cause of the positioning issue was most likely patient condition related. the root cause of the arrhythmia could not be determined without additional clinical details. corrections to the initial mfr report: b1-should have selected product problem. g1 mdr reporting contact email.
13881 | 1220648-2024-23991 | 20763098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763098&pc=OZD | 11/06/2024 | 01/08/2025 |  | 11/22/2024 |  |  | 11/06/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533570 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an unknown primary indication was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the pump could not cross the iliac artery due to a calcified stenosis. the implant was aborted, and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since patient had calcified stenosis.
13882 | 1220648-2024-23953 | 20763099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763099&pc=OZD | 11/03/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557337 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella 5.5 support had run of ventricular tachycardia (vt) requiring one shock, amiodarone bolus and drip. three days later, the patient went into vt twice in the morning requiring shocks both times. amiodarone and lidocaine drip were added. at the end of impella support the patient was bridged to left ventricular assist device. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
13883 | 1220648-2024-23936 | 20763133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763133&pc=PYX | 11/03/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025562134 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. prior to impella rp flex being implanted, the patient began having episodes of ventricular tachycardia (vt) and ventricular fibrillation (vf). it was reported that during the impella rp flex insertion the patient continued to be in vt/vf. the patient was defibrillated multiple times but unable to sustain a normal sinus rhythm. the patient became more unstable and vasoactive support was increased. the physician decided to abort impella rp flex placement and place patient on venoarterial extracorporeal membrane oxygenation for stabilization of lethal heart rhythm. | During delivery of the impella, the patient was noted to be unstable, and the physician aborted the implant and placed the patient on venoarterial extracorporeal membrane oxygenation (va ecmo). no further information is available. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the delivery and ventricular tachycardia (vt) has been completed. the impella was not returned for evaluation. however clinical details were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the patient suffered vt/vf during insertion and the pump was unable to pass through the vasculature and the insertion was aborted. | Additional information was added to b1, b5, g3, g6, h2. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. additional information for the initial mfr 1220648-2024-23936 was added to h6 and h10. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
13884 | 1220648-2024-23929 | 20763297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763297&pc=OZD | 11/07/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/07/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524983 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | Other | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, patient had ventricular tachycardia (vt) on day 6. the patient was shocked 3 times. then on day 8, the patient again had a ventricular fibrillation (vf). the patient additionally was seizing, and an eeg was planned. the family made the choice to withdraw care. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
13885 | 1220648-2024-23907 | 20763484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763484&pc=OZD | 11/01/2024 | 12/27/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547608 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.5 device for mechanical circulatory support. the 5.5 pump failed to deliver despite all measures and troubleshooting. a second impella 5.5 was attempted and after multiple attempts, the decision was made to remove the second impella 5.5 due to the patient's tortuous anatomy and place an impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the patient had tortuous anatomy. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added h6 component code 4627 corrections to the initial mfr report: g1 mdr reporting contact email.
13886 | 1220648-2024-23987 | 20763516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763516&pc=OZD | 11/09/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/09/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. after successful explant of an impella cp pump, a thrombus was discovered to be causing a flow defect in the artery. the patient was referred to vascular surgery to assess the condition. the resolution for this issue is unknown. it was reported that the patient survived normal explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email.
13887 | 1220648-2024-23945 | 20763662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763662&pc=OZD | 11/02/2024 | 12/27/2024 |  | 11/22/2024 |  |  | 11/02/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Bradycardia | Cardiac Arrest | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the doctor attempted to reposition the impella with echocardiography as the impella was noted to have moved after patient transport. the physician had difficulty unscrewing the lock to reposition the catheter. at this time the patient began to experience bradycardia and lost pulses. the patient coded for 25 minutes before the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue and ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue and vt could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13888 | 1220648-2024-23990 | 20763779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763779&pc=OZD | 11/10/2024 | 12/31/2024 |  | 11/22/2024 |  |  | 11/10/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558880 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 79-year-old female patient was on impella cp support and after the implant, a femoral artery perforation was noted. the impella was removed, and a covered stent was placed successfully. it was reported that the patient survived normal explant. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13889 | 1220648-2024-23909 | 20763788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20763788&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552422 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the bilateral femoral access was complicated and deemed inappropriate for impella insertion. left axillary access was unsuccessful. the decision was made to abort impella insertion. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, bilaterally femoral access complicated and deemed inappropriate for impella insertion.
13890 | 1220648-2024-23954 | 20764254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764254&pc=OZD | 04/30/2021 | 11/22/2024 |  | 11/22/2024 |  |  | 04/30/2021 | 11/22/2024 |  |  | 05/31/2022 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED. INC. |  |  | US |  | 1435327 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced a hematoma and heparin was withheld. it was also reported that the device was suctioning at p2 which indicates low flow. the resolution for this issue is unknown. it was noted that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13891 | 1220648-2024-24004 | 20764291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764291&pc=OZD | 11/13/2024 | 01/07/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285254 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the automated impella controller (aic) was turned on the battery level showed as 0%, even when plugged in, and an error message occurs. there was no report of patient involvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the automated impella controller (aic) battery failure has been completed. data logs were reviewed finding that the reported battery level 0% issue was confirmed via data analysis. an instance of battery failure (transport connection blown/missing) alarm was observed in the logs. device analysis of the aic was not performed because the issue was reportedly resolved onsite and was not returned. the reported battery issue was determined to be use related. it is likely that the aic was booted up while connected to alternating current power and the battery switch was disengaged since the issue occurred upon first boot of the console after routine functional check was performed on it. a.1 and a.3 revised to reflect no patient was involved as they were left blank on the initial manufacturer device report number 1220648-2024-24004 and should not have been. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24004. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-24004.
13892 | 1220648-2024-23993 | 20764433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764433&pc=OZD | 11/12/2024 | 01/08/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568407 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella 5.5 device for mechanical circulatory support. during implantation, a tortuous patient anatomy caused the 5.5 catheter to give next to the motor housing causing lack of pushing capability and was unable to progress that pump any further, so it was necessary to open new 5.5 and attempt again. the physician noted that there were no complaints on the device and blamed it on the patient¿s anatomy. no further information is available. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issue has been completed. the pump was returned for investigation and no damage was noted on the pump. the root cause of delivery issue is determined to be use issue because the second pump was successfully inserted with manipulation through anastomosis. corrections to the initial mfr report: h6 impact code 4627 changed to 4629
13893 | 1220648-2024-23969 | 20764435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764435&pc=OZD | 11/01/2024 | 12/27/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516907 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 5.5 pump support ongoing for a patient admitted in cardiogenic shock and with cardiomyopathy. the pump was supporting when the purge pressure rose. the team made the choice to not explant the pump and to add tissue plasminogen activator lysis to the purge. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. neither the impella device nor the data logs were returned for evaluation. additionally, clinical details provided were limited to determine the root cause. therefore, the root cause of the high purge pressure could not be determined. added h6 impact code 4644.
13894 | 1220648-2024-23995 | 20764437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764437&pc=OZD | 11/12/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and is now on dialysis. the patient remained on support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed after pump started and run without issue for 183.5 hours, motor current (mc) suddenly spiked followed low flow that triggered auto suction response and suction alarms. pump then ran on lower p level without issue to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely biomaterial ingestion.
13895 | 1220648-2024-23998 | 20764443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764443&pc=OZD | 11/12/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  | 06/30/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373475 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, due to difficult imaging, the impella was pulled back too far into the aorta and was unable to place. the decision was made to explant the impella and replace it with a new one. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
13896 | 1220648-2024-24001 | 20764504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764504&pc=OZD | 11/12/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614416 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported for an automated impella controller (aic) that during a function test the aic was not priming. the alarm log showed ¿unexpected controller shut down¿. this issue occurred during a function test, not during patient support. | [5 empty] | The investigation for the console (aic) shutdown issue has been completed. console logs from the reported event show an unexpected shutdown followed by another shutdown just seconds later. both shutdowns triggered "unexpected shutdown" alarms after the system rebooted. after each shutdown, the console's software automatically restarted. the first reboot occurred before the imcgui failed, prior to the sw watchdog being established. the second reboot occurred after the watchdog was reset. the logs did not show any errors related to priming the purge system. after removing and reinserting the purge cassette, the representative restarted the case start wizards. according to the logs, the priming was successfully completed and luers was detected 100% the case start was later cancelled. the console was returned for evaluation, but the issue could not be reproduced, and no unexpected shutdowns or irregularities in performance were observed. no power supply interruptions were detected between the pbm and the 4-in-1 assembly. the most likely cause of the shutdown was a software issue. the console's purge unit was tested with a known good purge cassette, and visual inspection of the purge assembly revealed no excessive purge fluid residue. the root cause of the aic priming issue was not related to any problems with the aic. no issues were found within the aic purge unit. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
13897 | 1220648-2024-24005 | 20764520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764520&pc=OZD | 11/14/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567209 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support and during the implant, the pump kinked at several points including cannula and juncture of cannula/white cable. after the implant, motor current dramatically reduced and pulsatility ceased, flows decreased to 0.0-0.1. position was confirmed and troubleshooting efforts did not resolve the issue. the pump was explanted and replaced with a new impella 5.5. there was no patient harm reported. | [5 empty] | The investigation of low pump flow has been completed. the returned compliant 5.5 pump visually inspected. cannula kink and slight catheter kink at connection point observed. no biomaterial and no broken blade observed. the clinical details stated: tortuous left axillary implant of 5.5. impella kinked at several points including cannula and juncture of cannula/white cable. the root cause of the low pump flow was patient condition related due to tortuous anatomy of left axillary that kinked the device. d9 date device returned to manufacturer added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
13898 | 1220648-2024-24007 | 20764523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764523&pc=OZD | 11/13/2024 | 12/20/2024 |  | 11/22/2024 |  |  | 11/13/2024 | 11/22/2024 |  |  | 07/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371625 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.5 device for mechanical circulatory support due to cardiomyopathy. the pump supported for over 49 days when it was explanted as the patient had successfully bridged to transplant. later in (b)(6) 2024 the medical team published upon this patient and another in which the pump used sodium bicarb in the purge. the publication noted that during the impella support there were 2 adverse events attributed to the use of the pump. the accepted manuscript from the university of iowa pharmacy team entitled " short-term use of an impella ventricular assist device sterile water¿based bicarbonate purge solution for positron emission tomography scanning" details the 2 patients and events. one of whom developed ae attributed to the use of impella. the 5.5 supported patient had thromboyctopenia and infection. the patient survived the reported harms. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. schumacher, s. m., & hobbs, r. a. (2024). short-term use of an impella ventricular assist device sterile water¿based bicarbonate purge solution for positron emission tomography scanning. american journal of health-system pharmacy, 81(21). https://doi.org/10.1093/ajhp/zxae160. | The investigation for the reported thrombocytopenia and sepsis has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombocytopenia and sepsis could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1 mdr reporting contact email.
13899 | 1220648-2024-24000 | 20764535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764535&pc=OZD | 11/12/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515182 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient suffered an embolic stroke. a thrombectomy was performed and the pump was subsequently explanted in response to the condition. the patient survived the operation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of stroke/cva has been completed. the returned product was visually inspected and no abnormalities were found. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
13900 | 1220648-2024-24018 | 20764548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764548&pc=OZD | 11/14/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568405 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.5 device for mechanical circulatory support. while on support there were low flows/suction that were persistent. the team attempted to troubleshoot but the low flows persisted. additionally, the patient experienced ventricular tachycardia (vt) that was treated by defibrillation. the pump was repositioned due to the vt and low flows and suction. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow and arrhythmia has been completed. the impella pump was returned for evaluation and visually inspected. biomaterial observed around impeller and purge gup. a cannula kink was observed near outlet. no broken bladed found. the root cause of the low pump flow issue was most likely biomaterial ingestion, and cannula kink may occurred during removal. the root cause of the arrhythmia could not be determined without additional information.
13901 | 1220648-2024-24009 | 20764612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764612&pc=OZD | 11/13/2024 | 02/20/2025 |  | 11/22/2024 |  |  | 11/13/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556366 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a purge alarm, and the team observed a purge leak. the critical patient could not come off therapy to replace the pump, so the issue was troubleshot by sealing the leak with bone wax and performing a purge bypass. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation has been completed. the pump and purge cassette were returned for investigation. data logs were not provided for the case run. the severed sidearm was inspected and a crack along the length of the yellow luer was observed. the sidearm was connected to the purge fluid via the purge cassette and the leak was reproduced. as per the given clinical information, a leak was observed at the luer lock connection of the purge arm. the sidearm bypass was performed to resolve the leak. also, sidearm retainer was reported to be cleaned during pump use. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc., per supplier evaluation. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ d.9 date device returned/information was added. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-24009. e.1 revised first name, last name, facility name and city as they were entered incorrectly on the initial manufacturer device report number 1220648-2024-24009. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-24009 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2024-24009 and were added.
13902 | 1220648-2024-24006 | 20764633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764633&pc=OZD | 11/15/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/15/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568844 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the physician reported that the pump was placed via the left axillary using the direct approach. the physician reported a dissection of the left axillary artery, possibly related to the failed insertion. no further information was available other than plans were in place to wean the patient on (b)(6) 2024. | [5 empty] | D6b explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13903 | 1220648-2024-24011 | 20764646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764646&pc=OZD | 11/13/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/13/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. the patient's groin was oozing and systemic heparin was withheld. support continued and heparin was later started. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values; heparin was stopped to achieve hemostasis as per the clinical description. corrections to the initial mfr report: d3 should have stated danvers and not aachen. incorrect zip code reported on initial. g1 mdr reporting contact email.
13904 | 1220648-2024-24014 | 20764675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764675&pc=OZD | 11/14/2024 | 12/23/2024 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207017 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a pump support ongoing when there was a purge disc/aic failure. the aic was replaced to resolve the reported issue. the aic failure occurred during patient use but no harm or injury was noted from the interruption in the support. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | E1, e2, e3: the name and occupation of the initial reporter is unknown. the impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed.
13905 | 1220648-2024-23986 | 20764692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764692&pc=OZD | 11/01/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501264 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 73-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, vascular was consulted for distal leg perfusion. staff could not find pedal pulse on right foot. no additional information was provided regarding intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
13906 | 1220648-2024-24013 | 20764713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764713&pc=OZD | 11/14/2024 | 01/16/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530080 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that roughly five days into impella cp support, the patient experienced a drop in hemoglobin levels combined with some light bleeding from their impella access site. the dressing was changed, but the bleed persisted. due to the persistent bleed and suspicions of hemolysis, the decision was made to explant the device. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The investigation for the hemolysis and the access site bleed has been completed. data logs were reviewed finding suction occurred. no motor current spike was observed outside of the p-level changes. no positioning issue was observed. the impella was not returned for evaluation. no additional clinical details were provided. the root cause of the hemolysis issue could not be determined since the complaint device not returned for analysis and insufficient clinical data provided. the root cause of access site bleeding - major issue most likely was patient condition related due to hemophilia a. d.3 revised us city and us zip code as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25044. d.6b revised explant date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-25044. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-25044 was submitted in accordance with updated procedures. h.6 code 4641 was reported incorrectly on initial manufacturer device report number 1220648-2024-25044 and a new code was added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-25044 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm wheth er a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
13907 | 1220648-2024-24016 | 20764724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764724&pc=OZD | 11/14/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569230 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.5 device for mechanical circulatory support. while on support, low pump flows were noted. flows were 0.0 after pulling venous cannula. the physician suspects a clot was ingested. additionally, a left ventricle thrombus was noted. the pump was explanted, and the patient was stable after the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow and thrombus has been completed. the cause of the low pump flow issue was most likely improper positioning of the device per log analysis and with issue not reproducing on return product. as the necessary information was not provided, the root cause of the thrombus was not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-24016 in accordance with updated procedures.
13908 | 1220648-2024-23959 | 20764845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764845&pc=OZD | 10/31/2024 | 12/10/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562019 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support high purge pressure was observed. tissue plasminogen activator was implemented in the purge. support was continued and staff monitored. the patient expired after 234.72 hours of impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | H6 code 4582 was reported incorrectly on manufacturer device report 1220648-2024-23959. new code was added to health effect - clinical code.
13909 | 1220648-2024-23997 | 20764973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764973&pc=OZD | 11/12/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during implantation, the cp pump was unable to be delivered fully to the left ventricle, despite all troubleshooting measures and the cp was removed due to the patient¿s anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, cp was removed due to patient anatomy.
13910 | 1220648-2024-23999 | 20764985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20764985&pc=OZD | 11/12/2024 | 01/06/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, upon explant and closure, the perclose failed. the patient continued to bleed and vascular was consulted. a 16fr gore dry seal was used to obtain hemostasis. the decision was made to go to the operating room and let vascular surgically repair the vessel. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined.
13911 | 1220648-2024-23967 | 20765080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765080&pc=OZD | 11/03/2024 | 12/19/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532915 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced some oozing from their groin. heparin was withdrawn temporarily and the patient was repositioned to manage the condition. during planned explant the following day, a clot was discovered at the site. forceps were utilized to remove the clot and manual pressure was held at the site to achieve hemostasis. the impella was successfully weaned and the impella was explanted. the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into thrombus and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely anticoagulation management since groin oozing ceased with heparin stoppage. the root cause of the thrombus could not be determined since the product was not returned and insufficient clinical details were provided.
13912 | 1220648-2024-24019 | 20765168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765168&pc=OZD | 11/14/2024 | 01/21/2025 |  | 11/22/2024 |  |  | 11/15/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025511315 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support. while on support, a red alarm "purge pressure high" displayed. all purge lines were checked, and attention was drawn to the risk of a pump stop. a recommendation was made to exchange the pump but was declined. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the pump was returned for investigation. the cause of the high purge pressure issue was biomaterial, based on data log review and returned product. d9 date device returned to manufacturer added. b3 date of event was inadvertently reported incorrectly on manufacturer device report 1220648-2024-24019. g1 reporting contact email revised on manufacturer device report 1220648-2024-24019 in accordance with updated procedures.
13913 | 1220648-2024-23960 | 20765174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765174&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 10/31/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523270 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia with pulses lost in the impella leg. distal perfusion catheter was placed and the pulses returned to the impella leg. the patient¿s cardiac status was reported to have improved, resulting in the impella being removed. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13914 | 1220648-2024-23963 | 20765424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765424&pc=OZD | 11/03/2024 | 12/30/2024 |  | 11/22/2024 |  |  | 11/03/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474068 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella 5.5 for mechanical circulatory support. it was reported that during the impella 5.5 placement, the impella 5.5 was unable to pass the innominate. attempts were made with multiple wires but unsuccessful. the decision was made to place an impella cp in the right axillary. the patient was stable and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the impella 5.5 was unable to be delivered through the right axillary and the impella cp was being inserted later through the right axillary.
13915 | 1220648-2024-23994 | 20765511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765511&pc=OZD | 11/12/2024 | 11/22/2024 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  | 07/31/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387251 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, following planned shockwave therapy, the optical signal of an implanted impella cp pump became corrupted and the ao placement monitoring was lost. the doctor felt uncomfortable with continuing support without placement signal monitoring. a slow wean of the impella was performed and eventually the team were comfortable with the patient¿s improved hemodynamics to proceeded with the explant of the impella. it was reported that the patient survived explant. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined.
13916 | 1220648-2024-23996 | 20765729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765729&pc=OZD | 11/12/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542553 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma above the access site. pressure was held and hematoma was monitored. the pump was explanted (not related to the hematoma) and staff held pressure on the access site. the hematoma was expressed after the explant. the physician noted that hematoma was possibly caused from previous cath. | A hematoma above the right femoral access site was noted. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the peel away sheath was left inside till the end of the case. b.5 information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23996 and was added. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23996 and was added. h.6 health effect - clinical code 1884 was inadvertently not included on the initial manufacturer device report number 1220648-2024-23996 and was added.
13917 | 1220648-2024-23964 | 20765774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765774&pc=OZD | 11/01/2024 | 01/03/2025 |  | 11/22/2024 |  |  | 11/01/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516948 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient was on impella 5.5 support and while on support, there was a pump stop reported. the patient was mobilizing, and the pump stop occurred. the pump was explanted and replaced. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. data logs were not returned for review. returned product was investigated and the pump failed electrical testing as all three motor cable phases were open. the red pump plug was opened and it was found that all three cables were broken at the joint with the kink protector/catheter. they were necked and the glue holding all the cables together was pulled out of the kink protector. these are all signs of excessive force. clinical details report that the pump stop occurred during patient movement, which aligns with the findings on returned product. based on clinical details and returned product, the root cause of the pump stop was determined to most likely be an electrical open due to excessive force during patient transport.
13918 | 1220648-2024-23984 | 20765794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765794&pc=PYX | 11/09/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/09/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560608 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 75-year-old female patient with caridomyopathy and other systemic comorbidities was implanted with an impella rp device for mechanical circulatory support. during impella support, the purge flow began to decrease suddenly. the staff was advised on the tissue plasminogen activator protocol to manage the purge rate. it was further reported that the family elected to withdraw care and make patient comfort measures only and patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the data log showed a 3.5-hour gradual rise in purge pressure close to 1000 mmhg, and this trend was resolved after changing the purge bag/cassette, with a decreasing trend over two hours. no high purge pressure or low pure flow alarms were triggered during the purge pressure rise and purge flow decrease trend. the cause of the high purge pressure issue was most likely biomaterial, based on data log review trend on purge pressure. e.1 revised facility name as it was previously reported incorrectly on initial manufacturer device report number 1220648-2024-23984. g.1 revised reporting contact email since initial manufacturer device report 1220648-2024-23984 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report 1220648-2024-23984 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
13919 | 1220648-2024-23992 | 20765906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765906&pc=OZD | 11/10/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/10/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448019 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 81-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, impella crossed valve but in attempting to pull back into good position, impella bounced across valve and out of left ventricle. impella was removed and new pump opened. no further information is available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
13920 | 1220648-2024-24002 | 20765926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765926&pc=OZD | 11/12/2024 | 01/14/2025 |  | 11/22/2024 |  |  | 11/12/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504733 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | White | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion (hrpci) was implanted with an impella cp device for mechanical circulatory support. while on support, the long peel away sheath split from the tip and the patient was bleeding. the pump was explanted and hrpci was stopped and rescheduled. the patient was stable after the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and introducer damage has been completed. only the introducer was returned for investigation. a data log review was not required for this case. the returned introducer has a scoreline split, and there was a kink on it. according to the clinical details, the complaint was reported regarding the introducer, where the long peel-away sheath split from the tip, causing the bleeding event. the hrpci was rescheduled two days later. the cause of the access site bleeding issue was an introducer sheath split along the scorelines most likely due to patient's diseased vessels condition.
13921 | 1220648-2024-24003 | 20765935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765935&pc=OZD | 11/14/2024 | 01/02/2025 |  | 11/22/2024 |  |  | 11/14/2024 | 11/22/2024 |  |  |  |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558091 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while during impella cp support, there was an impella in aorta alarm. when the physician attempted to reposition the impella, the sterile repositioning sheath was found to be torn. the physician decided to clean catheter and then repositioned the impella. the sterile repositioning sheath was patched with tegaderms. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the pump was returned for investigation. a data log review was not required for this case. damage to the sterile sleeve was found on the returned product. no other damage was found on the returned product. as per the clinical details, the sterile sleeve was found torn during repositioning. the cause of the product damage issue was most likely use related.
13922 | 1220648-2024-24012 | 20765960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20765960&pc=OZD | 09/28/2023 | 01/08/2025 |  | 11/22/2024 |  |  | 11/13/2024 | 11/22/2024 |  |  | 06/30/2025 |  | 11/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368176 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 89 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. a protect iv study reported a patient who was on impella cp support had a fever that was possibly related to the impella procedure. after the patient's percutaneous coronary intervention (pci), they had a fever up to 38.6 ° c with an unknown origin. the patient was prescribed tazobac. the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of infection has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was reported incorrectly on manufacturer device report 1220648-2024-24012. g1 reporting contact email revised on manufacturer device report 1220648-2024-24012 in accordance with updated procedures.
13923 | 1220648-2024-23744 | 20743031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743031&pc=OZD | 09/06/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/06/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464033 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.5 device for mechanical circulatory support. during implantation, the axillary sheath was inserted without issue. the guidewire and al1 were placed. however, there was difficulty due to the patient's anatomy. the patient had a contrast shot to ensure the patency and the wire was exchanged for an 0.18 wire. the impella backloaded onto the wire and was kinked. the impella 5.5 was out of the body and flushed under heparin saline and re-crossed with al1 and guidewire. a buddy wire placed and v18 was placed into the left ventricle. after multiple failed attempts, the impella 5.5 implantation was aborted and replaced with an impella cp. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not returned for evaluation. however, information provided was sufficient to determine a root cause. the root cause of the delivery issues was due to patient condition as the impella 5.5 was unable to pass through due to patient anatomy. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13924 | 1220648-2024-23745 | 20743054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743054&pc=OZD | 10/28/2024 | 12/18/2024 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558659 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp for mechanical circulatory support and after two hours of support, the device was explanted due to access site arterial bleeding. the patient's outcome is noted as unknown. however, there is no report or indication of any further adverse events or effects to the patient. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13925 | 1220648-2024-23747 | 20743245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743245&pc=OZD | 11/06/2024 | 12/23/2024 |  | 11/21/2024 |  |  | 11/06/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522181 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock and cardiomyopathy was implanted with an impella cp for mechanical circulatory support, and during support, an optical sensor system error with the impella cp pump was experienced. despite changing the automated impella controller, the problem was not resolved. the pump was replaced. the patient was reported to be stable following this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the optical signal issue has been completed. data logs and product not returned. the cause of the optical signal issue was not determined as no product or data logs returned.
13926 | 1220648-2024-23748 | 20743277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743277&pc=OZD | 11/06/2024 | 01/02/2025 |  | 11/21/2024 |  |  | 11/06/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced hemolysis one day after start of support. the impella cp pump was repositioned to address the hemolysis. however, the physician elected to remove the pump after the hemolysis had not improved after reduction of performance level. the physician felt that the patient was stable enough to explant the pump and wait for coronary artery bypass graft surgery in two days. the patient was successfully weaned, and the pump was removed without complications. the patient was reported to be stable following the event. | [5 empty] | The investigation for the hemolysis has been completed. the product was not returned for review. data logs showed the pump ran without issue during support. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23748 was submitted in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
13927 | 1220648-2024-23749 | 20743406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743406&pc=OZD | 11/06/2024 | 01/07/2025 |  | 11/21/2024 |  |  | 11/06/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556095 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was documented that the pump was unable to cross into the left ventricle due to the tortuosity at the patient's aortic arch. the implant was aborted as a result of the resistance. there was no report or indication of any adverse event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the impella was unable to cross into the left ventricle due to patient anatomy. b.7 information was added as it was inadvertently omitted on the initial manufacturer device report number. g.1 revised reporting contact email since the initial manufacturer device report number was submitted in accordance with updated procedures. h.8 revised as it was accidently left blank on the initial manufacturer device report number and should not have been.
13928 | 1220648-2024-23750 | 20743474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20743474&pc=OZD | 11/06/2024 | 12/19/2024 |  | 11/21/2024 |  |  | 11/06/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564680 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device and experienced access site bleeding along with a concern for limb ischemia. the patient had access site bleeding and a hematoma. the groin was redressed and a stitch was placed but the hematoma persisted. additionally, the team was unable to find pedal pulses in either foot. the feet were cool and noted to be likely due to poor overall perfusion combined with peripheral artery disease. the patient was taken back to cardiac catheterization lab to re-assess the left anterior descending artery. due to concern for bleeding and possible limb ischemia, the impella and repo sheath were removed and 14 french sheath was placed in its place with hemostasis. the patient continued with extracorporeal membrane oxygenation support. the physician proceeded with the left heart catheterization and percutaneous interventions. the bleeding required eight units packed red blood cells and no blood products prior to the time of the impella cp removal. it was planned to take the patient to the cardiovascular operating room with vascular surgery for a right femoral cut-down, exploration and closure. final right lower extremity angiogram showed flow down the common femoral artery with extravasation of contrast. the patient was reported to be critical following this event. | [5 empty] | The investigation for the reported access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and limb ischemia could not be determined. revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23750 was submitted. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13929 | 1220648-2024-23751 | 20744125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744125&pc=OZD | 10/24/2024 | 12/18/2024 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546211 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 72-year-old male patient was on impella 5.5 support due to st elevated myocardial infarction. while on support, it was noted that the impella appeared malpositioned on the chest x-ray. transthoracic echocardiogram was then performed and the device was noted to be likely too deep or malpositioned in some aspect. however, the device was not re-positioned at that time. the device was removed due to the patient recovering adequate heart function. the doctor commented that coronary artery bypass graft with impella 5.5 may have contributed to aortic valvular damage/aortic insufficiency. the patient survived the explant and was stable. | [5 empty] | The failure mode positioning issue was added to address the positioning issues the customer had. positioning issues - due to lack of clinical details provided, the root cause of the positioning issues cannot be determined. heart valve damage - due to lack of product returned and insufficient clinical details provided, the root cause of the heart valve damage cannot be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23751. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
13930 | 1220648-2024-23752 | 20744224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744224&pc=PYX | 10/28/2024 | 12/10/2024 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025543813 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Pumping Stopped |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient was on impella rp flex support and during support, the pump alerted, ¿impella stopped. motor current high¿. it was followed by a ¿purge pressure high¿ alarm.the staff performed troubleshooting by looking for kinks in the purge line and disconnecting the purge arm. the pump was restarted and there were low flows with high purge alarms. the pump was explanted and a thrombus was visualized in the inlet. | [5 empty] | The investigation for the low pump flow, high purge pressure, thrombus, and pump stop has been completed. the failure mode "pump stop" was changed to "temporary pump stop" because the pump was able to be restarted. the returned pump was inspected and biomaterial was found in the purge gap beneath the impeller in the inlet. the pump was run in the lab and neither the pump stop nor the high purge pressure issues reproduced. the data logs confirm "impella stopped motor current high" alarms. the logs show a motor current spike with a speed deviation concurrent with a rise in placement signal and a drop in pump flows. there was also a rise in purge pressure. this pattern is characteristic of biomaterial ingestion. of note is the clinical mention of no anticoagulation since implant. the root cause of the temporary pump stop was determined to be biomaterial. the root cause of the high purge pressure was determined to be biomaterial. the root cause of the low pump flow issue was determined to be biomaterial. the root cause of the thrombus could not be determined without additional information. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13931 | 1220648-2024-23704 | 20744237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744237&pc=OZD | 07/04/2021 | 11/20/2024 |  | 11/21/2024 |  |  | 07/04/2021 | 11/21/2024 |  |  | 05/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1486998 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support an episode of bradycardia was noted during weaning. the pump was adjusted back up to p6 weaning was continued the next day. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13932 | 1220648-2024-23753 | 20744586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744586&pc=OZD | 10/23/2024 | 01/06/2025 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547364 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 80-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, purge fluid was leaking from the filter or pressure reservoir. the leak was small. there were no alarms for the purge flow rate or purge pressure. the pump was running until removal. the pump was removed due to hemodynamics being stable. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. the product was returned for investigation. data logs review was not required for this case. as per the clinical details, fluid was observed from the side arm without a reported purge-related alarm. doctor could not find exact location of the purge leak. no physical damage was observed on the returned product. the pump and purge cassette were primed successfully, with purge pressure/flow values within the specified range. further investigation revealed no leaking even after blocking the side arm exit for 20 minutes. the cause of the purge leak issue was most likely use related, based on returned product investigation. corrections to the initial mfr report: g1 mdr reporting contact email.
13933 | 1220648-2024-23754 | 20744747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744747&pc=OZD | 09/13/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/13/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484458 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced oozing at the impella access site. it was noted that the patient had chronic thrombocytopenia, and that bleeding was expected. the physician placed an additional suture at the impella access site to achieve hemostasis. the patient remained on impella cp support until they were successfully weaned and explanted. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was most likely use issue since adding additional suture stopped the bleeding. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13934 | 1220648-2024-23755 | 20744892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744892&pc=OZD | 09/14/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/14/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516139 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a loss of their dorsalis pedis artery pulse. there was no reported intervention for the limb ischemia. the impella was not explanted early due to the limb ischemia. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. as all the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13935 | 1220648-2024-23756 | 20744901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20744901&pc=OZD | 10/28/2024 | 12/16/2024 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528293 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 69-year-old female patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the staff were unable to doppler pulses to the right leg. the staff replaced a swan and adjusted the position to the impella 5.5. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13936 | 1220648-2024-23758 | 20745003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745003&pc=OZD | 10/27/2024 | 01/24/2025 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528748 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 65-year-old male patient was on impella cp support due to having cardiogenic shock. the impella cp was placed and flows were only 2 liters per minute in auto mode. volume was given and the pump was repositioned to no avail. the pump was explanted and no thrombus was observed in the inlet or outlet cage. the patient expired after just 1.78 hours of impella cp support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the low pump flow has been completed. no product returned. data logs show positioning alarms but clinical states pump repositioned to good position and did not resolve low issue. there are suction alarms and placement signal (ps) trending down. clinical states patient decompensating and expired on support (not impella related). no motor current (mc) spikes outside of p level changes and clinical states act good with no signs of thrombus. the pump is on auto and struggling to flow at high p levels. the flows are low from the start. the cause of the low flow was most likely patient condition related. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email. | B3 date of event was inadvertently reported incorrectly on manufacturer device report 1220648-2024-23758. d6b explant date was inadvertently reported incorrectly on manufacturer device report 1220648-2024-23758. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
13937 | 1220648-2024-23757 | 20745006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745006&pc=OZD | 09/14/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/14/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced distal pulses that were unable to be dopplered, however the medical team noted the patient's systemic vascular resistance was still high at around 4000 dynes x sec/cm5. the patient was monitored and the medical team continued to assess for pulses, and no intervention was reported. the physician noted that they did not believe the pump caused the limb ischemia. the physician noted that the patient was n profound shock when the pump was placed, and the medical team was unable to doppler pulses on any extremity because the patient was receiving high doses of pressors and the patient's shock state. the day following noted lack of distal pulses, the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the root cause of the limb ischemia cannot be determined since the product was not returned and insufficient clinical details were provided. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13938 | 1220648-2024-23759 | 20745080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745080&pc=OZD | 10/31/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 10/31/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025479583 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient with a cardiac history. the team placed the cp, but there were issues at the access site. there was bleeding and limb ischemia bilaterally. the patient was taken to the operating room for surgical repair of the site. additionally, the team gave 2 units of red blood cells. the pump remained on for support when the team and family making choices. after 46 hours the care was withdrawn and patient expired. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. h6 impact code added 1802. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
13939 | 1220648-2024-23760 | 20745256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745256&pc=OZD | 09/15/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/15/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544078 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced limb ischemia in their right lower limb. the patient also experienced hemolysis, that was persistent after repositioning of the pump. due to the patient's limb ischemia and hemolysis the decision was made to wean and explant the pump. the patient was successfully weaned and the pump was explanted and was placed on an intra-aortic balloon pump and venoarterial extracorporeal membrane oxygenation. | [5 empty] | The investigation into the limb ischemia and hemolysis is complete. the impella device was not received from the customer, therefore, investigation of the device was not possible. the root cause of the hemolysis and limb ischemia cannot be determined since product was not returned and insufficient clinical details were provided. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
13940 | 1220648-2024-23761 | 20745311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745311&pc=OZD | 10/10/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/10/2023 | 11/21/2024 |  |  | 02/28/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321874 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported they attempted to place the impella cp, but the patient's blood vessels in the lower limbs were extremely tortuous making it difficult to insert. the procedure was aborted and there was no patient harm. | [5 empty] | The investigation into the impella cp has been completed. no product was returned. per the clinical investigation, the root cause of the delivery issue was most likely due to patient condition due to patient tortuous anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13941 | 1220648-2024-23762 | 20745385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745385&pc=OZD | 10/25/2024 | 12/16/2024 |  | 11/21/2024 |  |  | 10/26/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553011 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention |  |  |  |  | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and while on support experienced a global anoxic brain injury. the patient presented to the hospital three days prior to the procedure with a non-st elevation myocardial infarction and decompensated into multi-pressor shock. the patient was taken to the catheterization lab, and the physician was unable to place an impella cp device as the physician had concern for the diseased groin vasculature. the shock team was called, and the decision was made to implant an intra-aortic balloon pump for mcs support at such time. the patient continued to decompensate and veno-arterial extracorporeal membrane oxygenation (va ecmo) mcs was started. thereafter the patient was taken to the cardiovascular operating room for and impella 5.5 device which was successfully placed. the iabp was removed, and va emco remained. the patient was transferred to the unit on ecpella (ecmo and impella) support. a following morning, two days after start of mcs, the patient was taken for a computed tomography scan and an electroencephalogram which confirmed global anoxic injury. discussion was made with the family with a decision to withdraw care. the patient passed away the next day following explant of mcs. the cause and whether the impella mcs may have cause or contributed to the anoxic brain injury are unknown. | [5 empty] | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
13942 | 1220648-2024-23763 | 20745673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745673&pc=OZD | 10/31/2024 | 12/31/2024 |  | 11/21/2024 |  |  | 10/31/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547323 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced tract bleeding and a hematoma at their impella cp access site during support. manual pressure and a pressure dressing were applied to manage the condition, and one unit of packed red blood cells was given to the patient. around the same time, a thrombus was noted in the left ventricle via echocardiogram. the decision made to remove the impella device and an intra-aortic balloon pump was placed for ongoing support. the patient's status was later updated to expired. it is unknown if the complications caused or contributed to the patient's passing. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombus could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
13943 | 1220648-2024-23764 | 20745909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20745909&pc=OZD | 11/03/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 11/04/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552635 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and after the pump removal, a thrombus was noted in the right iliac. a 6 french fogerty catheter used and thrombus removed. arteriotomy closed and incision closed. doppler pulses present post closure. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
13944 | 1220648-2024-23765 | 20746029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746029&pc=OZD | 04/12/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 04/12/2022 | 11/21/2024 |  |  | 06/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022142687 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support and during impella closure, one of the two placed perclose sutures broke. manual pressure was held to achieve hemostasis and the device was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13945 | 1220648-2024-23766 | 20746068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746068&pc=OZD | 04/22/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 04/22/2022 | 11/21/2024 |  |  | 10/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173323 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the impella 2.5 was properly weaned and explanted, however the closure method was site preclosed with proglidex2 and manual pressure held. post angiogram done from r radial and slight oozing was noticed. the physician decided to do percutaneous transluminal angioplasty (pta) with 7mmx40mm balloon, and the device was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13946 | 1220648-2024-23767 | 20746172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746172&pc=OZD | 04/25/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 04/25/2022 | 11/21/2024 |  |  | 06/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022114873 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during impella closure, the perclose failed and manual pressure was held for one hour to manage the complication. the device was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13947 | 1220648-2024-23768 | 20746227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746227&pc=OZD | 09/15/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/15/2024 | 11/21/2024 |  |  | 07/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380939 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported while on (b)(6) 2024 while impella cp support, the patient experienced bleeding at the impella access site. the patient received 3 units of red blood cells and a femstop was applied to the site and heparin was withheld. the following day, (b)(6) 2024, the patient was still requiring the femstop at the impella access site to prevent profuse arterial bleeding. the patient received three additional units of red blood cells and heparin continued to be withheld until hemostasis was achieved. on (b)(6) 2024, the patient was noted to have cold bilateral lower extremities and they were noted to have petechiae and mottling. it was also reported that the patient had several episodes of ventricular tachycardia overnight that required defibrillation and an increase of pressor support to resolve. on (b)(6) 2024 the patient was experiencing self-limiting ventricular tachycardia and was intermittently having episodes of atrial fibrillation and sinus bradycardia. amiodarone and lidocaine were administered to the patient to treat the arrhythmias. it was noted that the patient had a known hematoma on the anterior wall of the myocardium which was thought to be the cause of the arrhythmias. the patient was critically ill, and expired on support later that day. | [5 empty] | The investigation into the reported issues has been completed. the impella device was discarded by the customer, therefore, investigation of the device was not possible. as all the necessary information was not provided, the root causes of the limb ischemia, access site bleeding, and cardiac arrhythmias were not able to be determined. should any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13948 | 1220648-2024-23769 | 20746275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746275&pc=OZD | 04/29/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 04/29/2022 | 11/21/2024 |  |  | 05/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022090459 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during impella closure, the perclose failed. oozing was observed and a femstop was placed for thirty minutes to manage the complication. the device was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13949 | 1220648-2024-23770 | 20746322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746322&pc=OZD | 09/15/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/15/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479926 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported while on impella 5.5 support, the patient experienced a gastrointestinal bleed and black/tarry stool. heparin was withheld and the patient received a unit of packed red blood cells. support continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | Product and data files were not returned for review. the root cause of gastrointestinal bleed was unable to be determined as limited clinical details were provided and no product was returned for review. should the device or any new information be received, a supplemental mdr will be filed.
13950 | 1220648-2024-23771 | 20746329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746329&pc=OZD | 11/13/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 11/13/2022 | 11/21/2024 |  |  | 10/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170303 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the purge system was blocked, and the purge pressure increased, and the purge line was removed and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13951 | 1220648-2024-23772 | 20746390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746390&pc=OZD | 09/16/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/16/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532969 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella was weaned and removed, the perclose devices were tightened but the closure site did not achieve hemostasis. a balloon tamponade of the site entry was deployed, protamine was given and femostop was placed to achieve hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. due to lack of product return and sufficient clinical information, the root cause of the access site bleeding cannot be determined. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13952 | 1220648-2024-23773 | 20746516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746516&pc=OZD | 11/14/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 11/14/2022 | 11/21/2024 |  |  | 11/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210783 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced an episode of ventricular tachycardia. the patient was cardioverted, and the right coronary artery was stented to manage the complication. the patient needed more hemodynamic support, and the decision was made to escalate from an impella 2.5 to a impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13953 | 1220648-2024-23774 | 20746664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746664&pc=OZD | 01/11/2022 | 09/24/2025 |  | 11/21/2024 |  |  | 01/11/2022 | 11/21/2024 |  |  | 09/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151098 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support the patient experienced bleeding of the access site. the amount of bleeding and actions to take to resolve the bleeding are unknown as the complainant did not provide any further information. additional information was provided that care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined g1 reporting contact email revised on manufacturer device report 1220648-2024-23774 in accordance with updated procedures.
13954 | 1220648-2024-23775 | 20746966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20746966&pc=OZD | 07/06/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 07/06/2022 | 11/21/2024 |  |  | 03/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210612 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced a gastrointestinal bleed. heparin was withheld. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
13955 | 1220648-2024-23777 | 20747158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747158&pc=OZD | 09/26/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 08/13/2024 | 11/21/2024 |  |  | 04/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351149 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the physician operator feels the axillary artery was just too small to accommodate the impella cp. the operators were unable to advance the impella cp beyond the anastomosis given multiple tries with viperslide lubrication, multiple torquing attempts, multiple adjustments to graft angle, etc. the implantation of the impella cp was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned, and the issue was confirmed. visual inspection found no issues on the pump, gw & red lumen. per the results of the clinical review and investigation, the root cause was determined to be most likely use related. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13956 | 1220648-2024-23779 | 20747168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747168&pc=OZD | 10/10/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 08/13/2024 | 11/21/2024 |  |  | 08/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390766 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.5 device for mechanical circulatory support. during implantation, the medical staff was having difficulty delivering an impella 5.5 pump due to the patient¿s vascular anatomy. after several failed attempts they decided to abort impella 5.5 and keep the patient on impella cp support instead. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine the root cause, which is most likely patient condition related. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13957 | 1220648-2024-23776 | 20747195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747195&pc=OZD | 10/03/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/03/2023 | 11/21/2024 |  |  | 04/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337957 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient scheduled for a coronary artery bypass graft was being implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the medical team had difficulty delivering the impella pump due to the patient's vascular anatomy issue. after failed attempts the procedure was aborted. there was no harm or patient injury reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely due to patient condition, the pump could not pass through due to vascular anatomy issue. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
13958 | 1220648-2024-23780 | 20747214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747214&pc=OZD | 02/04/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 02/04/2022 | 11/21/2024 |  |  | 10/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2021019999 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient experienced hemolysis and the device was repositioned. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13959 | 1220648-2024-23778 | 20747267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747267&pc=OZD | 08/14/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 08/14/2022 | 11/21/2024 |  |  | 04/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support ct bleeding was at 30 milliliters per hour. heparin in purge was put on hold by surgeon. the patient continued on support until bridged to left ventricular assist device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13960 | 1220648-2024-23781 | 20747299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747299&pc=OZD | 01/28/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 01/28/2022 | 11/21/2024 |  |  | 11/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED INC. |  |  | US |  | 2022148724 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced access site bleeding. the resolution for this issue was reported as surgically hemostasis (use of hemostatic agent). survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported bleeding issue has been completed. the impella device was not received from the customer. the clinical details provided were sufficient to determine a root cause. the cause of the bleeding was a patient condition related resulting in an arterial bleed around the guidewire repositioning unit sheath or an arteriotomy hematoma. best practices were followed, and no use related issue was reported. includes large abdominal pannus resulting in the repo unit unable to reach the vessel. this report is being filed as part of a retrospective review of historical records.
13961 | 1220648-2024-23784 | 20747365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747365&pc=OZD | 01/24/2022 | 10/07/2025 |  | 11/21/2024 |  |  | 01/24/2022 | 11/21/2024 |  |  | 06/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098182 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | Asian | Death | Yes | Yes | Fracture | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.5 for mechanical circulatory support. the devices sidearm broke, the device was explanted and replaced with another impella 5.5. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Complainant reported the patient was on impella cp support due to the primary indication of ami/csg (acute myocardial infarction/cardiogenic shock). while on support, on (b)(6) 2022, it was reported that the purge sidearm broke on original impella 5.5. the decision was made today to swap the 5.5 out for a new 5.5 device. care was withdrawn and the patient expired. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
13962 | 1220648-2024-23786 | 20747379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747379&pc=PYX | 08/09/2024 | 03/06/2025 |  | 11/21/2024 |  |  | 08/09/2024 | 11/21/2024 |  |  | 08/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024392313 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing treatment for a pulmonary embolism was attempted to be implanted with an impella rp device for mechanical circulatory support. during the implantation, the physician was unable to advance the pump past the pulmonary valve after multiple attempts. the impella was removed, and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. rp smartassist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-23786 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-23786 in accordance with updated procedures.
13963 | 1220648-2024-23783 | 20747412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747412&pc=OZD | 09/29/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 09/29/2023 | 11/21/2024 |  |  | 08/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390223 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with previous heart transplant experiencing graft rejection and coded was being implanted with an impella cp device for mechanical circulatory support. it was reported upon inserting the impella, there was a previous graft site on aortic arch that physician could not advance the impella through ,despite multiple attempts. the decision was made to abort the case and send patient to cardiovascular operating room for an axillary impella placement. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue is most likely patient condition, upon inserting impella there was a previous graft site on aortic arch that physician could not advance impella through despite multiple attempts as per the clinical description. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
13964 | 1220648-2024-23787 | 20747425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747425&pc=OZD | 01/22/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 01/22/2022 | 11/21/2024 |  |  | 09/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022139383 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient received an electrical shock due to fibrillation, and it was reported that the device was exhibiting suction alarms indicating low flow and no flow. an echocardiogram was performed and confirmed that the device was placed correctly. no further information is available. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, no abnormalities were observed, and the pump was run in a basin of water and a high motor current initially occurred. upon restart, biomaterial kicked out of the pump and the pump operated to specification with an average flow of 5.2 l/min. therefore, the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
13965 | 1220648-2024-23782 | 20747511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747511&pc=PYX | 10/24/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 08/07/2024 | 11/21/2024 |  |  | 01/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024329037 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient scheduled for post-operative left ventricular assist device (lvad) placement was attempted to be implanted with an impella rp for mechanical circulatory support. during delivery of the impella rp, the team had difficulty maintaining the position above the pulmonary valve due to the patient's large right ventricle, and the procedure was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was most likely patient condition related. rp smartassist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
13966 | 1220648-2024-23788 | 20747521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747521&pc=OZD | 01/21/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 01/21/2022 | 11/21/2024 |  |  | 07/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121652 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, it was discovered that the patient suffered a major stroke. it was reported that the procedure was successful, however additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13967 | 1220648-2024-23785 | 20747588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747588&pc=OZD | 10/20/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/25/2023 | 11/21/2024 |  |  | 09/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024401066 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity in post cardiotomy cardiogenic shock / low cardiac output syndrome was attempted to be implanted with an impella 5.5 for mechanical circulatory support. during implantation, the pump could not pass the aortic arch due to the patient's abnormal anatomy. after several unsuccessful attempts, the case was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was most likely patient condition related; the pump could not pass the aortic arch due to abnormal anatomy as per the clinical description provided. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
13968 | 1220648-2024-23791 | 20747649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747649&pc=OZD | 10/10/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/10/2023 | 11/21/2024 |  |  | 08/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393084 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaint reported the impella 5.5 was not able to traverse into ascending aorta into left ventricle. during the process the wire came out of left ventricle. multiple attempts were made over several hours in the operating room. implant was aborted and there was no reported patient harm. | [5 empty] | The investigation into impella 5.5 has been completed. the impella device was discarded by the customer. per the clinical investigation, the root cause of the delivery anatomy issue was most likely patient condition related. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13969 | 1220648-2024-23790 | 20747650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747650&pc=OZD | 01/13/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 01/13/2022 | 11/21/2024 |  |  | 06/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022142227 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella ld for mechanical circulatory support. during support, the device exhibited a purge leak from the hub, along with blood slowing leaking from between the clear side arm and red impella plug. the staff performed a purge reservoir bypass which resolved the problem. the patient continued on support, and it was reported that they survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13970 | 1220648-2024-23789 | 20747659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747659&pc=OZD | 09/09/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/09/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532960 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as a high-risk percutaneous coronary intervention was needing an impella cp device for mechanical circulatory support. while trying to advance the impella through the 14 x 25 centimeter sheath, the motor housing was unable to make the final turn into the aorta. after multiple attempts, the impella cp was aborted and the case proceeded without the impella cp. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. the impella was unable to make the final turn into aorta as per the clinical description provided.
13971 | 1220648-2024-23792 | 20747735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747735&pc=OZD | 10/10/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/10/2023 | 11/21/2024 |  |  | 08/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388732 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella 5.5 could not pass through a calcified shelf in subclavian. crossing was unsuccessful despite multiple attempts and impella 5.5 placement was aborted. there was no known harm or injury to the patient. | [5 empty] | The investigation into the impella 5.5 has been completed. no product was returned. per the clinical investigation, the root cause of the delivery issue was most likely patient condition related, the impella could not pass through a calcified shelf in subclavian as per the clinical description provided. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13972 | 1220648-2024-23794 | 20747745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747745&pc=OZD | 11/13/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/13/2021 | 11/21/2024 |  |  | 07/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121616 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the patient was positive for heparin-induced thrombocytopenia (hit), the resolution for this issue is unknown. the patient was bleeding internally, and heparin was removed from the purge. additionally, the device was exhibiting low purge flows and high pressures. the medical staff did not want to give tissue plasminogen activator (tpa) due to bleeding risk. it was reported that the patient was made do not resuscitate, after impella 5.0 support of 190.33 hours, care was withdrawn and the patient expired thereafter. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13973 | 1220648-2024-23795 | 20747862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747862&pc=OZD | 10/20/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/23/2023 | 11/21/2024 |  |  | 01/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024309845 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. during placement, the impella was met with significant resistance, so it was removed and replaced with another impella device. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided the root cause of the delivery issue was patient condition related due to vessels stenosis and %50 lesion. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13974 | 1220648-2024-23793 | 20747864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747864&pc=OZD | 11/15/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/15/2021 | 11/21/2024 |  |  | 06/30/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1502902 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the device exhibited a purge leak, and the staff replaced the purge cassette. the device exhibited high purge pressure alarm after the cassette was replaced and the purge cassette again leaked. the purge cassette was replaced again, the purge flow rate continues to rise, and the purge pressure continues to fall. the console was replaced; however, the alarm did not resolve. the resolution for this device is unknown. additional information noted that patient outcome was death due to respiratory failure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned and evaluated, and the investigation for the reported high purge pressure has been completed. the device was returned, and the leak was confirmed. adding additional dextrose to the purge, the viscosity of the purge fluid increased therefore increasing the purge pressure. the root cause of the first purge cassette leak was found to be a piston cylinder leak/insufficient piston/cylinder seal. this report is being filed as part of a retrospective review of historical records.
13975 | 1220648-2024-23796 | 20747939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747939&pc=OZD | 11/11/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/11/2021 | 11/21/2024 |  |  | 10/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022026616 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella ld for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced bleeding complications due to an extensive procedure. it was reported that the patient's aorta was repatched, and the chest was open post operation. it was stated that seven units of red blood cells were provided, and heparin was removed from the purge in an effort to manage the complication. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13976 | 1220648-2024-23799 | 20747964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747964&pc=OZD | 10/07/2022 | 02/13/2025 |  | 11/21/2024 |  |  | 10/07/2022 | 11/21/2024 |  |  | 06/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247503 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a nurse on orientation attempted a cassette change. the nurse initially damaged the first cassette by improperly placing the cassette into the automated impella controller and damaging the blue tubing on the cassette. immediately, purge fluid began leaking out of the cassette door. the same nurse attempted to remove the damaged cassette and broke the blue plastic piece off the air sensor. a backup automated impella controller was required. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. no product was received for evaluation. the cause of the purge leak was most likely a damaged valve caused by improper use technique. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
13977 | 1220648-2024-23798 | 20747990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747990&pc=OZD | 11/08/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/08/2021 | 11/21/2024 |  |  | 07/31/2021 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1442186 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. the complainant reported that during impella placement, there were recurring problems with the red curve. the curve disappeared and a significant decrease in calculated flow of the pump was observed. it was reported that the calculated flow did not coordinate with the hemodynamic of the patient or echocardiogram. the issue was resolved by recalibration of the pump. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13978 | 1220648-2024-23805 | 20747998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20747998&pc=OZD | 09/10/2024 | 07/18/2025 |  | 11/21/2024 |  |  | 09/10/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534812 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was needing an impella 5.5 device for mechanical circulatory support. while on support, the patient was noted to have hemolysis. the patient was given blood products. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Pt was noted to have hemolysis while on cp support. after escalating to 5.5, the hemolysis was noted to have continued. after 6 days on support, it was noted the hemolysis remained present. it is unknown if the hemolysis resolved. the pump was not returned. data log review showed no abnormalities. no suction alarms, mc spike, or ps trends. the cause of the hemolysis was not determined since product was not returned, no abnormalities shown in the data logs, and insufficient clinical details.
13979 | 1220648-2024-23801 | 20748056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748056&pc=OZD | 08/12/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/12/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461087 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing elective coronary artery bypass graft and valve placement was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. during impella implantation, several attempts were made with different wires, catheters, and an un-wired impella pump, but the physician was unable to pass the device into the proper position. the impella placement was aborted, and the physician proceeded with chest closure and placement of an intra-aortic balloon pump. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through the vasculature per the clinical description provided. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
13980 | 1220648-2024-23802 | 20748071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748071&pc=OZD | 11/04/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/04/2021 | 11/21/2024 |  |  | 04/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078642 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be implanted with an impella 5.0 for mechanical circulatory support. during delivery, the motor would not pass the patient's brachiocephalic trunk area. the device was removed and replaced by an impella cp. additionally, it was reported that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition, specifically the patient's small vasculature. this report is being filed as part of a retrospective review of historical records.
13981 | 1220648-2024-23803 | 20748074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748074&pc=OZD | 07/17/2024 | 07/14/2025 |  | 11/21/2024 |  |  | 07/19/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487664 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had ventricular tachycardia. the pump was repositioned and support continued. additionally, the patient had renal failure on support. it is unknown if the impella contributed to the renal failure. furthermore, the automated impella controller (aic) had a suction alarm. the investigation for console determined there was likely a pump issue. the investigating engineer determined that the pump should also be investigated for any product malfunctions. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Vt/vf/at/af: in order to make a root cause determination, all of the clinical details outlined in (b)(4) would have to be provided. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. renal failure: the cause of the renal failure was not determined due to insufficient clinical details. low pump flow: no product was returned for analysis. the data logs show flows below expected range for p4 on 7/17. there were no motor current spikes or positioning issues. the low pump flow resolved and returned to expected range for p4. there were suction alarms which resolved. limited clinical information was provided. the root cause of the low pump flow was not determined as no product was returned for analysis, no data log patterns or trends were identified and limited clinical information was provided.
13982 | 1220648-2024-23797 | 20748089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748089&pc=OZD | 10/20/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/25/2023 | 11/21/2024 |  |  | 04/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371336 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity in post cardiotomy cardiogenic shock / low cardiac output syndrome was attempted to be implanted with an impella cp for mechanical circulatory support. during implantation, the staff had difficulty with wiring/accessing the left ventricle through the patient's bioprosthetic aortic valve. the procedure was aborted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was most likely patient condition related. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
13983 | 1220648-2024-23804 | 20748117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748117&pc=OZD | 11/03/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/03/2021 | 11/21/2024 |  |  | 12/31/2021 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1463041 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Material Puncture/Hole | Device Damaged by Another Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the guidewire pierced the pump, another 5.0 pump was used, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. the root cause of the delivery issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
13984 | 1220648-2024-23806 | 20748134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748134&pc=OZD | 11/02/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/02/2021 | 11/21/2024 |  |  | 07/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024958 |  | 005066 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported high purge pressure and a purge cassette leak during support. there was no adverse impact to the patient and support continued with this pump. no additional information provided. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. a root cause for the high purge pressure could not be determined through this evaluation as no product or data logs were provided. this report is being filed as part of a retrospective review of historical records.
13985 | 1220648-2024-23800 | 20748155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748155&pc=OZD | 11/05/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 11/05/2021 | 11/21/2024 |  |  | 12/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021016284 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. the complainant reported a pump stop occurred during use of the device. the pump stop occurred after 14 days, which is outside the instructions for use intended use life. no report of patient harm or injury. | [5 empty] | Explant date unknown. the investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. this report is being filed as part of a retrospective review of historical records.
13986 | 1220648-2024-23807 | 20748160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748160&pc=OZD | 07/25/2024 | 09/28/2025 |  | 11/21/2024 |  |  | 07/25/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474692 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the impella exhibited high purge pressure and low purge flow, which was addressed by adding purge lysis to the purge. the patient remained on impella support, and there was no reported patient harm. | The complainant also reported that the pump triggered several yellows 'placement signal unreliable' alarms, with no placement signal (ps) visible. however, the issue resolved on its own during a call with the representative. the impella was not explanted as a result of the event. |  |  |  |  | The investigation for the reported high purge pressure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. per log review purge pressure rise with motor current spike and increase in flows seen. high purge pressure and purge system blocked alarms were found. later gradual resolving of purge pressure observed after tissue plasminogen activator administration. the cause of the high purge pressure issue was most likely biomaterial ingestion. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. impella pump was not returned for investigation. upon review of data logs it is apparent that the console is the potential cause of the optical signal issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. g1 manufacturing site name and address corrected.
13987 | 1220648-2024-23808 | 20748280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748280&pc=OZD | 11/01/2024 | 01/08/2025 |  | 11/21/2024 |  |  | 11/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508756 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced limb ischemia. the physician considered performing a femoral to femoral bypass. however, the vascular surgeon reported that the leg was still sufficiently perfused. therefore, a femoral to femoral bypass was not necessary. the patient was monitored and the leg was deemed sufficiently perfused and did not require intervention. support was continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
13988 | 1220648-2024-23810 | 20748597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748597&pc=PYX | 11/09/2024 | 12/31/2024 |  | 11/21/2024 |  |  | 11/09/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551607 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient with right heart failure and other systemic comorbidities was implanted with an impella rp flex device for mechanical circulatory support. while on support, there was a sudden drop in pump flow from 10.9 to 3.5. volume was given. the pump was removed due to low pump flow. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the data logs confirmed the failure mode and clinical narrative. based on clinical description and data review, the root cause of low flow was most likely biomaterial ingestion. g1 reporting contact email revised on manufacturer device report 1220648-2024-23810 in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
13989 | 1220648-2024-23809 | 20748631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748631&pc=OZD | 01/28/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 01/28/2022 | 11/21/2024 |  |  | 12/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023162538 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced thrombocytopenia and was treated with sodium bicarbonate. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
13990 | 1220648-2024-23811 | 20748793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20748793&pc=OZD | 10/24/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 10/30/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553867 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the pump was supporting for a week when the pump was removed in the operating room for left ventricular assist device delivery. the nurse observed valve damage and new aortic insufficiency. the valve was replaced. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the heart valve damage has been completed. product was not returned for investigation. data logs were investigated and there were no positioning alarms that would suggest the pump was improperly positioned near the valve. clinical details do not provide any evidence that would explain why the heart valve damage occurred. the medical team is unsure if the issue was present before impella insertion. due to limited clinical details and product not being returned, the root cause of the heart valve damage could not be determined.
13991 | 1220648-2024-23812 | 20749297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749297&pc=OZD | 09/16/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/16/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515254 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient was escalated to the impella 5.5 from an impella cp. it was reported while on impella 5.5 support, the patient experienced an episode of ventricular tachycardia [vt]. patient was shocked once and cardiopulmonary resuscitation was administered and return of spontaneous circulation [rosc] was achieved. the impella position was verified with echocardiography. support continued with the implanted pump following the intervention. the patient was successfully weaned and explanted. | [5 empty] | The investigation into the arrythmia issue is complete. the impella device was not received from the customer, therefore, investigation of the device was not possible. as the necessary information was not provided, the root cause of the ventricular tachycardia was not determined. should the device or any new information be received, a supplemental mdr will be filed.
13992 | 1220648-2024-23814 | 20749828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749828&pc=OZD | 10/18/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 08/13/2024 | 11/21/2024 |  |  | 09/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396490 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.5 device for mechanical circulatory support. during implantation, the placement was unsuccessful due to the patient having axillary stenosis. the insertion for the impella 5.5 was aborted and a impella cp was placed. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related due to patient anatomy as angio revealed acute almost parallel to the floor innominate artery before it came to almost a 90 degree angle off the carotid and continued on to the arch distally and left axillary stenosis was noted too. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13993 | 1220648-2024-23815 | 20749886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749886&pc=OZD | 08/13/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/13/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482870 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, multiple echocardiograms revealed that the impella was slightly deep, and it was noted that it may have been interacting with the leaflets on the calcified aortic valve. the device was then explanted and replaced with another pump. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
13994 | 1220648-2024-23816 | 20749910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749910&pc=OZD | 08/18/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/18/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025526490 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the access site. the patient was successfully weaned and the impella was explanted, and the site was secured. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
13995 | 1220648-2024-23813 | 20749959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749959&pc=OZD | 09/20/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 09/20/2023 | 11/21/2024 |  |  | 11/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294523 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was not able to be inserted due to the patient¿s anatomy and the case was aborted. there was no report of patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely due to patient condition. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions: never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
13996 | 1220648-2024-23817 | 20749986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20749986&pc=OZD | 10/06/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/06/2023 | 11/21/2024 |  |  | 11/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196793 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. it was reported the complainant reported that impella met resistance in the stenosed area of the iliac artery. after multiple attempts to pass the stenosis the case was aborted and the 2.5 was explanted through the peel away sheath. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related due to the stenosis in iliac.
13997 | 1220648-2024-23818 | 20750018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750018&pc=OZD | 08/20/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/20/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468637 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was attempted to be implanted with an impella 5.5 for mechanical circulatory support. during placement, the physician met resistance at the axillary artery and was unable to pass through the vasculature. the physician elected to abort the impella implantation, and no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition since insertion of pump was aborted due to patient's anatomy even after trying multiple times. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
13998 | 1220648-2024-23819 | 20750104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750104&pc=OZD | 03/08/2022 | 03/10/2025 |  | 11/21/2024 |  |  | 03/08/2022 | 11/21/2024 |  |  | 10/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023161429 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding. heparin was being withheld. additional information noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported vascular damage - peripheral vessel issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
13999 | 1220648-2024-23820 | 20750199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750199&pc=OZD | 04/05/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 04/05/2021 | 11/21/2024 |  |  | 10/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2021017101 |  | 0048-0034-JP |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the patient experienced a hematoma and access site bleeding. the staff administered a blood transfusion. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14000 | 1220648-2024-23822 | 20750333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750333&pc=OZD | 04/08/2021 | 03/17/2025 |  | 11/21/2024 |  |  | 04/08/2021 | 11/21/2024 |  |  | 01/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED. INC. |  |  | US |  | 2022049231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 83 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/08/2021 |  |  |  |  | Us complainant reported the patient was on impella cp for hemodynamic support. while on support patient was taken to the operating room due to ischemic leg on impella. cp was explanted via vascular surgery. patient expired on support. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. section b2, b5, h6 was updated to reflect the correct information.
14001 | 1220648-2024-23821 | 20750394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750394&pc=OZD | 03/25/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 03/25/2022 | 11/21/2024 |  |  | 12/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023176880 |  | 0550-0008 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding from all sites and bilateral hematomas. there was a loss of three units of blood. actions taken to resolve are unknown. additional information noted that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
14002 | 1220648-2024-23823 | 20750599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750599&pc=OZD | 03/25/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 03/25/2022 | 11/21/2024 |  |  | 12/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175367 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced runs of ventricular tachycardia during support. the patient was cardioverted and received chest compressions. additional information noted that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the issue could not be determined as insufficient clinical details were provided.
14003 | 1220648-2024-23826 | 20750825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750825&pc=OZD | 10/26/2024 | 12/24/2024 |  | 11/21/2024 |  |  | 10/26/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025562023 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient has many cardiac comorbidities. after the placement of cp, upon transferring the patient from the cath lab table to the bed, impella was accidentally pulled into the aorta. physician attempted to recross the aortic valve but was unsuccessful. wire was placed through the reaccess port and impella cp was then removed and physician did not want to replace with the same pump. a new pump was placed without any issue. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issues most likely was patient movement related. g1 reporting contact email revised on manufacturer device report 1220648-2024-23826 in accordance with updated procedures.
14004 | 1220648-2024-23827 | 20750997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20750997&pc=PYX | 10/27/2024 | 12/20/2024 |  | 11/21/2024 |  |  | 10/27/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025448722 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with a history of mixed cardiomyopathy was implanted with an impella rp for mechanical circulatory support, and three days after start of support, the patient developed bleeding from the access site, mouth, and nose. bivalirudin was withdrawn in response to the bleed, and support remained ongoing following the intervention. however, three days thereafter, the impella rp had a "dramatic" decrease in flows. the pulmonary artery pressure was increased on the automated impella controller, and the flows became less than 1.0. discussion was made regarding possible clot entrainment (no information at hand confirming however) as patient was off anticoagulation. the decision was made to explant the device which was successfully completed. the patient's outcome was noted as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation into access site bleeding, vascular damage - peripheral vessel, and low pump flow has been completed. based on limited clinical details and product not being returned, the root cause of the vascular damage and access site bleeding could not be determined. product was not returned for investigation. data logs were investigated, and as reported, the low pump flows were confirmed on 10/30. leading up to the decrease in flows there are two clear motor current spikes and speed deviations without any change in p-level. this is followed by a step-wise increase in purge pressure up to 800 mmhg and drop in purge flows without triggering any purge alarms. for these reasons, the failure mode was changed from low pump flows to placement signal issues. based on data logs, the root cause of the placement signal issues was determined to most likely be dp sensor drift.
14005 | 1220648-2024-23829 | 20751071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751071&pc=OZD | 10/28/2024 | 02/07/2025 |  | 11/21/2024 |  |  | 10/28/2024 | 11/21/2024 |  |  | 12/31/2022 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022027305 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant in 2021 had a cp support for a patient admitted in with a stemi. the patient was enrolled into the stemi dtu study. the support allowed for the pci and was explanted after 11.5 hours. later in 2024 the adverse event was documented that a week later after the procedure in which the team states there was left ventricle, thrombus developed with a "possible" relationship to the impella device and procedure. the thrombus was treated with medication and was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." corrections to the initial mfr report: g1 mdr reporting contact email.
14006 | 1220648-2024-23825 | 20751100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751100&pc=OZD | 03/29/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 03/29/2022 | 11/21/2024 |  |  | 11/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023164800 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient had an episode of ventricular fibrillation requiring successful cardioversion once. the bedside registered nurse had impella position placement questions. the patient had ventricular fibrillation again. the patient was given cardiopulmonary resuscitation (cpr) and shocked three times. they were taken to the cardiac catheterization laboratory for left heart catheterization, but no abnormalities were seen. the patient was returned to the unit. an episode of ventricular tachycardia occurred which required cardioversion. the manufacturer reviewed recommendations for pump management during cardioversion or cpr with the staff. there was no extracorporeal membrane oxygenation as the patient was on continuous renal replacement therapy. additional information noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
14007 | 1220648-2024-23830 | 20751138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751138&pc=OZD | 07/06/2022 | 05/22/2025 |  | 11/21/2024 |  |  | 07/06/2022 | 11/21/2024 |  |  | 03/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206568 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient implanted with an impella cp device for mechanical circulatory support experienced a ventricular fibrilation (vf) / ventricular tachycardia (vt) issue. during patient support, it was reported that the patient developed an ischemic event followed by vf/vt. the patient received defibrillation six times and adrenaline and amiodarone intravenously. an external pacemaker was inserted but there was no output. the decision was made to stop the procedure and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14008 | 1220648-2024-23831 | 20751185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751185&pc=OZD | 11/01/2024 | 01/02/2025 |  | 11/21/2024 |  |  | 11/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515598 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.5 device for mechanical circulatory support. while on support, there was a ¿purge system blocked¿ alarm. there were no kinks detected, and the motor current was still in range. protocol given to apply the lysis and allow to run. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d6b explant date reported incorrectly on manufacturer device report 1220648-2024-23831. g1 reporting contact email revised on manufacturer device report 1220648-2024-23831 in accordance with updated procedures. h10 instructions for use missing on manufacturer device report 1220648-2024-23831.
14009 | 1220648-2024-23828 | 20751262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751262&pc=OZD | 04/08/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 04/08/2022 | 11/21/2024 |  |  | 01/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183583 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the pump experienced suction alarms that were resolved by giving volume. additionally, the patient experienced gastrointestinal bleeding as well as access site bleeding with no resolution specified in either case. the patient also experienced limb ischemia that was resolved by surgical repair of the right femoral artery, but with no resolution specified for the left side. the patient expired. additional information noted that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14010 | 1220648-2024-23824 | 20751374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751374&pc=OZD | 06/03/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 06/03/2021 | 11/21/2024 |  |  | 01/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022052313 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2021 |  |  |  |  | The user facility reported a male in cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. it was reported the patient experienced clotting in the abdominal aorta and lower limbs. due to the patient's circumstances the decision was made to withdraw care and eventually the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | B2: the date of patient death is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of this issue and its relation to the impella device was not determined since the product, data logs, and sufficient clinical information were not provided.
14011 | 1220648-2024-23834 | 20751495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751495&pc=OZD | 11/15/2024 | 01/16/2025 |  | 11/21/2024 |  |  | 11/15/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557686 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported after 15 days on impella cp support, the pump experienced a motor current high red alarm and the pump then suddenly stopped. the medical team attempted to restart the pump three times without success. the pump was also disconnected and reconnected to the automated impella controller, but this also did not resolve the problem. the pump was explanted and the patient remained on extracorporeal membrane oxygenation. there was no reported patient harm or injury. | [5 empty] | The investigation for the pump stop has been completed. no product or logs were returned for evaluation. clinical details noted that the pump experienced a pump stop after an increase in motor current. the root cause of the pump stop could not be definitively determined as no product or logs were returned for evaluation. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23834 was submitted in accordance with updated procedure. h.6 codes 2199 and 4115 were reported incorrectly on the initial manufacturer device report number 1220648-2024-23834. it was unknown if the device was discarded. the device was not returned. new codes were added. h.10 additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-23834 stated that the device was discarded but the device was not returned. type of investigation codes has been updated to reflect not returned. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-23834 in accordance with updated procedures. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
14012 | 1220648-2024-23833 | 20751555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751555&pc=OZD | 11/05/2024 | 10/20/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1251201 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A user facility reported that following impella cp implant, no placement signal waveform appeared on the automated impella controller (aic). placement signal not reliable alarms subsequently appeared coinciding with controller failure alarms. the aic was replaced and no further issues were noted. there was no patient harm. | [5 empty] | The impella device has not been returned by the customer for evaluation. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results.
14013 | 1220648-2024-23839 | 20751600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751600&pc=OZD | 11/02/2024 | 08/14/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1849988 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Display/Image | Display or Visual Feedback Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a user facility hospital team/biomed found a console with a screen power up issue. despite multiple troubleshooting steps, the screen would not illuminate. the controller itself was verified to be powered on via illuminated buttons and a green light above the grey selector wheel. there was no patient involvement or harm at the time of discovery of the noted issue. | [5 empty] | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. | A.1 added as information was inadvertently omitted from the initial manufacturer device report number: 1220648-2024-23839. a.2 age should have been left blank on the initial manufacturer device report number: 1220648-2024-23839 as no patient was involved. d.2 revised common device name as it was entered incorrectly on the initial manufacturer device report number: 1220648-2024-23839. d.4 revised model number as it was entered incorrectly on the initial manufacturer device report number: 1220648-2024-23839. e.1 phone number was reported incorrectly on initial manufacturer device report number: 1220648-2024-23839. e.2 revised as it was entered incorrectly on the initial manufacturer device report number: 1220648-2024-23839. g.1 revised mdr reporting contact email as it was entered incorrectly on the initial manufacturer device report number: 1220648-2024-23839. h.3 revised to reflect device evaluation anticipated as previous selections were entered incorrectly on the initial manufacturer device report number: 1220648-2024-23839. h.6 codes 3221 and 2199 were reported incorrectly on the initial manufacturer device report number: 1220648-2024-23839. a new code was added.
14014 | 1220648-2024-23835 | 20751649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751649&pc=OZD | 08/20/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/20/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519976 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient undergoing an electrophysiology (ep)/ablation procedure was attempted to be implanted with an impella cp for mechanical circulatory support. during implantation, the physician could not get the catheter and wire to advance across the valve, and they ultimately decided to abort the impella placement. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because the catheter was unable to cross the valve due to anatomy of the patient and the impella implantation was aborted. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14015 | 1220648-2024-23836 | 20751658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751658&pc=OZD | 08/23/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/23/2024 | 11/21/2024 |  |  | 01/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024313915 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with a pulmonary embolism was implanted with an impella rp device for mechanical circulatory support. during support, bleeding was noted at the insertion site. this was addressed with mattress suture x2 and an additional purse stitch. there was no further reported harm, and the patient remained on support and was successfully weaned. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was most likely use related, mattress suture attempted x2 and then modified purse string to achieve hemostasis as per the clinical description provided. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
14016 | 1220648-2024-23837 | 20751734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751734&pc=OZD | 11/05/2024 | 01/08/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025535354 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was implanted with an impella cp pump for mechanical circulatory support. upon attempted placement of a second perclose following implant, difficulty was noted and an arterial tear was suspected. placement of the peel-away sheath temporarily resolved the bleed, however up upon its eventual removal, a hematoma was noted coinciding with active bleeding around the sheath. attempts to tighten the perclose failed and the decision was made to explant the pump. a third perclose was placed following explant to achieve hemostasis. | [5 empty] | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was returned for evaluation. no issues were found with the returned product. clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since staff attempted to deploy 2 perclose without success. corrections to the initial mfr report: d4 updated model number. g1 mdr reporting contact email.
14017 | 1220648-2024-23838 | 20751747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751747&pc=OZD | 11/05/2024 | 01/03/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025534117 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance through the subclavian due to calcium in the patient's vasculature. the implant was aborted as a result of the resistance. the patient remained supported via ecmo and balloon pump. there was no patient harm. | [5 empty] | The impella device has not been returned for evaluation. upon conclusion of the investigation, a supplemental report will be submitted. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, unable to pass pump past subclavian due to vascular anatomy and calcium. the following have been reported incorrectly on manufacturer device report 1220648-2024-23838. d4 model number. g1 reporting contact email revised in accordance with updated procedures.
14018 | 1220648-2024-23840 | 20751783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751783&pc=OZD | 11/05/2024 | 01/06/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025512540 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had a faint pedal pulse with doppler in their right lower extremity (rle). the patient lost perfusion to the rle later that night prompting discussions on the next course of action. based off hemodynamics, labs, and impella values, it appeared that the patient needed to be escalated to impella 5.5 or ecmo support. a decision was made by the family to withdraw care. the impella was turned off and the patient passed away shortly after. it is unknown if the condition contributed to the patient's passing. | [5 empty] | The impella device was not returned by the customer for evaluation. upon conclusion of the investigation, a supplemental report will be submitted. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23840. d.4 revised model number as it was reported incorrectly on the initial manufacturer device report number 1220648-2024-23840.
14019 | 1220648-2024-23842 | 20751924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20751924&pc=OZD | 10/12/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 10/12/2023 | 11/21/2024 |  |  | 05/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365522 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp insertion was unsuccessful. the patient¿s vessel was calcified and the impella could not cross the iliac. the impella implant was aborted, there was no patient injury reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related due to vessels tortuosity and calcification. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
14020 | 1220648-2024-23841 | 20752012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20752012&pc=OZD | 07/19/2024 | 02/26/2025 |  | 11/21/2024 |  |  | 08/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Renal Failure | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported on (b)(6) the patient was on impella 5.5 support and during support, the patient developed arrythmia. the staff debated pulling back impella since it has been deep since placement, but suction self-resolved and patient pressures somewhat stabilized with slight increase in pressor support. later on october 16th a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening of aki on ckd with a "possible" relationship to the impella device used: ¿aki on ckd. unable to assess uremia as subject is sedated. will start cvvhd today with fluid removal as tolerated¿. it was reported that the patient/event has not recovered/not resolved. loqi registry also reported that the patient endured sinus tachycardia with a "possible" relationship to the impella device used: ¿pt now sinus tach 105-115 bpm. pt's map dipping to mid 50s. unable to remove volume via crrt as pt's blood pressure decreases with attempts. pt's hr decreased to 70s and crt-d failed to pace (no pacing spikes noted when this happened). pa (b)(6) made aware. pads placed on pt. norepinephrine started per pa. chattering noted on ecmo cannulas. albumin 5% x1 ordered and administered. map still noted to be low 60s. said pa aware. - 1l ns ordered and administered. 2 units prbcs ordered and administered. the second infusing.¿. the patient recovered with sequelae. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the renal failure, positioning issue, and the arrhythmia has been completed. no product was returned for analysis. the data log analysis cannot determine purge leak - cassette location. the root cause of the purge leak - cassette was not determined as no product was returned and limited clinical information was provided. the root cause of the positioning issues was not determined as no product was returned and limited clinical information was provided. the root cause of the renal failure and arrhythmia issues were not determined due to insufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email.
14021 | 1220648-2024-23843 | 20752121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20752121&pc=OZD | 09/14/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/14/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504154 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support, the patient experienced a massive ischemic stroke which was confirmed with head computed tomography. the patient suffered a loss of neurological function due to the stroke, and as a result of the patient's neurological status it was decided to withdraw care. the patient's care was withdrawn and the patient expired. | [5 empty] | The investigation for the reported stroke/cva has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. clinical information was not provided, and therefore the root cause of the stroke/cva could not be determined. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14022 | 1220648-2024-23844 | 20752400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20752400&pc=OZD | 09/14/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/14/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525914 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported while on impella 5.5 support, the patient experienced ventricular fibrillation which required two rounds of defibrillation to resolve. support continued with the implanted pump following the intervention. the patient was successfully weaned and the impella explanted. | [5 empty] | The investigation for the reported v-fib has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14023 | 1220648-2024-23845 | 20752994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20752994&pc=OZD | 06/09/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 06/09/2021 | 11/21/2024 |  |  | 03/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022063638 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella cp for mechanical circulatory support. it was reported that the delivery of the impella cp was unsuccessful due to the patient's vasculature. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
14024 | 1220648-2024-23846 | 20753264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753264&pc=OZD | 11/05/2024 | 12/27/2024 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025566738 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient presenting for acute myocardial infarction and cardiogenic shock was scheduled for impella cp implant for mechanical circulatory support. the femoral anatomy was seen to have peripheral artery disease with calcification and narrowing. a 14 fr introducer was placed but the doctor admitted that a incision was made in the sheath and it dissected. the groin had bleeding that they tried to remedy with 2 perclose and pressure, but the site eventually needed surgical intervention to achieve hemostasis. | [5 empty] | The impella device was not returned by the customer for evaluation. upon conclusion of the investigation, a supplemental report will be submitted. per the impella ifu: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution:: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage has been completed. no product was returned for analysis. clinical details had no reports of excessive force. the root cause of the vascular damage - peripheral vessel was not determined as the introducer was not returned for analysis and limited clinical information related to failure was provided. added h6 component code 925 corrections to the initial mfr report: d4 updated model number.
14025 | 1220648-2024-23848 | 20753539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753539&pc=OZD | 11/05/2024 | 08/27/2025 |  | 11/21/2024 |  |  | 11/05/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC |  |  | US |  | 1090689 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A user facility biomed reported that during preventative maintenance, an automated impella controller powered up with a self check failure. there was no patient involvement or harm associated with the failure. | [5 empty] | The device was returned and an investigation of the device is ongoing. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. | Additional information: the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause for the system self-check failure was not determined due to the intermittent failure mode. corrections to the initial mfr report: d2 device name has been updated. updated h3 to device evaluation anticipated. h6 type of investigation and conclusion code added.
14026 | 1220648-2024-23849 | 20753576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753576&pc=OZD | 09/19/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/19/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535310 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion, the device was prevented from proper insertion due to calcification in the ascending aorta. the impella was then removed, and the implant was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because the patient's vessels were calcified preventing to advance the impella. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action.
14027 | 1220648-2024-23847 | 20753619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753619&pc=OZD | 08/19/2024 | 07/07/2025 |  | 11/21/2024 |  |  | 08/19/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471284 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the device produced impella in aorta alarms, but echocardiograms confirmed that the pump was in the correct position. staff also noted that placement monitoring was disabled at the time. the staff continued to monitor for any changes. it was further reported that the patient was made a dnr and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Fm changed from positioning issues to optical signal issue because pump position was confirmed at the time of the false position in aorta alarm. no imc log available but the aic software version is likely >v11.1. data logs reviewed and the optical metrics were confirmed to be stable throughout the run. no mc spike at the time of the reported complaint false alarm issue. the cause of the optical signal issue was a false pump in aorta alarm patient condition related since pump position was confirmed to be good.
14028 | 1220648-2024-23853 | 20753700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753700&pc=OZD | 08/22/2024 | 08/19/2025 |  | 11/21/2024 |  |  | 08/22/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483422 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in preparation for high-risk surgery was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient was found to be septic and have a gastrointestinal bleed. no additional information regarding these issues was provided, but the patient ultimately expired while on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of infection/sepsis and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
14029 | 1220648-2024-23854 | 20753703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753703&pc=OZD | 11/07/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 11/07/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692221 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during preparation for impella implant, the automated impella controller (aic) had a purge disc issue. the aic was replaced successfully. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added d9 device return date.
14030 | 1220648-2024-23850 | 20753799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753799&pc=OZD | 11/07/2024 | 01/03/2025 |  | 11/21/2024 |  |  | 11/07/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505856 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. complications were encountered during attempted delivery. following insertion, it was documented that the pump was unable to advance beyond through the subclavian artery due to the patient's anatomy. during the fourth failed attempt, the patient had two episodes of ventricular tachycardia requiring cardioversion. the decision was subsequently made to abort the implant. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues and arrhythmia has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the pump was unable to advance beyond through the subclavian artery due to the patient's anatomy. the root cause of the arrhythmia could not be determined without additional clinical details. b.2 revised as not all the selections on the initial manufacturer device report number 1220648-2024-23850 were correct. d.6a and d.6b dates were revised as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23850. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23850 was submitted in accordance with updated procedures.
14031 | 1220648-2024-23859 | 20753865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753865&pc=OZD | 10/30/2024 | 12/13/2024 |  | 11/21/2024 |  |  | 10/30/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025566204 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Tachycardia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support the patient experienced atrial fibrillation/supraventricular tachycardia (svt) with rapid ventricular response. cardioversion was required. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the arrhythmia could not be determined.
14032 | 1220648-2024-23860 | 20753893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753893&pc=OZD | 05/24/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 05/24/2023 | 11/21/2024 |  |  | 12/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324648 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during introducer insertion, it was reported that the distal tip of long impella sheath had become kinked from aortic tortuosity/spinal stenosis surgery. impella would not pass kink. the sheath was exchanged for short sheath and long 7fr destination was used for guide support. impella was successfully implanted and there was no reported patient harm. the patient was successfully weaned and explanted. | [5 empty] | The investigation for the reported introducer damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the introducer damage issue was patient condition related to aortic tortuosity.
14033 | 1220648-2024-23857 | 20753897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753897&pc=OZD | 10/30/2024 | 11/19/2025 |  | 11/21/2024 |  |  | 10/30/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532372 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. following insertion, resistance was met while removing the guidewire. after manipulating the devices, the guidewire was successfully removed; however, it was discovered that the catheter was kinked at the aortic arch. the pump was replaced as a result of the noted damage. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14034 | 1220648-2024-23851 | 20753910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753910&pc=OZD | 11/07/2024 | 01/13/2025 |  | 11/21/2024 |  |  | 11/07/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560399 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during the implant, the pump failed to pass the aortic valve (av). the impella became too flexible overtime started to bend. the implant was aborted and a new impella was used successfully. no patient harm or injury was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the pump was returned for investigation and guide wire was noted to be kinked at multiple locations. the root cause of delivery issue is determined to be use issue because the second pump was placed without any issues.
14035 | 1220648-2024-23862 | 20753961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753961&pc=OZD | 11/23/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 11/23/2022 | 11/21/2024 |  |  | 08/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269640 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support it was noted that the patient had ventricular tachycardia and required defibrillation. in addition, vascular surgery was consulted for surgical repair of right femoral artery. chest and sternal closure was performed. impella was removed, and site repaired. impella ld was placed to wean from bypass. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14036 | 1220648-2024-23856 | 20753971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753971&pc=OZD | 05/19/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 05/19/2023 | 11/21/2024 |  |  | 11/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312049 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that impella cp with femoral access successfully placed, procedure completed without incidence. after the removal of 14f sheath, the pre-placed perclose device failed. manual compression held , but patient required 1 unit prbcs and protamine administration. hemostasis then achieved with manual compression, then femostop placed. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was discarded by the customer, however, clinical details were sufficient to determine the root cause. the root cause of the access site bleeding issue was lack of vessel recoil due to patient condition related as the patient had calcified vessels condition. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14037 | 1220648-2024-23861 | 20753988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20753988&pc=OZD | 10/30/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 10/30/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed for a patient with multiple cardiac and systemic comorbidities. the patient was in need of a high-risk percutaneous cardiac insertion (hrpci) with impella support. the team had difficulty with delivering but, the sheath and pump were placed. upon explant a vascular dissection was observed in the right femoral artery (rfa). team performed balloon tamponade and placed covered stents. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the vascular damage has been completed. no product was returned for evaluation. clinical details noted that patient had hereditary hemorrhagic telangiectasia (hht) that caused abnormally formed blood vessels. upon insertion of introducer, the physician noted difficulty when advancing the sheath but was able to eventually place sheath. after percutaneous coronary intervention procedure, an angiogram showed dissection in the right femoral artery, but no vascular repair was needed. as the customer reported, the vascular damage against the 14 french introducer, the product investigated was changed to the 14 french introducer. the root cause of the vascular damage - peripheral vessel was most likely patient condition related as the patient had hht. b.7 information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23861 and was added. h.5 revised as it should have been yes on the initial manufacturer device report number 1220648-2024-23861.
14038 | 1220648-2024-23863 | 20754009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754009&pc=OZD | 05/05/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 05/05/2023 | 11/21/2024 |  |  | 11/30/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during prep time the impella readied, but it was contaminated. a new device was prepped and inserted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14039 | 1220648-2024-23864 | 20754018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754018&pc=OZD | 05/27/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 05/28/2023 | 11/21/2024 |  |  | 03/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327665 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported since the first day of insertion, the impella had purge system blocked and high purge pressure alarms. the medical team ran tissue plasminogen activator in the purge solution but reported there was no change to purge flow or purge pressure. it was advised by abiomed support staff to change the purge cassette as part of troubleshooting. the plan was to wean and explant the patient. the patient remained on impella 5.5 support for 5 more days before successfully being weaned and explanted. there was no information provided how the purge issue was resolved. there was no harm to patient reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the high purge pressure could not be determined. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14040 | 1220648-2024-23865 | 20754050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754050&pc=OZD | 06/07/2023 | 11/21/2024 |  | 11/21/2024 |  |  | 06/07/2023 | 11/21/2024 |  |  | 02/28/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331662 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the clear ring on the 14fr dilator did not come off the peel away sheath following insertion of the impella catheter. the sheath was subsequently peeled away, and the repositioning sheath was placed through the disc and sutured into place. there was no patient harm resulting from the noted issue. support continued with the implanted pump following the issue, and was successfully weaned and explanted. | [5 empty] | The investigation for the reported introducer damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the introducer damage issue was not determined since product was not returned for investigation.
14041 | 1220648-2024-23852 | 20754053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754053&pc=OZD | 09/28/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 09/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510288 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 had high purge pressure and purge system blocked alarms. there were no kinks observed, the purge cassette was changed which did not resolve the issue, and there was no thrombus observed with imagining. the abiomed support team suggested the medical team to perform a tissue plasminogen activator procedure. after the lysis procedure was complete, the purge alarms resolved. there was no harm to the patient. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the high purge pressure was most likely due to biomaterial as it was able to be resolved with tissue plasminogen activator. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14042 | 1220648-2024-23855 | 20754104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754104&pc=OZD | 03/25/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 10/30/2024 | 11/21/2024 |  |  | 02/28/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an unknown primary indication was implanted with an impella cp device for mechanical circulatory support. during implantation and support, the patient exhibited ventricular tachycardia that was attributed to the impella, which required defibrillation. the patient remained in sinus tachycardia, but they stabilized and ultimately survived the procedure. | [5 empty] | The investigation into the reported issue has been completed. no product was returned. after reviewing the clinical information, the cause of the ventricular fibrillation and its relation to impella were unable to be determined.
14043 | 1220648-2024-23858 | 20754121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754121&pc=OZD | 07/01/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 07/01/2021 | 11/21/2024 |  |  | 02/28/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2022064812 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis and multiple organ failure (cerebral hemorrhage), the patient expired from multiple organ failure. no further information was provided. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14044 | 1220648-2024-23867 | 20754142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754142&pc=OZD | 08/25/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/25/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523833 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, an echocardiogram showed a thrombus in the left ventricle. the impella was removed and the physician opted for a different mechanical circulatory support (mcs) strategy. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. to determine the true root cause of the thrombus, the following clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined" thrombus: "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth.
14045 | 1220648-2024-23866 | 20754184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754184&pc=OZD | 10/31/2024 | 12/23/2024 |  | 11/21/2024 |  |  | 11/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555889 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. while on support the patient developed a hematoma. a fem stop was placed in the icu. the patient received two units of blood prophylactically and was then taken to the cath lab for impella removal. the surgeon attempted to tamponade the access site from the contralateral side. an attempt was made at post closure, and it was unsuccessful. a 14 french sheath was subsequently placed into the artery and the patient will go to the operation room for hematoma, evacuation, and site closure. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23866: b3 date of event. g1 reporting contact email revised from the initial submission in accordance with updated procedures. h6 health effect - impact code 4627 was missing.
14046 | 1220648-2024-23868 | 20754189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754189&pc=PYX | 08/26/2024 | 06/30/2025 |  | 11/21/2024 |  |  | 08/26/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025458200 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing post-operative cardiothoracic surgery was implanted with an impella rp flex device for mechanical circulatory support. during support, the patient had access site bleeding, which was resolved with an additional suture. in addition, the device's motor current temporarily tripled, and pump flows decreased by 1l/sec, suggesting clot ingestion. the team then explanted the pump, and after explant, the patient started on dobutamine and is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation of access site bleeding, low pump flow, and thrombus has been completed. the impella device was not returned for evaluation. access site bleeding - major: clinical noted the access site began oozing and required a suture. no additional information related to how bleeding occurred was provided. the root cause of the access site bleeding - major was not determined as insufficient clinical information was provided low pump flow: clinical mentioned reduced pump flows with suspect clot ingestion. the data logs show a motor current spike on (b)(6) with speed deviation and a shift in placement signal. there was also a drop in flow at instant of motor current spike. the identified pattern is characteristic of an ingestion. the root cause of the low pump flow was most likely biomaterial as evidenced by ingestion pattern observed in the data logs thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined".
14047 | 1220648-2024-23869 | 20754226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754226&pc=OZD | 08/27/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/27/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528695 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was a hematoma in the right femoral artery. heparin was put on hold, the hematoma was compressed out, and no further bleeding was noted. there was no further reported harm, and the patient survived. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14048 | 1220648-2024-23870 | 20754235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754235&pc=OZD | 08/28/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513982 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient developed hemolysis and a gastrointestinal (gi) bleed and was given a unit of packed red blood cells (prbc). the hemolysis continued, and the patient was given lasix as well as more prbc and platelets. the patient was also found to have a new gi bleed, and heparin was discontinued. however, the patient's urine output later cleared, and the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14049 | 1220648-2024-23871 | 20754238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754238&pc=OZD | 08/28/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/28/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520019 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed an infection at the access site due to a graft. the patient was to be started on antibiotics, and the impella was weaned and explanted. there was no further reported harm, and the patient survived. | [5 empty] | The investigation for the reported infection/sepsis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14050 | 1220648-2024-23872 | 20754242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754242&pc=OZD | 11/01/2024 | 12/26/2024 |  | 11/21/2024 |  |  | 11/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545133 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. the complainant reported that an impella cp was being inserted and was advanced over the wire into the left ventricle (lv) without difficulty. the wire was removed and impella started but unable to achieve flow greater than two. position on fluoroscopy appeared correct, but kink noted in cannula just distal to motor housing/outflow. the impella was pulled back into the aorta and kink was noted to be present. a new impella was successfully placed. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the product damage kinked cannula has been completed. data logs were reviewed which confirmed low pump flow when pump started and remained to the rest of the case. visual inspection found a kink on cannula about 15 mm from outflow cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely kinked cannula. added h6 impact code 4629 corrections to the initial mfr report: g1 mdr reporting contact email.
14051 | 1220648-2024-23873 | 20754248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754248&pc=OZD | 08/28/2024 | 11/21/2024 |  | 11/21/2024 |  |  | 08/28/2024 | 11/21/2024 |  |  | 07/31/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379825 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was found to have limb ischemia and was taken for right leg fasciotomy, with placement of a reproduction sheath in the right femoral artery. the patient ultimately expired, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14052 | 1220648-2024-23874 | 20754253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754253&pc=OZD | 11/01/2024 | 12/30/2024 |  | 11/21/2024 |  |  | 11/01/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555632 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.5 device for mechanical circulatory support. while on support, there was a hematoma at the access site. the patient was brought to the operating room and a washout was performed. in addition, the patient experienced atrial fibrillation. the p level was reduced to p7 due to the atrial fibrillation. the patient remained on impella support and was successfully weaned. | Additional information received. the patient condition deteriorated and the decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 patient outcome updated to death. b5 updated with new information. d6b added. h1 type of reportable event updated. h6 health effect - impact code 1802 added g1 reporting contact email revised on manufacturer device report 1220648-2024-23874 in accordance with updated procedures.
14053 | 1220648-2024-23875 | 20754295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754295&pc=OZD | 08/28/2024 | 02/14/2025 |  | 11/21/2024 |  |  | 08/28/2024 | 11/21/2024 |  |  | 11/30/2025 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient developed oozing and hematoma at the access site. a saline bag and pressure were applied over the site, and the patient was not given systemic anticoagulants. in addition, the patient had suspected hemolysis. the pump's p-level was lowered and the patient was started on nipride, and the issue resolved. the patient remained on impella support and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and access site bleed has been completed. no product returned. the cause of the access site bleeding was not determined due to lack of clinical information and no product returned. data logs show brief suction alarm that resolves near end of case. no mc spikes outside p level changes. no other abnormalities. the cause of the hemolysis was not determined since product was not returned and lack of clinical details.
14054 | 1220648-2024-23876 | 20754307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754307&pc=OZD | 07/26/2021 | 11/21/2024 |  | 11/21/2024 |  |  | 07/26/2021 | 11/21/2024 |  |  | 05/31/2023 |  | 11/21/2024 |  | LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED. INC. |  |  | US |  | 2022088118 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | Ectopic Heartbeat | [10 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.5 for mechanical circulatory support. during support, the patient experienced atrial fibrillation, rapid ventricular response, ectopy, and ventricular fibrillation. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14055 | 1220648-2024-23877 | 20754326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754326&pc=OZD | 12/17/2021 | 06/02/2025 |  | 11/21/2024 |  |  | 12/17/2021 | 11/21/2024 |  |  | 03/31/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064518 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 for hemodynamic support. while on support an unknown source of abdominal bleeding was noted. patient received ten units of blood. it was also noted that flows were only able to be increased to p6 due to suction. patient expired on support. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14056 | 1220648-2024-23878 | 20754339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754339&pc=OZD | 02/05/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 02/05/2022 | 11/21/2024 |  |  | 11/30/2023 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding from the venous line adjacent to the reposition sheath. the resolution for this issue is unknown. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14057 | 1220648-2024-23879 | 20754390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754390&pc=OZD | 08/27/2024 | 10/23/2025 |  | 11/21/2024 |  |  | 08/31/2024 | 11/21/2024 |  |  |  |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025506827 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | Peripheral Nervous Injury | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was limb numbness/tingling and skin discoloration therefore the pump was explanted. upon the explant, was thrombosis seen and the team performed thrombectomy. the patient remained stable and was successfully weaned. | [5 empty] | The investigation for the reported optical signal issue, peripheral nerve injury and thrombus issue has been completed. the pump was returned and evaluated. no abnormalities were found with the 5s parameters. the pump also passed the laser continuity test for sensor functionality. the impeller was covered in biomaterial and there was biomaterial near the optical sensor. log review showed an increase in placement signal for two hours along with some snr and contrast deviation during that time. given the biomaterial found on the pump and the placement signal abnormalities in the logs being intermittent, the root cause of the optical signal issue was most likely biomaterial. per the results of the clinical review and investigation, the root cause of the optical signal issue was most likely biomaterial. as the necessary information was not provided, the root cause of the thrombus was not determined.
14058 | 1220648-2024-23880 | 20754466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20754466&pc=OZD | 07/17/2022 | 11/21/2024 |  | 11/21/2024 |  |  | 07/17/2022 | 11/21/2024 |  |  | 01/31/2024 |  | 11/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186170 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Death | Yes | No | Positioning Failure | Restricted Flow rate | Increase in Pressure |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support the patient was trached and there was noted bleeding from trach site. heparin was removed from purge and restarted a few days later. it was noted that there were concerns for sepsis. cultures were sent and continuous veno-venous hemodialysis (cvvhd) fluid removal was stopped. in addition, the patient had a gastrointestinal bleed, and systemic heparin was stopped, and purge heparin was decreased. the patient underwent esophagogastroduodenoscopy (egd) and was noted hemoglobin was stabilized and no further episodes of melena; however, the egd showed no active bleeding but "friable tissue with a slow ooze so the patient received blood products. in addition. the impella was alarming "purge pressure high" which was resolved after they checked for kinks and changed the purge cassette. the patient was also started on antibiotics for a positive sputum culture. it was also noted an episode of suction and transesophageal echocardiography confirmed impella position. on echocardiogram the impella was orientated away from apex. the ct surgery was unable to torque at bedside, however the suction alarms were resolved with volume. the patient was on impella support for 80+ days; however, care was withdrawn, and the patient expired. | [5 empty] | The impella automated controller device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14059 | 1220648-2024-23606 | 20733519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20733519&pc=OZD | 11/03/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/03/2021 | 11/20/2024 |  |  | 09/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022022701 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the impella flows had dropped and there was suction. an echocardiogram that was performed that same evening, showed the position of device pointing away from the apex and within the mitral apparatus. however, the flow did not increase even with positioning. in addition, clots were suspected. the impella 5.0 was removed and replaced with an impella 5.5. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the product was returned, and the issue was confirmed. the field long term data log was reviewed and a gradual decrease in pump flow was observed on 03nov2021. additionally, numerous suction alarms were captured throughout the case. the returned pump was visually inspected, and blood/clotting was observed at the pump outflow. clinical details indicated that the pump was pointing away from the apex and was within the mitral apparatus. additionally, the patient had a previous embolectomy despite heparin. per the results of the clinical review and investigation, the root cause was determined to be most likely ingested biomaterial initially due to poor pump positioning. this report is being filed as part of a retrospective review of historical records.
14060 | 1220648-2024-23607 | 20733631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20733631&pc=OZD | 11/01/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/01/2021 | 11/20/2024 |  |  | 11/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022030228 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation and after several unsuccessful attempts to cross the stenotic aortic valve, the surgeon decided to abort placing the impella 5.0. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14061 | 1220648-2024-23608 | 20733713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20733713&pc=OZD | 09/09/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/09/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533187 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, access site bleeding issue was noted. during end of the case, impella was explanted and perclose suture was used, however, hemostasis was unable to be achieved during removal of the device. physician additionally had to use 18fr angioseal and hemostasis was then achieved. patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding - major has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by using 8fr angioseal.
14062 | 1220648-2024-23609 | 20733786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20733786&pc=OZD | 09/09/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/09/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508796 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, a hematoma at access site was noted. the hematoma at insertion site caused pain and discomfort. the patient was going to interventional radiology (ir) to assess surgical site hematoma. no additional information was provided. patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding - major has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding - major could not be determined.
14063 | 1220648-2024-23610 | 20733871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20733871&pc=OZD | 11/08/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/08/2021 | 11/20/2024 |  |  | 07/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024922 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support the patient experienced thrombocytopenia from insertion to removal. in addition, the patient had access site bleeding that required a blood transfusion. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14064 | 1220648-2024-23611 | 20734072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734072&pc=OZD | 12/02/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 12/02/2021 | 11/20/2024 |  |  | 09/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504585 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported an impella 5.0 was started and placed on p2 with very minimal flow. attempts were made to give volume, but unsuccessful. the decision was then made to place an rp to assist flow. it was reported that the patient expired after 5.73 hours of being on impella 5.0 support with no allegations made toward the impella 5.0 as the cause of death. however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14065 | 1220648-2024-23614 | 20734158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734158&pc=OZD | 11/07/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/07/2021 | 11/20/2024 |  |  | 06/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022109807 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient was oozing from the extracorporeal membrane oxygenation (emco) groin site. the patient received 3 units of packed red blood cells, 2 cryo and 2 units of platelets. it was also reported that the patient had significant upper and lower gastrointestinal bleeds. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14066 | 1220648-2024-23615 | 20734185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734185&pc=OZD | 11/29/2021 | 05/22/2025 |  | 11/20/2024 |  |  | 11/29/2021 | 11/20/2024 |  |  | 11/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022039181 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 2.5 for mechanical circulatory support. during support, the patient was bleeding around the foley and from the mouth. pt occasionally have runs of vt with coughing , doctor repositioned. a gastrointestinal bleed was also identified. heparin was discontinued. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. corrected information in sections a5, b2, b7, e1, g1, h1, and h2. | The date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14067 | 1220648-2024-23617 | 20734186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734186&pc=PYX | 10/22/2024 | 03/03/2025 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025548032 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | Hemolysis | Tachycardia | Thrombosis/Thrombus | [9 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | It was further reported via abiomed's long-term outcome and quality indicator impella registry (loqi) that three other adverse events (ae) occurred. the team reported upon hemolysis with definite relationship to the use of the impella and the thrombosis. loqi also documented the prior reported stroke (cva) with an unknown relationship to the use of the impella. the final ae was ventricular tachycardia (vt) with a possible relationship to the use of impella. | The complainant had a impella rp flex pump supporting a patient admitted in with history of a coronary artery bypass grafting and extracorporeal membrane oxygenation post bypass removal. the rp flex was placed without issue but, after 5 days the purge flow/pressure was alarming and did prompt the treatment with tissue plasminogen activator in the purge. the pump remains on for support despite the malfunction. later it was reported that the pump was explanted due to clot ingestion. it was also reported that the patient had a stroke and only moving right side. the family decided to withdraw all care and the patient expired. |  |  |  |  | Added-b5-additional information provided from loqi with 2 additional failure modes. h6 clinical code 1886, 2095. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure, thrombus, and the stroke has been completed. no product/logs were returned. the root cause of the high purge pressure, thrombus, and the stroke was not determined since no product/logs were returned and insufficient clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14068 | 1220648-2024-23616 | 20734192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734192&pc=OZD | 10/22/2024 | 12/24/2024 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525459 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the insertion of an impella 5.5 device was unsuccessful due to the patient's anatomy. the device was unable to go farther than the axillary to innominate artery, and the case was aborted. the outcome of the patient was indicated as unknown. however, there was no report or indication adverse effects to the patient. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23616 d6a and d6b missing e4 should have been left blank
14069 | 1220648-2024-23612 | 20734225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734225&pc=OZD | 11/04/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/04/2021 | 11/20/2024 |  |  | 04/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078642 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. during implantation the device would not advance due to the patient¿s vasculature. the impella 5.0 was removed and replaced with an impella cp. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14070 | 1220648-2024-23613 | 20734228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734228&pc=OZD | 12/01/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 12/01/2021 | 11/20/2024 |  |  | 07/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123386 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the clinical team reported that there was a "purge system failure" that eventually led to replacing the impella with a balloon pump. no further information was provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the purge issue was most likely misuse of the placement signal lumen, resulting in blood ingress due to lack of purge. the y-connector was improperly connected to the placement signal lumen. this report is being filed as part of a retrospective review of historical records.
14071 | 1220648-2024-23619 | 20734233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734233&pc=PYX | 10/22/2024 | 12/17/2024 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025537031 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support and during the swan placement, there was a spike in the motor current and a drop in flows. it was found that a clot was ingested. the pump was removed. the patient survived the event. | [5 empty] | The investigation into a low pump flow and thrombus has been completed. the returned product showed biomaterial in the inlet and on the impeller. there were no other noted physical abnormalities with the pump. the data logs show a motor current spike followed by a drop in flow levels from about 3.1l/min to 1.5l/min. the root cause of the low pump flow was determined to be due to biomaterial. as the necessary information was not provided, the root cause of the thrombus was not determined the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23619. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing. h6 code 4614 and 4641 were reported incorrectly. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
14072 | 1220648-2024-23621 | 20734244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734244&pc=OZD | 10/22/2024 | 12/20/2024 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539539 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 had purge pressure low alarm. there was a leak around the side arm leur lock connection. the team wants to manage the 5.5 leak instead of replacing pump. dextrose purge solution was increased. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the purge leak issues has been completed. no product was returned. log review showed multiple purge cassette/bag changes on the day of reported leak . the "purge pressure low" alarms were during or following these changes. the dextrose solution was increased as well, and the low purge flow resolved about 5 hours later. later in the case, the purge flow began to rise, and purge pressure began to decrease again slightly. a leak was observed in the field around the sidearm luer lock connection. the cause of the purge issue was most likely a damaged sidearm but the exact location of the leak was unknown since product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
14073 | 1220648-2024-23622 | 20734246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734246&pc=OZD | 10/14/2024 | 12/20/2024 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on impella cp support the patient experienced hematoma on the right groin with a "possible" relationship to the impella cp device. vessel closure and hemostasis were complicated due to heavily calcified vessels. the patient's left groin had a collagen plug that was half deployed but hemostatic. the right groin was repaired with suture, however a small hematoma developed with persistent bleeding. on arrival to the unit, manual pressure was held to the right groin for 15 minutes with hemostasis achieved and a pressure dressing was applied. the patient recovered with no sequelae. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient condition because hemostasis was complicated due to heavily calcified vessels. the following have been reported incorrectly on manufacturer device report 1220648-2024-23622. d4 model number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
14074 | 1220648-2024-23627 | 20734314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734314&pc=OZD | 10/24/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552615 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and during the entire case there were low purge pressure alarms. to address the issue, the team swapped to d10 without resolution and then swapped purge cassette still without resolution. the team also checked the line for loose connections or kinks. the case proceeded without incident with no significant changes in motor current or patient hemodynamics. the pci was successfully completed. the impella was weaned and explanted in the procedural area. there was no report of adverse effects to the patient. | [5 empty] | The product was not returned; therefore, an evaluation, analysis of the device was not possible. however, data logs were reviewed. low purge pressure was reported throughout support and purge cassette change did not resolve the issue. data logs confirm the low purge pressure alarms. there were no motor current or placement signal spikes or trends. product history review was completed, and the pump passed all sterile inspection checks. there are purge flow tick issues as well as variable purge pressure and flows. however, the cause of the purge leak pump was not determined as the product was not returned.
14075 | 1220648-2024-23629 | 20734333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734333&pc=OZD | 11/24/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/24/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024109 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. during implantation, it was not possible to place the pump in the left ventricle due to the vascular condition of the patient. the procedure was aborted and was implanted with an impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14076 | 1220648-2024-23625 | 20734370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734370&pc=OZD | 10/24/2024 | 01/02/2025 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536840 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced hemolysis one day after start of support. no treatment was given for hemolysis other than repositioning the pump in which there was "slight" improvement. dialysis was not performed. there was a deterioration in renal function (unknown value), however, it was determined that dialysis was not necessary. mcs continued for an additional five days. the patient was successfully weaned, and the impella cp pump was explanted with a survival outcome. | [5 empty] | The investigation for the hemolysis has been completed. data logs were reviewed finding many suctions occurred mostly in the three days after the case initiated. drop in motor current (mc) and flows were observed at p-8 on the first day of impella support. no mc spike was observed outside of p-level change. no position alarms were observed. the root cause of the hemolysis issue most likely was improper positioning related due to improper technique because the clinical details stated hemolysis improved slightly after position adjustment. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-23625 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the initial manufacturer device report number 1220648-2024-23625 and were added. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
14077 | 1220648-2024-23632 | 20734443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734443&pc=OZD | 11/27/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/27/2021 | 11/20/2024 |  |  | 09/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013243 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Asystole | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. during support the patient became bradycardic then asystolic, and a temporary pacer was placed. it was reported that after 1.33 hours of support the patient expired in the procedural area with no allegations made towards the impella 5.0 as the cause of death. however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14078 | 1220648-2024-23620 | 20734462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734462&pc=OZD | 10/22/2024 | 12/20/2024 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered a complete occlusion of the superficial femoral artery roughly seven days following impella cp explant. vascular consult was scheduled to address the condition. flows were obtained by doppler and the patient was made do-not-resuscitate by their family two days later. intervention was not performed and the patient expired. it is unknown if the complication contributed to the patient's passing. | [5 empty] | B.2 the exact date of death is unknown. the impella device has been discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. b.2 added date of death as it was not reported on the initial manufacturer device report number 1220648-2024-23620. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23620 was submitted in accordance with updated procedures.
14079 | 1220648-2024-23618 | 20734503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734503&pc=OZD | 10/22/2024 | 12/17/2024 |  | 11/20/2024 |  |  | 10/22/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536829 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance through the calcified iliac/aortic stents. the implant was subsequently aborted and an impella 5.5 pump was placed via axillary access for support. there was no patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since patient had calcification.
14080 | 1220648-2024-23623 | 20734507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734507&pc=OZD | 10/24/2024 | 12/16/2024 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555546 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and developed ischemia in the leg post implant same day. following the impella cp implant, the patient had no flow below the insertion site. the patient was noted to have severe peripheral vascular disease in bilateral femoral arteries at the time the condition developed. a distal perfusion catheter was placed to treat the condition. the patient outcome was indicated as unknown; however, support with the implanted pump continued following the intervention. | [5 empty] | The investigation into limb ischemia ¿ reversible has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-23623. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
14081 | 1220648-2024-23624 | 20734510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734510&pc=OZD | 10/24/2024 | 12/06/2024 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515159 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that multiple complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support (mcs) to a patient with cardiomyopathy. following the insertion, it was noted that the impella 5.5 pump was unable to advance through the axillary artery. propofol was then utilized as a lubricant, but the delivery still failed. after multiple unsuccessful attempts, the axillary artery began to bleed, and surgical repair was required. following the intervention, a final attempt was made to deliver the pump but was ultimately unsuccessful. the impella 5.5 pump implant was aborted and an impella cp pump was successfully placed for mcs. the patient was noted to be in stable condition. | [5 empty] | The investigation for the vascular damage and delivery issues has been completed. no product returned. clinical details were sufficient to determine the root cause. the cause of the delivery issue was most likely patient condition related as the 5.5 was unable to be inserted and a cp ultimately was able to be delivered. the cause of the vascular damage peripheral vessel issue was most likely use related due to multiple delivery attempts leading to vessel damage. corrections to the initial mfr report: added d6a/d6b g1-mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist: section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
14082 | 1220648-2024-23626 | 20734522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734522&pc=OZD | 10/24/2024 | 12/01/2025 |  | 11/20/2024 |  |  | 10/24/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and while moving the patient to the bed from the operating room table, the automated impella controller (aic) went black resulting in a pump stop. the pump was restarted, and a pump stop alarm appeared on the aic. the alarms persisted as support continued, and the decision was made to replace the aic. no further alarms were encountered following the replacement. there was no patient harm. | [5 empty] | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: the console was turned on and shut down successfully on 17 oct 2024. since the pump was started, after 3 hours and 48 minutes, though the user performed only the bag change procedure, the cassette was removed. after 20 seconds, the console displayed ¿complete the procedure (user interaction stopped) - complete the purge wizard ¿ alarm,¿ event#420 (imc_logs_ic9740.pdf). most likely, the user had an issue connecting the cassette back into the aic (unreported and unrelated to the reported failure mode). there was no evidence to confirm the reported failure mode. device analysis: the console was tested by initially, inserting the cassette during the case. start was difficult, and the purge motor spindle never rotated. there was no evidence of fluid ingress. after re-mounting the stepper motor, there was no issue in inserting the cassette. most likely, unmounting made the spindle less friction. no issue was found on the console hardware relate to the reported failure mode. it was found that the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). root cause: the reported failure mode was not related to console (b)(4) but was confirmed with the second console (b)(4). the root cause for the pump stop was the impellatronic board malfunction on (b)(4). h3 was the device evaluated by the manufacturer? h6 investigation findings and conclusions d4 model number updated.
14083 | 1220648-2024-23633 | 20734600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734600&pc=OZD | 11/26/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/26/2021 | 11/20/2024 |  |  | 06/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022099330 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, an echocardiogram indicated that the pigtail was in the pap. the surgeon tried to reposition without success as the pigtail was caught on a chord. the surgeon feared the device would slingshot out; therefore, the surgeon decided to leave the device in that position and the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14084 | 1220648-2024-23631 | 20734602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734602&pc=OZD | 10/29/2024 | 12/30/2024 |  | 11/20/2024 |  |  | 10/29/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. while on support, the patient developed hemolysis. the patient was given two units of red blood cells and two units of platelets. the patient was also noted to be in renal failure. the family withdrew care. the patient had multisystem organ failure. the patient expired. | [5 empty] | The investigation of hemolysis and renal failure has been completed. the impella device was not returned for evaluation. data logs were reviewed and long term log showed no suction alarms or positioning alarms noted. real time log at time of p-level change to p-4 showed a few instances of saturated pump flow, motor current, and motor speed with down spikes in placement signal. additionally, pump flows at p-4 are below normal for p-level. clinical details noted that patient was placed on extracorporeal membrane oxygenation soon after impella support was initiated. the pump was in proper position and hemolysis was resolved after blood products were given. additionally, the patient was noted to be in renal and liver failure. ultimately, multisystem organ failure was noted before care was withdrawn. the cause of the hemolysis was most likely patient condition related since giving blood volume resolved the hemolysis. the root cause of the renal failure could not be definitively determined. b.2 revised as required intervention was inadvertently omitted from initial manufacturer device report number 1220648-2024-23631.g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-23631 was submitted in accordance with updated procedures. | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14085 | 1220648-2024-23628 | 20734639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734639&pc=OZD | 10/25/2024 | 12/19/2024 |  | 11/20/2024 |  |  | 10/25/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp insertion the peel away sheath appeared to be cracked at the base of the hemostatic valve. small, but consistent amounts of blood were observed coming out at the mid-line recess where it¿s designed to split. the peel away sheath was removed, reposition sheath advanced and secured. the procedure continued successfully. the patient survived the event. | [5 empty] | The investigation into introducer damage ¿ mechanical has been completed. as per the clinical, bleeding was reported from introducer from creaked hemostatic valve. the cause of the access site bleeding issue was sheath scoreline split, based on given clinical details. introducer damage could not be verified without returned product or provided clinical image. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23628. e4 should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
14086 | 1220648-2024-23630 | 20734676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734676&pc=OZD | 11/28/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/28/2021 | 11/20/2024 |  |  | 04/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022052310 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient (pt) was being placed on impella 5.0 for hemodynamic support. it was reported that the patient had a hematoma at the access site around the sheath. heparin was in the purge for two hours then discontinued and two units of packed red blood cells were administered. it was reported that after explant, the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14087 | 1220648-2024-23634 | 20734678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734678&pc=OZD | 10/27/2024 | 12/23/2024 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553887 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during removal of the impella 5.5, the impella graft was de-clotted prior to removal and embolectomy of axillary artery was performed. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
14088 | 1220648-2024-23635 | 20734743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734743&pc=OZD | 09/27/2024 | 10/28/2025 |  | 11/20/2024 |  |  | 10/29/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148839 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support. the impella 5.5 was returned for continuous placement signal unreliable alarms. during investigation of the impella 5.5, a similar pattern was noted on the automated impella controller (aic). the patient survived the event. the related pump issue was reported in manufacturer device report number 1220648-2024-20804. | [5 empty] | The investigation for the console (aic) loss of opm data has been completed. data logs were reviewed finding there were immediate placement signal and snr errors. the lt logs for this case shows there was persistently low snr with out of bounds oscillating contrast. the aic was returned for evaluation. functional testing found the snr with a known-good optical cavity was 1886 which is below the acceptable threshold. the unreliable placement signal in the field was due to the low snr optical bench. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | H6 updated. b7 other relevant history, including preexisting medical conditions updated.d1 common device name updated.
14089 | 1220648-2024-23638 | 20734917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20734917&pc=OZD | 11/23/2021 | 05/22/2025 |  | 11/20/2024 |  |  | 11/23/2021 | 11/20/2024 |  |  | 09/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138733 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during placement, the right coronary artery began to dissect, and the procedure was stopped. transesophageal echocardiography (tee) was utilized due to the medical staffs concerns of an aortic dissection. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14090 | 1220648-2024-23636 | 20735057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735057&pc=OZD | 11/24/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/24/2021 | 11/20/2024 |  |  | 07/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022041427 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | Hospitalization | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the patient experienced a cerebral infraction and hemorrhage. it was reported that the patient did not recover consciousness due to this issue and eventually expired on support. there is no allegation against the impella, however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14091 | 1220648-2024-23637 | 20735081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735081&pc=OZD | 10/29/2024 | 12/27/2024 |  | 11/20/2024 |  |  | 10/29/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced bleeding from multiple sites during impella cp support. systemic heparin was put on hold in response to the bleed. support with the implanted pump continued following the intervention. it was later documented that the patient's oxygenation levels deteriorated and the family opted not to intubate. the patient was made do-not-resuscitate/do-not-intubate and passed away after care was withdrawn. it is unknown if the bleeding complications contributed to the patient's passing. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient condition because of bleeding from every stick site. g1 reporting contact email revised on manufacturer device report 1220648-2024-23637 in accordance with updated procedures. | B.2 the exact date of death is unknown. the impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14092 | 1220648-2024-23640 | 20735334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735334&pc=OZD | 10/30/2024 | 12/13/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered a watershed stroke during impella cp support. the patient was not moving the right side of their body at the time of discovery. the patient was hemodynamically stable off pressors. the pump was weaned and successfully explanted following the diagnosis. the patient survived the explant. | [5 empty] | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23640. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14093 | 1220648-2024-23639 | 20735471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735471&pc=OZD | 11/23/2021 | 03/07/2025 |  | 11/20/2024 |  |  | 11/23/2021 | 11/20/2024 |  |  | 02/28/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064694 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was reported that the patient was having lots of bleeding and received blood products, heparin was decreased in purge. it was reported that after 223.28 hours of support that care was withdrawn. the patient expired thereafter with no allegations made toward the impella as the cause of death, however there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
14094 | 1220648-2024-23641 | 20735499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735499&pc=OZD | 10/27/2024 | 12/24/2024 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560161 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was measuring deep at the beginning of the case. the 5.5 was pulled back to approximately 3.3. during manipulation the patient went into ventricular fibrillation requiring multiple shocks. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14095 | 1220648-2024-23642 | 20735647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735647&pc=OZD | 10/30/2024 | 12/23/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560885 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support continued to hemolyze. the p-level was decreased and two units of packed red blood cells were given. the patient was upgraded to an impella 5.5. the patient survived the explant. | [5 empty] | The investigation for the hemolysis has been completed. data logs were reviewed finding suction alarms occurred. no motor current (mc) spike observed outside the p-level change. no positioning issue alarms observed. no indication found to confirm the rc for reported issue from the logs. unclear if the repositioning solved the reported issue. the cause of the hemolysis issue cannot be determined since the complaint device not returned for analysis and lack of clinical details. | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14096 | 1220648-2024-23643 | 20735793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735793&pc=OZD | 12/05/2021 | 05/19/2025 |  | 11/20/2024 |  |  | 12/05/2021 | 11/20/2024 |  |  | 04/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075939 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. while on support bleeding was noticed. the patient was taken back to the operating room where the surgeon found a small tear on the medial side of graft. 1 suture was used to control the bleeding. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14097 | 1220648-2024-23646 | 20735917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735917&pc=OZD | 10/30/2024 | 12/24/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025566132 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had sustained ventricular tachycardia. antiarrhythmics were adjusted. transthoracic echocardiogram was obtained and impella 5.5 was repositioned. the physician stated that impella position is optimized at 4 centimeters with the patient¿s anatomy 5 centimeters caused the patient to go into ventricular tachycardia requiring cardioversion. additionally, the patient is heparin-induced thrombocytopenia (hit) positive. systemic heparin was removed and the patient was on direct thrombin inhibitor therapy. it was further reported that the family decided to withdrew care and the patient expired. | [5 empty] | The investigation thrombocytopenia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23646 in accordance with updated procedures. | B.2 the exact date of death is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14098 | 1220648-2024-23644 | 20735966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20735966&pc=OZD | 10/27/2024 | 12/16/2024 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487362 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed for care escalation. on day 11 of support the team observed left ventricle clot. the thrombus was treated/intervened with the addition of heparin. the pump remains on for support and its noted the patient remains stable. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14099 | 1220648-2024-23647 | 20736085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736085&pc=OZD | 06/24/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 06/24/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024085 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the systemic heparin was held due to bleeding at the site. the patient remained on support and the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14100 | 1220648-2024-23645 | 20736156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736156&pc=OZD | 06/05/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 06/05/2021 | 11/20/2024 |  |  | 06/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1494055 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient was returned to the operating room due to significant bleeding from the chest tubes. the surgeon evaluated and resolved any issues he found and requested to remove heparin from purge until the next morning to see if that would help slow bleeding issues in the chest. heparin was added into purge the next day but then removed due to bleeding concerns and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14101 | 1220648-2024-23648 | 20736181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736181&pc=OZD | 02/08/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 02/08/2022 | 11/20/2024 |  |  | 02/28/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1467868 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a hematoma and was sent for explant of the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14102 | 1220648-2024-23652 | 20736269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736269&pc=OZD | 10/27/2024 | 12/27/2024 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547659 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support under transesophageal echocardiogram, noted a large mobile thrombus in the left ventricle. it appeared to be attached to the inlet cage with pedunculated thrombus attached and floating towards the aortic valve. a decision was made to use therapeutic anticoagulation. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the thrombus has been completed. the impella was returned for evaluation and there were no abnormalities found with the pump. clinical details were not sufficient to determine the root cause. therefore, the root cause of the thrombus could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14103 | 1220648-2024-23649 | 20736408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736408&pc=OZD | 11/01/2024 | 12/26/2024 |  | 11/20/2024 |  |  | 11/01/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553862 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 began alarming for suction. metrics were consistent with clot entrainment. as there was decreased flow and then flow saturation. the impella 5.5 was explanted. imaging showed a clot at the tip of the impella 5.5. the patient survived the explant and continued on extracorporeal membrane oxygenation support. | [5 empty] | The investigation for the thrombus and saturated flow has been completed. the pump was not returned for investigation. data logs show a motor current spike followed by saturated pump flow at the end of the case run. no trend was seen on purge pressure. as per the given clinical information, the doctor noticed suction alarm and a trend of clot interaction. additionally, the imaging showed a clot on the pump. the pump was removed. the cause of the saturated pump flow issue was most likely biomaterial, based on given clinical and data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14104 | 1220648-2024-23654 | 20736489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736489&pc=OZD | 07/30/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 07/30/2021 | 11/20/2024 |  |  | 12/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022034454 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was oozing from the impella surgical site, and heparin was withheld from the purge to manage the complication. a hematoma was noted at the access site and right arm of the patient. a surgical repair was required, and four units of packed red blood cells were transfused. also, in the operating room while performing a surgical site washout, a stitch was identified on the graft. the anastomosis was found to be leaking and was repaired. it was recommended that the impella pump speed be decreased to reduce hemolysis. the patient remained on support and weaning was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14105 | 1220648-2024-23651 | 20736505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736505&pc=OZD | 06/29/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 06/29/2021 | 11/20/2024 |  |  | 04/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075864 |  | 0085062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient is to rest on the impella for a few days and systemic heparin was placed on hold until drainage from chest slowed, per the physician¿s request. the patient remained on support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14106 | 1220648-2024-23650 | 20736613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736613&pc=OZD | 10/27/2024 | 09/09/2025 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552952 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications occurred during impella cp implant for mechanical circulatory support. during delivery, vascular damage occurred at the impella insertion site resulting in bleeding and formation of a large hematoma. perclose sutures were applied to the site and an unspecified quantity of blood products were transfused to treat the condition. it was additionally documented that the patient suffered cardiac arrest during impella placement. compressions were performed to stabilize the patient. support with the implanted pump proceeded following the interventions until successful explant and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: product was not returned and logs are not applicable. it was noted that the vascular damage occurred during placement and active compressions while the patient was in cardiac arrest. the root cause of the vascular damage was most likely use-related since the bleeding resolved after adding sutures. vf/vt/af/at: the root cause of the vt/vf was unable to be determined due to insufficient clinical details. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23650.
14107 | 1220648-2024-23655 | 20736773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736773&pc=OZD | 08/18/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 08/18/2021 | 11/20/2024 |  |  | 04/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022073442 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient was observed to have heparin induced thrombocytopenia (hit). the impella purge solution was changed to argatroban to manage the complication. the patient remained on support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14108 | 1220648-2024-23656 | 20736775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20736775&pc=OZD | 10/27/2024 | 12/18/2024 |  | 11/20/2024 |  |  | 10/27/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541886 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during the implant there was a lot of ectopy, the pump appeared deep so the pump was pulled back slightly and it came across the valve. the impella was removed from the patient and it was inspected. the pump was implant again and it appeared deep. the physician repositioned the pump and it came back into the aorta again. the prior steps were repeated and the pump was placed into the left ventricle. the impella waveforms improved and flows increased to 2.8-2.9 lpm on p-7. support continued. shortly after, the pump appeared deep and the doctor tried to move the pump back. after multiple times attempting to pull back, waveforms were slightly improved so the doctor opted to leave the impella slightly deep due to concern of it pulling back across the valve again. the patient expired after 9.45 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23656. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14109 | 1220648-2024-23658 | 20737010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737010&pc=OZD | 11/30/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/30/2021 | 11/20/2024 |  |  | 09/30/2023 |  | 11/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022139328 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support the patient had access site bleeding around the sheath. the patient was sent to the operating room to resolve the tissue bleeding. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14110 | 1220648-2024-23661 | 20737025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737025&pc=OZD | 12/23/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 12/23/2021 | 11/20/2024 |  |  | 06/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022109774 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. it was reported that the patient was brought to cath lab for left atrial extracorporeal membrane oxygenation (ecmo) and removal of the impella due to a thrombus seen on an echocardiogram. the impella was removed and the patient continued ecmo support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14111 | 1220648-2024-23662 | 20737206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737206&pc=OZD | 01/28/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 01/28/2022 | 11/20/2024 |  |  | 11/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022148708 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. it was noted that the patient¿s left axillary was completely occluded, and right axillary was also calcified. the pump insertion was attempted, but it did not pass after the anastomosis. noted repositioning of the artificial blood vessel in femoral artery, but it still did not pass, so staff reported a change to impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14112 | 1220648-2024-23664 | 20737303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737303&pc=OZD | 02/11/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 02/11/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136244 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the pump would not pass stenosis at the innominate artery, turndown not visualized on initial angiogram; procedure aborted, and the patient will be monitored in the unit for possible impella cp support if needed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14113 | 1220648-2024-23668 | 20737311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737311&pc=OZD | 03/17/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 03/17/2022 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021008530 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was a "high purge pressure" alarm. the physician tried to solve the problem on his own by changing the ps, but to no avail. it was noted throughout the case that purge pressure was high, no further information provided. | [5 empty] | D6b explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14114 | 1220648-2024-23653 | 20737322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737322&pc=OZD | 07/15/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 07/15/2021 | 11/20/2024 |  |  | 12/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022034752 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was active internal bleeding over the coronary artery bypass graft operative site. blood products were transfused to top up hemoglobin (hb) level. suction alarm due to hypovolemia. latest hb level was 4.7. re-operation afterwards. the next day exploratory sternotomy was done and clots were found in mediastinum. hemostasis achieved. intra-op blood products were further given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14115 | 1220648-2024-23669 | 20737358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737358&pc=OZD | 09/11/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532387 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock attempted to be implanted with an impella cp device for mechanical circulatory support. during implant procedure, pump encountered delivery issues. the implant procedure was aborted following groin angiogram. per physician, vessel size was very small and calcified. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because the patient vessel size was small and calcified preventing the pump to insert. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14116 | 1220648-2024-23663 | 20737395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737395&pc=OZD | 02/09/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 02/09/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131608 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the impella 5.0 was loaded and inserted through the sheath and into the graft, where it was unable to make the turn into axillary artery and the implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14117 | 1220648-2024-23671 | 20737541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737541&pc=OZD | 10/28/2024 | 08/22/2025 |  | 11/20/2024 |  |  | 10/28/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1079313 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has system self-check failed error. there was no reported patient involvement. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
14118 | 1220648-2024-23665 | 20737604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737604&pc=OZD | 03/02/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 03/02/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022146983 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the impella was noted to be too deep and touching the mitral valve. the physician did not want to reposition at the time, would revisit the following day. the physician also suspected that the impella was causing hemolysis, but thought that it may be caused by a clot near the inlet and not placement. there were also some concerns with bleeding due to hemoglobin being low. the patient received a unit of blood overnight. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14119 | 1220648-2024-23666 | 20737631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737631&pc=OZD | 10/28/2024 | 12/20/2024 |  | 11/20/2024 |  |  | 10/28/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539552 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a surgeon was unable to manipulate the positioning of an impella 5.5 pump following a scheduled coronary artery bypass grafting (cabg) procedure. it was documented that despite pressing the catheter lock button, the surgeon was unable to shallow the catheter. after multiple unsuccessful attempts, the surgeon cut the heavy ties attached to the blue suture hub and graft which allowed for manipulation of the catheter. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email
14120 | 1220648-2024-23676 | 20737662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737662&pc=OZD | 09/09/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/09/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481779 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. implant procedure was unsuccessful due to difficult patient's anatomy. staff noted inability to pass through patient's vasculature at the axillary artery. surgeon reported catheter lost its ¿push-ability¿ to advanced towards aortic valve (av). pump was replaced with a new impella 5.5. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, implant unsuccessful due to difficult anatomy as per the clinical description provided. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
14121 | 1220648-2024-23677 | 20737715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737715&pc=OZD | 09/08/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/08/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500908 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after patient was transferred to intensive critical unit (icu), left groin impella site was noted to have developed a large hematoma and oozing was observed. physician elected to place femostop. vascular surgery was consulted to look at patient's groin and decided on next steps. staff noted that patient was volume resuscitated and received two units of packed red blood cells [prbc¿s]. patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding - major has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
14122 | 1220648-2024-23679 | 20737742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737742&pc=OZD | 10/24/2024 | 06/26/2025 |  | 11/20/2024 |  |  | 09/02/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502390 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the experienced an embolization of a branch of the left iliac artery, causing a retroperitoneal bleed. heparin was removed from the purge. in addition, high purge pressure was reported. tissue plasma activator was administered, and the patient remained on impella support . | [5 empty] | The investigation of vascular damage - peripheral vessel, thrombus, and high purge pressure has been completed. insufficient product was returned for analysis. the pump returned was severed off at the catheter leaving the cannula and part of the catheter separate from the rest of the catheter and patient cable. high purge pressure: the severed pump was inspected and no biomaterial was observed in the purge gap or on the impeller. clinical mentions high purge pressure event occurred after which tissue plasminogen activator (tpa) was added. clinical mentions issue resolved after hours of administration. the logs show a gradual rise in purge pressure starting at 20:54 on (b)(6). the purge pressure rises for about 90 minutes before remaining elevated in the 1000s. the elevated purge pressure remains for about 2 hours before gradually decreasing to normal range. the root cause of the high purge pressure was most likely biomaterial as evidenced by the gradual rise in purge pressure and clinical mention of tpa resolution. vascular damage - peripheral vessel: clinical reported during support, the patient had an embolization of a branch off the left iliac artery. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23679. h3 was erroneously selected on the initial manufacturer device report number 1220648-2024-23679. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2024-23679. h10 narrative erroneously stated the device was not received for investigation on the initial manufacturer device report number 1220648-2024-23679. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: high pressure purge: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14123 | 1220648-2024-23667 | 20737753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737753&pc=OZD | 10/30/2024 | 12/24/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562900 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and there was low pump flows. the impella was placed for life saving measures. the staff did not confirm prior to the impella placement that there was no thrombus in the left ventricle. after the placement, low flows were observed and blood products were given. the pump was also repositioned without resolution. a thrombus was never confirmed just suspected due to low pump flows. the patient expired after 117.45 hours of impella support. | [5 empty] | The investigation for the low pump flow has been completed. the impella device nor the data logs were returned for evaluation. therefore, without sufficient clinical details, the root cause of the low pump flow could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: b2 added date of death g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14124 | 1220648-2024-23673 | 20737794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737794&pc=OZD | 09/10/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/10/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534252 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. during implant procedure, the implant was aborted due to tortuosity and inability to advance this impella in left ventricle (lv). a new impella 5.5 pump was implanted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, impella was aborted due to tortuosity and inability to advance this impella in lv as per the clinical description provided.
14125 | 1220648-2024-23687 | 20737796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737796&pc=OZD | 10/27/2024 | 07/10/2025 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547041 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. during intensive care unit management, a fluid leak was confirmed near the pump reservoir. as there was no significant change in the purging fluid, the situation was observed. two days after insertion, the pump was removed without any problems. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". ¿clean the connector cable with 70% isopropyl alcohol.¿. | The investigation of purge leak: pump has been completed. the pump was returned for investigation. the pump was purged for over 30 minutes and there was a minor leak observed on the reservoir. no visual damage observed. the reservoir was put under keyence ta crack was seen. the root cause of the purge leak was a damaged sidearm reservoir as evidenced by returned product inspection and testing. d9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report 1220648-2024-23687 in accordance with updated procedures.
14126 | 1220648-2024-23675 | 20737800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737800&pc=OZD | 10/30/2024 | 12/23/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545517 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and there was a leak at the red hub. the fluid was pink tinged. the staff was aware the pump may stop. support was continued and the patient was successfully weaned. the patient survived the explant. | [5 empty] | The investigation product damage (hole in catheter) has been completed. the product was returned and evaluated. the red plug was opened and inspected. there was dried blood found inside the red plug. a puncture hole was found in the catheter between the 75 and 80 cm mark. the root cause of the hole in catheter was use-related. d9 date device returned to manufacturer added. g1 reporting contact email revised on manufacturer device report 1220648-2024-23675 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14127 | 1220648-2024-23680 | 20737804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737804&pc=OZD | 08/24/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 08/24/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021019960 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. it was reported that the patient was observed to be positive for heparin induced thrombocytopenia after the patient was taken off bipella. an external ventricular assist device with oxygenator was implanted. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14128 | 1220648-2024-23670 | 20737814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737814&pc=OZD | 09/11/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during implant procedure, pump encountered delivery issues due to patient's anatomy. implant procedure was ultimately aborted without any harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because the insertion was unsuccessful due to patient anatomy.
14129 | 1220648-2024-23681 | 20737818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737818&pc=OZD | 08/29/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 08/29/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025471268 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient's distal pulses on both feet were faint, but the right dorsalis pedis was not dopplerable. the physician then chose to wean and explant the impella, and the patient survived with no other reported harm. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia could not be determined. as noted in the impella ifu: limb ischemia: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14130 | 1220648-2024-23684 | 20737835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737835&pc=OZD | 10/30/2024 | 12/26/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531939 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella 5.5 implant. however, it was unsuccessful. the impella 5.5 got through the graft anastomosis. the pump would not advance due to tortuosity. the implant was aborted and an impella cp was inserted. the patient¿s outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23684 in accordance with updated procedures.
14131 | 1220648-2024-23682 | 20737970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737970&pc=OZD | 01/31/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 01/31/2022 | 11/20/2024 |  |  | 11/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023160028 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited outflow blocked alarms even though it was noted that the impella was well positioned in the left ventricle. it was reported that the patient was stable and survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14132 | 1220648-2024-23657 | 20737976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737976&pc=OZD | 08/17/2021 | 05/21/2025 |  | 11/20/2024 |  |  | 08/17/2021 | 11/20/2024 |  |  | 11/30/2021 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1448549 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient developed sepsis from an unknown origin. the medical staff changed the patient¿s medication to manage the issue. the patient remained on impella support and was successfully weaning. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14133 | 1220648-2024-23692 | 20737996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20737996&pc=PYX | 10/31/2024 | 12/30/2024 |  | 11/20/2024 |  |  | 10/31/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529486 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella rp flex. the pump was unable to pass tricuspid valve status post tricuspid band repair. the patient¿s outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the pump was unable to pass the tricuspid valve status post tricuspid band repair. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23692 was submitted in accoreance with updated procedures.
14134 | 1220648-2024-23689 | 20738048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738048&pc=OZD | 03/08/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 03/08/2022 | 11/20/2024 |  |  | 05/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1476801 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that complications were encountered during attempted delivery of an impella 5.0 pump for mechanical circulatory support. due to resistance met during the implant attempt, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14135 | 1220648-2024-23688 | 20738135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738135&pc=OZD | 02/22/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 02/22/2022 | 11/20/2024 |  |  | 07/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121630 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported no heparin in purge or systemically due to the patient bleeding previously. it was also noted that a limited echo was done due to a thrombus in left ventricle; transesophageal echocardiography was discussed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14136 | 1220648-2024-23659 | 20738148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738148&pc=OZD | 10/28/2024 | 12/01/2025 |  | 11/20/2024 |  |  | 10/28/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1853697 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had an automated impella controller (aic) at distribution center and being shipped when the aic was assessed and found to have the inability to shut down properly. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Investigation completed. data analysis: the console software (v11.0.1j) was reinstalled in an attempt to fix the pmd hardware revision. after the software reinstallation, the pmd hardware revision was as intended: 'j 257 v85_9 c.' the software reinstallation process removes previous logs, so no logs were available to confirm the reported failure mode. the ¿selector knob¿ issue is not related to the pmd version issue. most likely, the selector knob malfunction prevented console shutdown. device analysis: this console was tested in fs and was unable to reproduce the selector knob malfunction. the console was able to run the pump at each p-level from p1 to p9. no issue was found with the selector knob or rotatory knob¿s select function or rotate function. upon pressing the selector knob 50 times, measured 5v between tp100.0 and tp101.0 and heard the beep sound. root cause: the root cause for the inability to shut down the console was not determined due to the inability to reproduce. likely, the selector knob/rotary encoder malfunction prevented console shutdown. h6 investigation type, findings and conclusions g1 reporting contact email. | A2 age should have been left blank on the initial manufacturer device report number 1220648-2024-23659 as no patient was involved. d2 revised common device name as it was reported incorrectly on initial manufacturer device report number 1220648-2024-23659. e1 revised name prefix/title as it was reported incorrectly on initial manufacturer device report number 1220648-2024-23659. g1 revised mdr reporting contact email as it was reported incorrectly on initial manufacturer device report number 1220648-2024-23659. g2 foreign selected was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-23659. h6 codes 4114 and 3221 were reported incorrectly on initial manufacturer device report number 1220648-2024-23659.
14137 | 1220648-2024-23672 | 20738154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738154&pc=OZD | 09/11/2024 | 07/03/2025 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520017 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, cp had positioning issues due to being pulled out across the aortic valve (av) and unable to readvance. another impella cp pump was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the positioning issue was use related since the pump pulled out during repositioning.
14138 | 1220648-2024-23691 | 20738283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738283&pc=OZD | 09/29/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/29/2022 | 11/20/2024 |  |  | 05/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023215554 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. during implantation the impella, bleeding occurred at the anastomosis area which was managed with fibrin, felt, and (b)(4) unit of red blood cells. according to the physician the patient went for a ct scan which revealed the patient's vascular status was not possible for the implantation of the impella. the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14139 | 1220648-2024-23696 | 20738361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738361&pc=OZD | 07/15/2022 | 09/16/2025 |  | 11/20/2024 |  |  | 07/15/2022 | 11/20/2024 |  |  | 08/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1497958 |  | 005060 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician confirmed the patient had heparin-induced thrombocytopenia (hit) and was placed on argatroban in the purge. later, during ongoing 5.0 support, there was a purge-sidearm that needed a bypass done due to leaking. afterwards the pump remained on for support. the team was getting additional information from med affairs for sodium bicarb in the purge. the patient survived and was successfully weaned. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for purge leak- pump has been completed. the product was returned and a filter leak was confirmed at one of the venting holes. the root cause of the purge leak was determined to be a broken purge filter most likely due to ipa interaction.
14140 | 1220648-2024-23698 | 20738372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738372&pc=OZD | 10/30/2024 | 12/16/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548131 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient¿s right lower extremity was noted to be cold and mottled with minimal toe movement. a vascular surgeon was consulted. rewarmed and sensation improved as did motor function. the surgeon noted issue is likely due to low-flow state. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14141 | 1220648-2024-23660 | 20738386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738386&pc=OZD | 10/28/2024 | 01/02/2025 |  | 11/20/2024 |  |  | 10/28/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535449 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 88-year-old male patient on impella cp support developed hematoma on the impella access site. surgical intervention was planned. it was further reported that the patient expired on impella support. no other information is available. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23660 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14142 | 1220648-2024-23703 | 20738407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738407&pc=OZD | 08/22/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 08/22/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022047653 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support it was noted that heparin was turned off due to a nosebleed. continued dialysis was also noted. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14143 | 1220648-2024-23706 | 20738416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738416&pc=OZD | 10/30/2024 | 12/26/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562064 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a biventricular failure patient admitted in cardiogenic shock. the pump was placed but the team had to hold heparin for the bleeding at the access site. the patient additionally did receive blood product in form of platelets. it was also reported that the patient lost pulses in right leg, impella access side. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14144 | 1220648-2024-23705 | 20738542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738542&pc=OZD | 07/29/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 07/29/2022 | 11/20/2024 |  |  | 09/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504585 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support there was a drop in platelets that led to concern for heparin-induced thrombocytopenia. heparin drip was switched to argatroban with sodium bicarbonate in the purge. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14145 | 1220648-2024-23699 | 20738597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738597&pc=OZD | 11/29/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 11/29/2022 | 11/20/2024 |  |  | 08/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495914 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported that the patient was bleeding. blood products were required and heparin decreased. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14146 | 1220648-2024-23708 | 20738646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738646&pc=OZD | 11/01/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 11/01/2022 | 11/20/2024 |  |  | 11/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023207709 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia and was shocked several times to regain vitals. the patient also went into atrial fibrillation and the impella pump speed level was increased to help compensate for loss of cardiac output. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14147 | 1220648-2024-23720 | 20738717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738717&pc=PYX | 10/26/2024 | 12/12/2024 |  | 11/20/2024 |  |  | 10/26/2024 | 11/20/2024 |  |  | 10/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023279903 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiogenic shock from an acute myocardial infarction was implanted with an impella rp device for mechanical circulatory support and experienced bleeding at the access site. an additional suture was applied with forward tension on the catheter which stopped the bleeding. the patient was administered two units of packed red blood cells and noted to be in stable condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by adding additional forward suture. corrections to the initial mfr report: g1 mdr reporting contact email. g1 mfr site name and address updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
14148 | 1220648-2024-23711 | 20738797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738797&pc=OZD | 11/11/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 11/11/2022 | 11/20/2024 |  |  | 03/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022084638 |  | 005064 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during insertion in the right axillary artery, the physician was unable to advance the impella into the vessel. several attempts were made to advance the impella 5.0 through graft. the patient was stable, and the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14149 | 1220648-2024-23716 | 20738808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738808&pc=OZD | 08/25/2021 | 05/22/2025 |  | 11/20/2024 |  |  | 08/25/2021 | 11/20/2024 |  |  | 01/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022034806 |  | 004680-AU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. patient had 1 episode of ventricular tachycardia (vt). defibrillation was performed and arrhythmia resolved. pump re-positioned with echo. subsequently, the patient experienced ventricular tachycardia storm. ecmo was inserted and patient was intubated. vt runs continued. ecmo flows were increased and noradrenaline infusion discontinued. on day 14 post-implant/day 5 ecpella, the patient experienced hematoma evacuation to right arm/chest from ecmo insertion site. existing left chest drain noted with significant bleeding. blood products were transfused to patient. ptt increased to over 100 and systemic heparin infusion was titrated down accordingly. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14150 | 1220648-2024-23690 | 20738849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738849&pc=OZD | 10/30/2024 | 02/11/2025 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181672 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when plugging the impella catheter into the automated impella controller (aic), an error message appeared. the message stated that the optical sensor was not detected. staff attempted to trouble shoot by looking at the optical sensor and unplugging and plugging the pump back in to no avail. it was reported there was also a controller error (yellow alarm). the aic was replaced successfully and there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of optical signal issue and automated impella controller (aic) - controller error (yellow alarm) has been completed. according to the data and device analysis the root cause for the aic issues was a defective optical bench. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
14151 | 1220648-2024-23685 | 20738920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738920&pc=OZD | 09/27/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/27/2022 | 11/20/2024 |  |  | 05/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023226759 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.0 and extracorporeal membrane oxygenation (ecmo) support for mechanical circulatory support due to an unspecific retro sternal bleeding and tamponade. the purge fluid was changed to sodium bicarbonate. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14152 | 1220648-2024-23707 | 20738973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20738973&pc=OZD | 10/29/2024 | 12/18/2024 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562343 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump to support a patient admitted in for aortic valve replacement. the pump was placed but there were signs of hemolysis observed on day two of the support. the pump was repositioned but urine remained red. there was adjustments to the p level and the patient remained on support an additional three days. weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23707. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14153 | 1220648-2024-23709 | 20739021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739021&pc=OZD | 07/09/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 07/09/2021 | 11/20/2024 |  |  | 08/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1498424 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported that the patient was bleeding. it was noted that there was no heparin in purge until bleeding was controlled. staff later noted that the site looked good without hematoma/bleeding. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14154 | 1220648-2024-23712 | 20739095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739095&pc=OZD | 04/27/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 04/27/2021 | 11/20/2024 |  |  | 10/31/2021 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1455361 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the us complainant reported that, during patient support, there was no heparin due to bleeding concerns. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14155 | 1220648-2024-23713 | 20739134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739134&pc=OZD | 12/02/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 12/02/2022 | 11/20/2024 |  |  | 03/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023209492 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support there was high purge pressure alarms. no further information was provided. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14156 | 1220648-2024-23678 | 20739164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739164&pc=OZD | 10/30/2024 | 06/27/2025 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518511 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was inserted and placement signal low alarm occurred. the impella 5.5 was removed and exchanged. the patient¿s outcome is unknown. | [5 empty] | The investigation of optical signal issue has been completed. the returned pump was inspected and there were no physical abnormalities. the pump passed the light continuity test and all 5s parameters were within normal range. log review showed the following optical trends: signal not reliable (snr) decrease to 0, light decrease, lamp increase to 100, contrast decrease, and gain increase. upon initial insertion, the parameters were in spec, but about 50 minutes later, snr dropped with the other parameter changes. real time (rt) log data was unavailable at the time of the snr drop. this drop was accompanied by a placement signal unreliable alarm and occurred while advancing the pump over the wire. the replacement pump worked without complication. the root cause of the optical signal issue was not determined as insufficient data logs were returned for analysis - rt logs at time of issue occurrence was not returned, there were no abnormalities on returned product, and insufficient clinical details related to failure were provided. d9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report 1220648-2024-23678 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14157 | 1220648-2024-23674 | 20739166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739166&pc=OZD | 09/11/2024 | 09/17/2025 |  | 11/20/2024 |  |  | 10/29/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486513 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. it was reported that the sterile sleeve broke while attached to the patient. the sterile sleeve disconnected at the tuohy lock. the patient remained stable and continued on support therefore the explant date and patient outcome are unknown. | It was further reported that the patient was transferred to another facility and pump in aorta alarms occurred. echo done and impella was in the correct position. aortic pressures not correlated with the arterial line. aortic sensor disappeared. placement signal issue. |  |  |  |  | The investigation into the impella cp has been completed. no product was returned. per the clinical investigation, the root cause of the sterile sleeve damage was not determined since product was not returned. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. the data log shows that the placement signal monitoring was disabled after several impella position in aorta alarms were reported. the pump flow gradually became saturated before the aorta alarms were triggered, a mc spike occurred, and then flow remained saturated until the end of the case run, with less pulsatile motor current compared to the start of the case run. no abnormalities were reported in the optical signal parameters, or purge pressure. according to the provided clinical details, a hematoma was reported at the right radial access site only. additionally, the doctor noted placement signal discrepancy between the aic and the arterial line, accompanied by a pump position in aorta alarm, while confirming a good pump position. the cause of the access site bleeding issue was most likely patient condition related. the relationship between the radial access site bleed and the impella was unrelated, as the pump was inserted through the right femoral artery. fm placement signal issue was changed to fm optical signal issue since the impella cp has an optical sensor. the cause of the optical signal issue and the saturated pump flow was most likely biomaterial ingestion since pump flow trended up at the time of the position in aorta alarms when position was confirmed to be correct, followed by saturated flows. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. the cause of the sterile sleeve damage was not determined since product was not returned.
14158 | 1220648-2024-23697 | 20739190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739190&pc=OZD | 09/24/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/24/2022 | 11/20/2024 |  |  | 02/29/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023177085 |  | 004680-AU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support elevated partial thromboplastin time, and ventricular tachycardia (vt) episodes was noted. the patient continued on support until being bridged to a left ventricular assist device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14159 | 1220648-2024-23695 | 20739244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739244&pc=OZD | 10/31/2024 | 12/23/2024 |  | 11/20/2024 |  |  | 10/31/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530834 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed hemolysis. no intervention at this time. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23695 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14160 | 1220648-2024-23717 | 20739256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739256&pc=PYX | 10/26/2024 | 12/17/2024 |  | 11/20/2024 |  |  | 10/26/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025560984 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic |  | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp flows decreased to 2.5 l/m and pulmonary artery was not correlating. the patient had no path for escalation. the physician opted to remove the rp and placed a new one in an effort for right ventricle recovery. impella explanted and ultimately replaced with a new rp flex. patient expired within 3 hours in the intensive care unit. | [5 empty] | The investigation into low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-23717. e4: should have been left blank. g1: reporting contact email. g1: reporting contact fax number missing. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
14161 | 1220648-2024-23686 | 20739582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739582&pc=OZD | 10/30/2024 | 01/28/2025 |  | 11/20/2024 |  |  | 10/30/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513035 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional details state that high purge pressure and low purge flow were reported by the physician. the lysis protocol was initiated. the problem was resolved within hours after starting the tissue plasminogen activator (tpa) protocol. | The device exhibited low purge flow and high purge pressure. the resolution for this issue is unknown. | The complainant had a impella 5.5 pump placed for a transfer patient. a hematoma was reported and treated by two units of blood and coagulation of patient was reduced. the patient remained stable and the procedural outcome was the patient survived surgery. |  |  |  | The investigation for the access site bleed and the high purge pressure has been completed. the pump and data log were not returned for investigation. according to the clinical details, the patient had poor coagulation due to hypocoagulopathy caused by liver disease and excessive heparin at the time of the access site bleeding event. the issue was resolved quickly. additionally, a few days later, high purge pressure was reported and resolved after administering tpa. the cause of the access site bleeding issue was high patient act values after given too much heparin, based on give clinical details. the cause of the high purge pressure issue was biomaterial, based on provided clinical details. corrections to the initial mfr report: b1-selected product problem. b5-added 2nd failure mode. g1 mdr reporting contact email. h6-impact code 4644 added. h6-med dev prob code 1491 added. | The additional information is reflected in b2, b5, g3, g6, h1, h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14162 | 1220648-2024-23694 | 20739614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739614&pc=OZD | 05/10/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 05/10/2022 | 11/20/2024 |  |  | 06/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1486754 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. it was reported that the patient experienced 1500 milliliters of blood loss. the patient was administered volume and red blood cells. additionally, thrombus was identified and a thrombectomy performed. it was reported that the patient had normal distal pulses, and the site was free of bleeding. the patient survived normal device explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14163 | 1220648-2024-23683 | 20739632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739632&pc=PYX | 08/29/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 08/29/2024 | 11/20/2024 |  |  | 11/30/2025 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024425611 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella rp device for mechanical circulatory support. the team spent 90 minutes trying to place the device, but the proper position could not be accessed. the implant was aborted, and the procedure was continued with a different device. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because the rp flex was unable to pass through the vasculature and the patient was back on 5.5 support. as noted in the impella ifu: delivery issues: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
14164 | 1220648-2024-23722 | 20739661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739661&pc=OZD | 10/22/2024 | 06/26/2025 |  | 11/20/2024 |  |  | 10/31/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563691 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Disability | Required Intervention | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry upon a new adverse event of worsening acute kidney injury superimposed on stage 3b chronic kidney disease with relationship listed as possibly related to the pump. previously it was reported that it was an unknown (undetermined) relationship. | A report from the abiomed¿s long-term outcome and quality indicator (loqi) impella registry noted a patient with heart failure with reduced ejection fraction was implanted with an impella cp device for mechanical circulatory support. while on the impella cp support, the patient experienced acute kidney injury superimposed on stage 3b chronic kidney disease with an "unknown (undetermined)" relationship to the impella cp device. creatinine was 2.6 on day of implant and increased to 4 four days later. the patient was coughing up blood and had some blood in urine which resolved. continuous renal replacement therapy was started and continued until the patient's demise approximately four days later. the patient was on support for a total of eight days. of note, the patient was not a candidate orthotopic heart transplant given relevant history. however, the patient was possibly a candidate for left ventricular assist device, but such support was refused by the patient. |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry upon a new adverse event of worsening acute kidney injury superimposed on stage 3b chronic kidney disease with relationship listed as possibly related to the pump. previously it was reported that it was an unknown (undetermined) relationship. | The investigation for the renal failure and vascular damage has been completed. no product was returned for analysis. the data logs do not show any motor current spikes or trending placement signal. clinical details do not reveal positioning issues. the root cause of the renal failure was not determined as insufficient information related to failure was provided. the root cause of the vascular damage - peripheral vessel was not determined as insufficient information related to failure was provided and no product was returned corrections to the initial mfr report: d4 updated udi number | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
14165 | 1220648-2024-23700 | 20739663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739663&pc=OZD | 07/14/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 07/14/2022 | 11/20/2024 |  |  | 11/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023204419 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 5.0 device for mechanical circulatory support. later, an activated coagulation time of 360 seconds and purge without heparin were reported, along with anticoagulation issues and slight bleeding. a transfusion of ek/tk and human prothrombin complex (ppsb) were given. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14166 | 1220648-2024-23719 | 20739772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739772&pc=OZD | 12/01/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 12/01/2022 | 11/20/2024 |  |  | 03/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023209497 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Malposition of Device | [5 empty] | Renal Failure | Thrombosis/Thrombus | Unspecified Heart Problem | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 5.0 device for mechanical circulatory support. according to the transesophageal echocardiography [tee], the impella appears to be close to the posterior mitral leaflet [pml]. there was a papillary muscle rupture found to have resulted in a massive myocardial infarction [mi]. the surgeon decided to remove the impella due to the poor position and massive mi. in addition, a left ventricle [lv] thrombus was found and the patient was noted to be in critical condition with renal failure. the impella was explanted and the patient continued on extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14167 | 1220648-2024-23702 | 20739787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739787&pc=OZD | 08/27/2021 | 05/28/2025 |  | 11/20/2024 |  |  | 08/27/2021 | 11/20/2024 |  |  | 06/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1487001 |  | 005082 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support it was noted that purge flow was down, and urinary output was pink/red. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14168 | 1220648-2024-23721 | 20739818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739818&pc=OZD | 11/29/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 11/29/2022 | 11/20/2024 |  |  | 11/30/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161306 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient had heparin induced thrombocytopenia. the purge was changed to sodium bicarbonate and bivalirudin, and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14169 | 1220648-2024-23715 | 20739827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739827&pc=OZD | 10/26/2024 | 12/17/2024 |  | 11/20/2024 |  |  | 10/26/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531281 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following implant, an impella cp pump stopped working unexpectedly. it was documented that after a minute of support, the motor current spiked and the pump stopped. the pump was restarted multiple times, but stopped again each time. the pump was subsequently replaced as a result of the noted failure. there was no patient harm. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. the returned product was inspected and biomaterial was found by the impeller in the inlet cage. the pump stop did not reproduce in the lab and the biomaterial was pushed out of the cannula when the pump was run in the lab. the data logs confirm impella stopped motor current high alarms and show a motor current spike before pump stop and show failed restart attempts. the root cause of the pump stop was determined to be biomaterial.
14170 | 1220648-2024-23710 | 20739843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20739843&pc=OZD | 07/12/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 07/12/2022 | 11/20/2024 |  |  | 07/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1494066 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced tachycardia. the patient was in and out of atrial flutter. actions taken are unknown. the patient continued on support until the device was successfully weaned. the patient outcome was reported to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14171 | 1220648-2024-23693 | 20740022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20740022&pc=OZD | 10/01/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 10/01/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136286 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, it was noted that the partial thromboplastin time (ptt) was elevated and that a gastrointestinal bleed was observed. heparin drip was titrated down per protocols for elevated ptt. the patient was extubated and central venous pressure (cvp) decreased to 10 following extubating. the patient was noted to be improving. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14172 | 1220648-2024-23718 | 20740214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20740214&pc=OZD | 10/26/2024 | 12/20/2024 |  | 11/20/2024 |  |  | 10/26/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524059 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a patient experienced multiple complications during impella cp support. after one day of support, the patient's urine became pink tinged, and their hemoglobin levels had dropped to 8.6. the pump was noted to be too deep at the time of the developing hemolysis and the device was pulled back into proper position. two units of packed red blood cells were transfused to treat the drop in hemoglobin. the following day, the patient lost pulse in the leg. the pump was removed and both a thrombectomy and fasciotomy were performed to treat the condition. the patient was noted to be in critical condition. | [5 empty] | The investigation for the thrombus, hemolysis, and limb ischemia has been completed. data logs were reviewed finding no motor current (mc) spike observed outside p-levels change. no positioning issue observed. suction alarms occurred before/after reported hemolysis issue. no indications found in the logs. the cause of the hemolysis, thrombus, and the limb ischemia issue cannot be determined since compliant device not returned for analysis and lack of clinical data. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: impella catheter too far in the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine. if the impella catheter is too far into the left ventricle, echocardiography will likely show the following: catheter inlet area more than 4 cm below the aortic valve. catheter outlet area across or near the aortic valve.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
14173 | 1220648-2024-23714 | 20740460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20740460&pc=OZD | 10/26/2024 | 12/18/2024 |  | 11/20/2024 |  |  | 10/26/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560983 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed hemolysis during impella cp support. the patient's urine had a red tinge and the plasma free hemoglobin levels were measured to be 117. continuous renal replacement therapy was initiated in response to the condition with discussions of escalation should the condition not improve. support with the implanted pump continued following the intervention. the patient survived the event. | [5 empty] | The investigation for the hemolysis has been completed. data logs were reviewed which showed pump was in improper position corresponded with the time hemolysis was reported per clinical description but resolved quickly. logs also showed multiple intermittent suction alarms were triggered during the case. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14174 | 1220648-2024-23701 | 20740675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20740675&pc=OZD | 06/21/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 06/21/2021 | 11/20/2024 |  |  | 06/30/2021 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1429748 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella device for mechanical circulatory support. pump advanced to aortic valve [av] but had difficulty crossing. after multiple attempts decision was made to abort procedure. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14175 | 1220648-2024-23725 | 20741017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741017&pc=OZD | 09/11/2024 | 11/04/2025 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518573 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient noted with elective placement was implanted with an impella 5.5 device for mechanical circulatory support. dring support, impella pump had positioning issues in addition to false aortic position alarm. placement signal alarms were silenced. no patient harm was reported. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | G1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised on manufacturer device report in accordance with updated procedures. | The investigation of the reported failure mode has been completed. optical signal issue: per the clinical, false "impella position in aorta" alarms were observed and that the staff decided to silence placements signal alarm. the logs do not show any motor current spikes prior to impella position in aorta alarms. the 5s parameters were stable and within range. the cause of the false alarm optical signal issue is likely as a result of the patient's condition b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23725. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23725. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23725. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-23725. h6 health effect - impact code was erroneously selected on the initial manufacturer device report number 1220648-2024-23725.
14176 | 1220648-2024-23723 | 20741025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741025&pc=OZD | 09/11/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511452 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. impella insertion procedure was aborted due to patient anatomy. staff noted an inability to pass through patient's vasculature. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided. as noted in the impella ifu impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14177 | 1220648-2024-23724 | 20741071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741071&pc=OZD | 09/11/2024 | 11/20/2024 |  | 11/20/2024 |  |  | 09/11/2024 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529787 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, pump encountered delivery issues. during patient support, the wire-guided advancement revealed a severe calcified stenosis at the y-bifurcation. the impella could not be advanced to the left ventricle (lv) and was explanted. implant procedure was aborted due to patient anatomy. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition since patient had a severe calcified stenosis at the y-bifurcation. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14178 | 1220648-2024-23726 | 20741376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741376&pc=OZD | 12/23/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 12/23/2022 | 11/20/2024 |  |  | 10/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023268590 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support the patient developed a hematoma at the insertion site. the hematoma was surgical removed. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14179 | 1220648-2024-23727 | 20741405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741405&pc=OZD | 04/10/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 04/10/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2021019985 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella ld device for mechanical circulatory support. while on support, low purge flow and high pressure was noted, a decision was made to start heparin. blood in endotracheal tube was also noted. patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14180 | 1220648-2024-23738 | 20741456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741456&pc=OZD | 03/23/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 03/23/2022 | 11/20/2024 |  |  | 01/31/2024 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023174271 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the medical staff restricted the patient's anticoagulation medication due to a gastrointestinal bleed. the patient remained on support and weaning was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14181 | 1220648-2024-23741 | 20741502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741502&pc=OZD | 09/30/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/30/2022 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022047652 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support, bleeding was a concern, so heparin was withheld in both purge and systemic. the patient received fresh frozen plasma, cryo, packed red blood cell along with desmopressin acetate. successful coronary artery bypass grafting x 2 with intra-aortic balloon pump support was completed. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14182 | 1220648-2024-23728 | 20741569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741569&pc=OZD | 08/22/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 08/22/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2021020020 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support purge pressure close to 1000 and running plain d5 due to bleeding was noted. an episode of supraventricular tachycardia (svt) was also noted. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14183 | 1220648-2024-23729 | 20741581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741581&pc=OZD | 04/23/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 04/23/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1475275 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, it was noted that no systemic heparin was given due to chest tubes being pulled, there was some bleeding that subsided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14184 | 1220648-2024-23731 | 20741582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741582&pc=OZD | 05/26/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 05/26/2021 | 11/20/2024 |  |  | 07/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1494066 |  | 005082 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella ld device for mechanical circulatory support. while on support, the yellow hub of the purge side arm broke off inside the purge tubing. it was also noted that the systemic heparin turned off due to bleeding from site where epicardial pacing leads were removed. | [5 empty] | The investigation for the reported purge leak issue has been completed. the product was returned, and the issue was not confirmed. the product was returned for review, but no sidearm was returned. it is unknown where the sidearm was damaged. per the results of the clinical review and investigation, the root cause was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14185 | 1220648-2024-23735 | 20741584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741584&pc=OZD | 01/15/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 01/15/2022 | 11/20/2024 |  |  | 03/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022059073 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the device could not be delivered because the shaft was kinked. the impella 5.0 placement was aborted, and the patient was implanted with an impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14186 | 1220648-2024-23733 | 20741594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741594&pc=OZD | 07/20/2022 | 12/03/2025 |  | 11/20/2024 |  |  | 07/20/2022 | 11/20/2024 |  |  |  |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022148742 |  | 005082-RO |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. it was reported that the pump was not streaming to the aic (automated impella controller). the ifu was followed and the backup controller was used. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. data analysis: the rlm was booted up on 8/10 with an ssid, was put into maintenance mode, and the rlm rebooted a couple times. after 20:19:11 when the rlm rebooted, the ssid was blank for all of the logs through 9/1 so without an ssid, the rlm would never connect to a wi-fi. root cause: it could not be determined from the logs why the configuration was removed and the console was not returned for investigation. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | D2 common device name revised on manufacturer device report in accordance with updated procedures. d6a and d6b should have been left blank on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
14187 | 1220648-2024-23734 | 20741597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741597&pc=OZD | 11/23/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/23/2021 | 11/20/2024 |  |  | 03/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022116198 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella ld device for mechanical circulatory support. during support, the device exhibited high purge pressure alarms 3 hours after implant. the staff tried to de-air the device without success. the p-0 was reset, and this resolved the issue temporally, however the alarms resumed after 2 hours. the physician decided to only administer the patient a half dose of heparin due to this issue. the resolution is unknown. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14188 | 1220648-2024-23730 | 20741641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741641&pc=OZD | 09/22/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 09/22/2021 | 11/20/2024 |  |  | 10/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2021019997 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support there was hemolyzing and surgical bleeding. a dialysis line was placed at the bedside to start continuous renal replacement therapy. heparin was removed due to the surgical bleed. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14189 | 1220648-2024-23743 | 20741653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741653&pc=OZD | 04/28/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 04/30/2021 | 11/20/2024 |  |  | 09/30/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1504064 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella device for mechanical circulatory support due to cardiomyopathy. while on support it was noted that the patient was going to operating room for exploration for bleeding. patient was resting on support, however the chest remained open. the chest tube drainage was acceptable and prbcs were transfused for overall surgical bleeding. in addition, the impella was weaned. pump was explanted at the same time as closure of the chest. in addition the patient experienced a stroke. the procedure outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14190 | 1220648-2024-23740 | 20741672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741672&pc=OZD | 09/22/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/22/2022 | 11/20/2024 |  |  | 12/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023162528 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced hemolysis. three units of packed red blood cells were transfused to the patient and the impella device was repositioned, however, this effort to manage the complication was unsuccessful and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14191 | 1220648-2024-23742 | 20741679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741679&pc=OZD | 11/23/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 11/23/2022 | 11/20/2024 |  |  | 03/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022070990 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. during support it was reported that the patient had increased amounts of bleeding. the patient was given 8 units of packed red blood cells along with other blood products. heparin was withheld. in addition, the patient had numerous runs of ventricular tachycardia requiring defibrillation and medication. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14192 | 1220648-2024-23732 | 20741691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741691&pc=OZD | 11/16/2021 | 11/20/2024 |  | 11/20/2024 |  |  | 11/16/2021 | 11/20/2024 |  |  | 08/31/2022 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1502420 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support the patient experienced a subarachnoid hemorrhage. additionally, it was reported that the staff failed to disconnect the yellow luer. the staff believed that the purge cassette was leaking, and this issue could be resolved by changing the cassette, however the resolution for this issue is unknown. the patient was bridged to a left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14193 | 1220648-2024-23736 | 20741694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741694&pc=OZD | 09/19/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/19/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136260 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced persistent bleeding requiring an infusion of multiple blood products to be administered. the patient was transferred to unit for exploration of the chest and a clean out. it was reported that upon opening the chest, two to three liters of blood and thrombus were observed. it was stated that a warm saline performed to manage the complication. it was stated that no bleeding from the heart was noted by the medical staff. however, it was stated that a sternum fraction was the potential cause of the bleed. it was reported that various areas around the sternum were cauterized, the fracture was wired in an effort to stabilize. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14194 | 1220648-2024-23737 | 20741702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741702&pc=OZD | 02/22/2022 | 10/13/2025 |  | 11/20/2024 |  |  | 02/22/2022 | 11/20/2024 |  |  | 12/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023158660 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient was awake and responsive, slightly delirious, and pulled the impella out of position twice. the second time the patient damaged the repositioning sheath, and there was no option to exchange the impella pump as the sheath and device were fixed in position. the icu team moved all components of the impella out of the patient's reach so that this could not happen again. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine the root cause. based on clinical account, the patient¿s movement was the cause of the damaged repositioning sheath. this report is being filed as part of a retrospective review of historical records. | Based on clinical account, patient movement was the cause of the damaged repositioning sheath. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation findings and investigation conclusions was erroneously entered on the initial manufacturer device report number 1220648-2024-23737.
14195 | 1220648-2024-23739 | 20741703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20741703&pc=OZD | 09/08/2022 | 11/20/2024 |  | 11/20/2024 |  |  | 09/08/2022 | 11/20/2024 |  |  | 08/31/2023 |  | 11/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136290 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced nasal and oral bleeding, therefore, systemic heparin was held. the patient remained on support and the weaning was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14196 | 1220648-2024-23451 | 20717953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717953&pc=PYX | 09/04/2024 | 11/18/2024 |  | 11/19/2024 |  |  | 09/04/2024 | 11/19/2024 |  |  | 10/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274758 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, it was found that the patient's left and right groins were oozing. the left groin was addressed with mattress sutures x4, and the right groin was addressed with a femstop. the patient ultimately expired, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was most likely use related, mattress sutures were placed to achieve hemostasis as per the clinical description provided. as noted in the impella ifu: access site bleeding: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
14197 | 1220648-2024-23448 | 20718310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20718310&pc=OZD | 09/04/2024 | 11/18/2024 |  | 11/19/2024 |  |  | 09/04/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444576 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient receiving elective placement of a coronary artery bypass graft was attempted to be implanted with an impella cp for mechanical circulatory support. however, the impella implantation was aborted due to the patient's anatomy. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition since insertion of pump was aborted due to patient's anatomy. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14198 | 1220648-2024-23453 | 20718521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20718521&pc=OZD | 12/10/2021 | 11/18/2024 |  | 11/19/2024 |  |  | 12/10/2021 | 11/19/2024 |  |  | 05/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022094457 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that the pump would not work. no further information is available. | [5 empty] | The investigation into the reported issue has been completed. the impella 5.0 was returned for investigation. the visual inspection of the returned device found that the outflow cage was bent and pulled away from the motor housing. the pump started and ran without issue in the lab. the root cause of the issue was most likely due to use resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
14199 | 1220648-2024-23433 | 20718726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20718726&pc=OZD | 07/12/2022 | 05/30/2025 |  | 11/19/2024 |  |  | 07/12/2022 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXXX |  | 005060 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant called support and reported that a red controller alarm was generated and was not visible due to an automated impella controller (aic) connection failure. the local team was informed and assisted the end user. no further information was provided. there was no patient harm reported. | [5 empty] | D4 catalog number, expiration date, unique identifier and h4 manufacture date are unknown as the product model number was not supplied. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the automated impella controller (aic) device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. d6a / d6b implant and explant date should be blank since the automated impella controller (aic) is not an implantable device.
14200 | 1220648-2024-23452 | 20719446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20719446&pc=OZD | 12/16/2021 | 11/25/2024 |  | 11/19/2024 |  |  | 12/16/2021 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121649 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the device exhibited an outflow blocked alarm and catheter migration alarm. the resolution for this issue is unknown. it was reported that after 41.50 hours of impella 5.0 support the family decided to withdraw care. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Correction: the h1 ( type of reportable event ) was inadvertently reported as malfunction in the initial report which is corrected in this report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14201 | 1220648-2024-23454 | 20719494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20719494&pc=OZD | 12/09/2021 | 11/18/2024 |  | 11/19/2024 |  |  | 12/09/2021 | 11/19/2024 |  |  | 09/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022139393 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. the complainant reported that during impella placement, the pump experienced low pump flows. the purge cassette was replaced and the purge solution changed from dextrose with five percent water to dextrose with twenty percent water. weaning was successful, the device was explanted, and the procedural outcome was the patient survived explant. | [5 empty] | The investigation into the low pump flow/low purge pressure issue has been completed. the impella 5.0 was returned for investigation. the original cassette that was reportedly switched out to resolve the low purge pressure (s/n (b)(6) l/n 1604638) was not returned for analysis. the pump was returned and tested with a lab cassette and was able to purge with expected pressures. the root cause of the low purge pressure was most likely a leak in the first cassette. without cassette evaluation, the leak location is unknown. the clinical details provided were insufficient to determine a root cause and the cassette that was in use was not returned, therefore a root cause of the issue was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14202 | 1220648-2024-23457 | 20720171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20720171&pc=OZD | 12/25/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/25/2021 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119934 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on an impella ld for hemodynamic support. while on support the patient experienced bleeding issues and it was reported that the patient may have lost over 300 milliliters of blood, and the device exhibited a suction alarm. the patient was administered bicarb in the purge, 1 unit of packed red blood cells, 2 units of fresh frozen plasma, 7 units of albumin, 2000 fieba, and b12. additionally, the patient experienced, atrial fibrillation (a-fib) with no specified resolution. it was reported that the patient survived normal explant of the device and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14203 | 1220648-2024-23445 | 20720263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20720263&pc=OZD | 09/05/2024 | 11/18/2024 |  | 11/19/2024 |  |  | 09/05/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512438 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the external bypass became clotted, and the distal limb was found to be occluded after impella removal. no resolution was specified, but it was reported that the patient survived. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia could not be determined. as noted in the impella ifu: limb ischemia: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14204 | 1220648-2024-23458 | 20720348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20720348&pc=OZD | 08/18/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/18/2021 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022097242 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced retroperitoneal bleeding at the insertion site. it was explained to the medical staff that heparin could be temporarily stopped in an effort to control the bleeding. it was noted that there have been no particular problems with the pump and that the device was removed and replaced with an implantable ventricular assist device (vad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14205 | 1220648-2024-23455 | 20720506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20720506&pc=OZD | 12/27/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/27/2021 | 11/19/2024 |  |  | 08/31/2021 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131611 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the impella 5.0 migrated with an outflow blocked alarm intermittently activated. the patient is fully left ventricular assist device/right ventricular assist device dependent and has no pulse pressure. an echocardiogram was performed at bedside and the impella 5.0 was repositioned. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14206 | 1220648-2024-23456 | 20720760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20720760&pc=OZD | 12/26/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/26/2021 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022085150 |  | 005082 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella for mechanical circulatory support. the complainant reported that during an attempt to place the impella, the device cannula could not be advanced due to a kinked catheter. it was reported that the process of impella implantation was aborted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the delivery issues was determined to be the confirmed catheter kinks as a result of improper technique. this report is being filed as part of a retrospective review of historical records.
14207 | 1220648-2024-23459 | 20722150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722150&pc=OZD | 10/24/2024 | 07/09/2025 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559867 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The complainant reported a patient was on impella cp support. while on support, the patient developed hemolysis. continuous renal replacement therapy was started as creatinine continued to climb. hemolysis was resolved. the pump was successfully weaned and the patient survived the explant. | A notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that the patient had ventricular arrhythmia occur with a definite relationship to the impella cp used on this patient. no further details were provided on the treatment of the adverse event. |  |  |  |  | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the investigation results for hemolysis was previously reported on manufacturer device report number 1220648-2024-23459-1. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B5-additional information provided for new failure mode of ventricular arrhythmia. the impella device was received and has been evaluated. should any new information be received for the new failure mode of ventricular arrhythmia, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. data logs were confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in aorta. the positioning issue then was resolved & pump ran without issues to the rest of the case. pump was returned for analysis. visual inspection found no issues on the pump. based on clinical description stated that hemolysis was resolved after repositioning, product inspection found no issue & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email. | G2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23459. h6 health effect - clinical code 1721 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23459-2.
14208 | 1220648-2024-23460 | 20722239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722239&pc=OZD | 10/24/2024 | 12/20/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562069 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella 5.5. activated clotting time was 202. the 0.018 wire was exchanged and impella was attempted to rail up over the 0.018 abiomed wire with difficulty. the decision was made to exchange the .018 wire for an abiomed .027. a clot was noted inside the impella catheter and 10 millimeter graft. the decision was made to abort the procedure and place an intra-aortic balloon pump via the femoral artery. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since insertion was aborted due to patient's anatomy. the root cause of the thrombus could not be determined without additional clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist : section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14209 | 1220648-2024-23461 | 20722305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722305&pc=OZD | 10/24/2024 | 01/02/2025 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525844 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, the pump showed suction alarm two hours after submission to the intensive care unit. reposition, echocardiography, volume status check was done, but pump had only 0.6 liters per minute on p8. it was decided to explant the pump. the patient survived the explant. | [5 empty] | The investigation for the low pump flow has been completed. the cp pump was returned for evaluation and was visually inspected. biomaterial observed around impeller and purge gap. no broken blade found. no cannula kink observed. no other abnormality was noticed. data logs were reviewed finding motor current (mc) spikes were observed in the data logs prior to the low pump flow. suctions occurred at the reported time of the issue, then mc dropped with drop in flows without change in p-level. the root cause of the low pump flow issue was biomaterial ingestion. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14210 | 1220648-2024-23463 | 20722455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722455&pc=PYX | 10/25/2024 | 11/20/2024 |  | 11/19/2024 |  |  | 10/25/2024 | 11/19/2024 |  |  | 07/31/2025 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024381174 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient post coronary artery bypass graft surgery was undergoing an impella rp implant for right ventricular failure. during the impella implant procedure after the sheath was inserted, the physician noticed a hematoma. the physician made a decision not to continue in the groin with the impella rp and aborted the delivery. manual pressure was applied to the hematoma with a switch to an impella rp flex for mechanical circulatory support via the right internal jugular vein. the patient was noted to be in stable condition. | [5 empty] | Corrections to the initial mfr report: added d6a/d6b. | According to clinical, the impella rp was being placed via the 23fr sheath for an elderly patient with multiple cardiac and systemic comorbidities. the patient had emergent coronary artery bypass graft surgery and found to need the impella rp for right ventricular failure. there was vascular damage done at the insertion of the 23fr sheath. this vessel damage prompted the team to abort the impella rp delivery. the pump, introducer kit was not returned as it was discarded, and therefore, an evaluation of the device was not possible. product history review was completed, and the pump passed all post sterile investigation checks. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided, and no product was returned for analysis. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
14211 | 1220648-2024-23465 | 20722497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722497&pc=OZD | 08/25/2021 | 11/23/2024 |  | 11/19/2024 |  |  | 08/25/2021 | 11/19/2024 |  |  | 11/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022034452 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced gastrointestinal bleed and anticoagulation levels were low. the impella migrated to the aorta and was explanted, and a new 5.0 device was implanted. | It was additionally reported that while on support the patient had an episode of atrial fibrillation with rapid ventricular response. the patient required cardioversion to resolve. it is unknown what the relationship is between the arrhythmia and the impella. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information for the initial mfr 1220648-2024-23465 was provided in sections b5 and h6. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14212 | 1220648-2024-23466 | 20722518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722518&pc=OZD | 09/13/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 09/13/2021 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078619 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt) which required direct current cardioversion (dccv). a bedside echocardiogram noted that the impella inlet was one centimeter away from the aortic valve. through imaging, the pigtail appeared to be bent and potentially moving towards the posterior wall. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14213 | 1220648-2024-23462 | 20722538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722538&pc=OZD | 10/26/2024 | 12/19/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  | 09/30/2025 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399708 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella 5.5 was inserted wirelessly via direct aortic approach. they were uable to achieve appropriate flow. the impella 5.5 was noted to be facing upwards towards the mitral valve. the physician was attempting to torque and advance the catheter further in the left ventricle (lv) space when the tip had poked through the posterior ventricular wall. the impella 5.5 was removed, irrigated, and placed in surgical mode while the lv was repaired. the impella 5.5 was then reinserted without issue via catheter/wire approach and therapy was resumed. the patient remained stable the procedural outcome was successful with this device. | [5 empty] | The investigation for the ventricle perforation has been completed. no product was returned and logs are not applicable. positioning issues led to the lv perforation. the root cause of the ventricle perforation was improper positioning. added h6 med dev prob code 1158, h6 component code 925 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
14214 | 1220648-2024-23467 | 20722574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722574&pc=OZD | 09/02/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 09/02/2021 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022067984 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the impella 5.0 was unable to pass through the ascending aorta. the decision was made to abort the process of placing the impella 5.0 and to implant an impella cp to continue the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14215 | 1220648-2024-23464 | 20722575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722575&pc=OZD | 08/19/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/19/2021 | 11/19/2024 |  |  | 12/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021016277 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. the impella was added for left ventricular venting, with difficult implantation due to kinking in the aortic arch. unable to pass through vasculature and only femoral access was possible. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14216 | 1220648-2024-23468 | 20722683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722683&pc=OZD | 10/18/2024 | 12/20/2024 |  | 11/19/2024 |  |  | 10/22/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552675 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and upon removal of the 5.5, there were no clots noted on the pump. two days after the removal, an embolic clot was found in the left anterior descending artery. the doctor had a suspicion that the clot may have been from the impella removal. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14217 | 1220648-2024-23469 | 20722709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722709&pc=OZD | 10/26/2024 | 12/17/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025563664 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for 17 days when the pump stopped. the stop occurred when the patient stepped on the cable while ambulating down the hall. the pump was explanted. the patient survived the explant. | [5 empty] | The investigation into pump stop has been completed. based on clinical details, data logs, and returned product, the root cause of the pump stop was determined to be a electrical open. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23469. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14218 | 1220648-2024-23471 | 20722824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722824&pc=OZD | 09/21/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 09/21/2021 | 11/19/2024 |  |  | 05/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022083126 |  | 005062 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the patient experienced ventricular tachycardia (vt) as the guidewire entered the left ventricle and cardioversion was required. the patient developed heart block when the impella device entered the left ventricle and a transvenous pacemaker was implanted. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14219 | 1220648-2024-23470 | 20722858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722858&pc=OZD | 12/09/2022 | 01/06/2025 |  | 11/19/2024 |  |  | 10/23/2024 | 11/19/2024 |  |  | 10/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272742 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu in (b)(6) 2022 had a cp pump placed to support a patient admitted in for a high-risk percutaneous coronary intervention. the pump was supporting for just under 2 hours and then was explanted. no harm or injury or malfunction noted. the team had enrolled the patient in the protect iv protocol and in (b)(6) 2024 the documentation was forwarded alleging an adverse event (ae) attributed to the impella. the definite ae was the ck/ck-mb rising and falling which is indicative of acute myocardial infarction. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. b.2 revised selection as it was previously selected incorrectly on the initial manufacturer device report number 1220648-2024-23470. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23470 was submitted in accordance with updated procedures.
14220 | 1220648-2024-23472 | 20722914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722914&pc=OZD | 10/24/2024 | 12/06/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521546 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Puncture/Hole | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support with acute myocardial infarction and cardiogenic shock. it was reported that bleeding was noted at the connection of the purge tubing and the red impella plug, at the junction of the grey plastic to red plastic. blood products were given to the patient. removal and replacement of impella pump was recommended however the medical doctor placed bone wax around the leak. the leak was partially improved but there was still a slow leak. the patient was transferred to the intensive care unit with the impella remaining in place. it was further reported that the patient was taken to the operating room to have the impella removed due to suspected thrombus. thrombectomy was performed via right femoral cut down. per the charge nurse, patient was intubated and on 3 inotropes/vasopressors. the patient outcome is unknown. | [5 empty] | The investigation for the thrombus, product damage, and access site bleed has been completed. product not returned. bleeding reported from the red handle. the cause of the access site bleeding issue was not determined since product was not returned and a lack of clinical details. the cause of the product damage hole in catheter was not determined since product was not returned to confirm location of bleed/damage and due to a lack of clinical details. the root cause of the thrombus could not be determined without additional information. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14221 | MW5162620 | 20722941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20722941&pc=OZD | 11/14/2024 | 11/15/2024 |  | 11/19/2024 |  |  |  | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 0048-0014 |  |  | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Not Available | Unknown | [6 empty] | 75 YR |  | Female | 55 KG | Hispanic | White | Death | Yes | No | [7 empty] | Exsanguination | Cardiac Perforation | [11 empty] | Death | [28 empty] | Initial submission | [9 empty] | An impella cp was placed via the left femoral artery for cardiogenic shock. the device perforated the left ventricle. the patient was taken to the operating room for emergency surgery and repair. she expired from uncontrollable bleeding despite repair.
14222 | 1220648-2024-23473 | 20723265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723265&pc=OZD | 10/24/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552321 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support an acute myocardial infarction and cardiogenic shock patient. the cp was pulled back across the valve in error and the medical doctor was unable to re-deliver to the left ventricle. the pump was explanted and replaced. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique as the impella pulled back across valve by the physician and was unable to recross. corrections to the initial mfr report: g1 mdr reporting contact email. h6 impact code 4627 changed to 4629. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
14223 | 1220648-2024-23474 | 20723445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723445&pc=OZD | 10/24/2024 | 12/24/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480603 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. it was reported that the doctor attempted to implant cp, but could not get impella past calcium of aortic bifurcation/iliac. the impella was removed and shockwave treatment was completed at the bifurcation, then ballooned. multiple attempts to redeliver the cp performed with no success. the case was aborted and no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella could not get past calcium of ao bifurcation/iliac as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14224 | 1220648-2024-23476 | 20723547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723547&pc=OZD | 03/07/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/07/2022 | 11/19/2024 |  |  | 12/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023167605 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from gastrointestinal area. the patient was closely monitored. liver failure was a noted contributing factor. anticoagulants were optimized. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14225 | 1220648-2024-23475 | 20723565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723565&pc=OZD | 10/25/2024 | 12/20/2024 |  | 11/19/2024 |  |  | 10/25/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559757 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had severe bleeding in the access site. the sutures were removed and redone. the bleeding was resolved. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by redoing sutures. corrections to the initial mfr report:g1 mdr reporting contact email.
14226 | 1220648-2024-23267 | 20723728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723728&pc=OZD | 12/30/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/30/2021 | 11/19/2024 |  |  | 02/28/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022053035 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during insertion a purge line blockage was suspected. in addition, the patient required blood transfusion due to a vascular injury after removal of the 5.0. it was noted that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14227 | 1220648-2024-23477 | 20723819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723819&pc=OZD | 10/24/2024 | 12/16/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560037 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed to support a acute myocardial infarction and cardiogenic shock patient. the pump was placed but, limb ischemia occurred. pulses were lost and so distal perfusion catheter was placed, and when this did not fully resolve, the team explanted the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23477. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14228 | 1220648-2024-23478 | 20723830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723830&pc=OZD | 02/10/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/10/2022 | 11/19/2024 |  |  | 12/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2023161636 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella ld device for mechanical circulatory support. while on support, there was bleeding from the red impella plug, and a placement signal failure was noted. the pump was removed due to blood slowly leaking from the red impella plug and placement signal failure. patient was successfully weaned. | [5 empty] | The investigation for the reported placement signal and hole in catheter issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14229 | 1220648-2024-23479 | 20723952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723952&pc=OZD | 10/24/2024 | 12/20/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. the procedure was aborted due to patient anatomy. the patient had severe diseased and calcified vascular disease in the femoral and iliac arteries. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since insertion was aborted due to patient's anatomy. g.1 revised reporting contact email since the initial manufacturer device report was submitted in accordance with updated procedures. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use:impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states)¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controllersection: warnings & cautions: cautions¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿section: warnings & cautions: warnings¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿impella cp with smartassist system with automated impella controllersection: warnings & cautions: warningssection: pre-support evaluationsection: axillary insertion of the impella 5.5 cathetersection: direct aortic insertionsection: use of impella with transcatheter aortic valves¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿section: warnings & cautions: warnings¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14230 | 1220648-2024-23480 | 20723956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723956&pc=PYX | 08/14/2021 | 10/27/2025 |  | 11/19/2024 |  |  | 08/14/2021 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1231459 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA RP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported while on impella rp support, the automated impella controller optical signal issues that were resolved by taking x-rays to check impella positioning. there was no reported patient injury, and the patient was successfully weaned and explanted. | [5 empty] | D4: the automated impella controller serial information was not provided, therefore serial number, lot number, product number, and udi are unknown. h4: the automated impella controller serial information was not provided, therefore the manufacturing date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Console logs from the reported day of event do not contain any information that would either confirm the reported issue or allow to identify any other issue the complaint might have referred to. presumably, the issue was related to impellaconnect portal (data from the console being unavailable), but aic logs show that usb data transfer from aic to rlm was unaffected. there were no rlm journals from rl03968 available for analysis. there were also no issues with placement signal during the case, and there¿s no evidence that would support theory of it not being visible on aic screen. this console¿s rlm was replaced during pm on (b)(6) 2023. there were no complaints involving either of the rlms on this console. b7 other relevant history, including preexisting medical conditions updated. d2 common device name updated. d4 serial number, lot number, and unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated. d6a and d6b should have been left blank. g1 reporting contact email updated. h4 device manufacture date updated. h6 corrected and updated.
14231 | 1220648-2024-23481 | 20723961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20723961&pc=OZD | 03/03/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/03/2022 | 11/19/2024 |  |  | 10/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022047656 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella ld device for mechanical circulatory support. during implantation, the delivery failed due to the patient¿s anatomy. a new impella device was used without any harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient anastamosis anatomy related. replacement impella 5.5 pump was inserted successfully. this report is being filed as part of a retrospective review of historical records.
14232 | 1220648-2024-23482 | 20724050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724050&pc=OZD | 10/08/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 10/08/2021 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075904 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a purge system open alarm was observed, and the purge pressure and flow decreased significantly. the purge system was inspected, and it was noted that the purge side arm was damaged. upon troubleshooting and the purge pressure increasing, the decision was made by the physician to proactively explant the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14233 | 1220648-2024-23484 | 20724134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724134&pc=OZD | 07/20/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 07/20/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022148753 |  | 005082 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella device for mechanical circulatory support. while on support, after extubation, an outflow blocked alarm suddenly sounded. they were unable to tell if the pump was deep or shallow. the following day the impella pulled back; however, still showing deep. the outflow blocked alarm intermittently. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14234 | 1220648-2024-23483 | 20724299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724299&pc=OZD | 10/24/2024 | 12/12/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump to support a patient admitted in with cardiogenic shock and cardiomyopathy. the patient developed bleeding and hematoma at the impella site. the blood loss prompted intervention in form of washout and stopping heparin. improvement after giving fresh frozen plasma. the patient remains on support therefore explant date and outcome is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management since bleeding resolved with stopping systemic heparin. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14235 | 1220648-2024-23488 | 20724327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724327&pc=OZD | 08/31/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/31/2022 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022085148 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Atrial Fibrillation | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the device exhibited purge pressure high alarms, this issue was not resolved. additionally, the patient went into atrial fibrillation and was administered an amiodarone iv bolus. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14236 | 1220648-2024-23485 | 20724394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724394&pc=OZD | 10/24/2024 | 12/09/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546744 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had a run of ventricular tachycardia (vt) when trying to sit up. bedside echocardiography showed the impella clear from any structures and in the proper position but a little bit shallow. vt resolved after lying the patient back down in bed. about three weeks later, the patient begn to acutely decompensate. the patient went in to vt arrest and cardiopulmonary resuscitation was started. the patient did not achieve return of spontaneous circulation and eventually passed away. | [5 empty] | H6 health effect - impact code 1802 missing from manufacturer device report 1220648-2024-23485. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14237 | 1220648-2024-23493 | 20724404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724404&pc=OZD | 09/03/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 09/03/2022 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119935 |  | 005082-RO |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient developed a gastrointestinal bleed and heparin was withheld. additionally, the patient was taken to the operating room for a wash out and a large posterior clot was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14238 | 1220648-2024-23486 | 20724423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724423&pc=OZD | 10/14/2021 | 03/17/2025 |  | 11/19/2024 |  |  | 10/14/2021 | 11/19/2024 |  |  | 05/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022085045 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. during delivery of the device, implant was aborted because it was not possible to advance the impella into the patient's left ventricle due to anatomical conditions. no further information is available. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14239 | 1220648-2024-23489 | 20724445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724445&pc=OZD | 10/16/2021 | 03/17/2025 |  | 11/19/2024 |  |  | 10/16/2021 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022064736 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced access site bleeding. the staff applied pressure to the site and withheld heparin to help resolve the issue. it was reported that the patient was stable, survived normal explant and the procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14240 | 1220648-2024-23487 | 20724484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724484&pc=OZD | 08/13/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/13/2022 | 11/19/2024 |  |  | 02/28/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2023201083 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient developed a nosebleed and heparin was restarted. additionally, the patient experienced atrial fibrillation and was administered amiodarone. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14241 | 1220648-2024-23496 | 20724503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724503&pc=OZD | 10/11/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 10/11/2022 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119932 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced a large left hemothorax and was bleeding, heparin was removed from the purge and was administered 7 units of packed red blood cells. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14242 | 1220648-2024-23492 | 20724510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724510&pc=OZD | 07/08/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 07/08/2021 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471970 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) produced a "control unit failure" alarm. the alarm quickly resolved on its own, though the position waveform was not available when the alarm appeared. there was no patient harm reported for the reported issue, and no further intervention was needed for the aic. the patient later expired due to acute heart failure. no allegations were made towards the impella for cause of death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. the device logs confirmed the reported failure mode. console was tested thoroughly by running a known-good pump and purge unit, and the failure was not reproduced. the event history logs were reviewed and confirmed the reported failure modes. aside from the purge-related issues which were due to the purge cassette, the console also issued a controller error due to bad opm digital communication, resulting in an invalid placement signal. real time event logs confirmed the opm-related signals went invalid to a value of -16k. the cause of the controller error (opm communication crc error) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
14243 | 1220648-2024-23490 | 20724562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724562&pc=PYX | 10/24/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529631 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support when there was a clot ingested. there was evidence by increased plasma free hemoglobin, increased motor current and decreased flows. the clinical rep confirmed the patient had tea colored urine and the plasma free hemoglobin was 800. the doctor decided to start milrinone, monitor and explant the rp. the pump was weaned, explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the hemolysis, thrombus, and low pump flow has been completed. product was not returned for evaluation. data logs show numerous motor current (mc) spikes throughout support with the biggest. with the mc spikes there is drop in flow, speed deviations, and increased ps shifts. activated clotting time (act) was said to not have been maintained with subtherapeutic partial thromboplastin time (ptt). the root cause of the low pump flow was most likely biomaterial. the root cause of the hemolysis was most likely biomaterial. the root cause of the thrombus could not be determined without additional clinical details. the instructions for use for the impella rp smartassist system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14244 | 1220648-2024-23491 | 20724632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724632&pc=OZD | 11/25/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 11/25/2021 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022114002 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the pump stopped and was impossible to restart. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14245 | 1220648-2024-23495 | 20724682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724682&pc=OZD | 09/29/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 09/29/2022 | 11/19/2024 |  |  | 12/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2023161638 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient developed bleeding at the access site, heparin was withheld, and the patient was administered multiple blood products. it was reported that the device was normally explanted and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14246 | 1220648-2024-23499 | 20724700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724700&pc=OZD | 10/23/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 10/23/2022 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119938 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella for mechanical circulatory support. during support, the patient experienced clots found in the chest and active bleeding at the graft sites. blood products were administered. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14247 | 1220648-2024-23498 | 20724701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724701&pc=OZD | 12/11/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/11/2021 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022095582 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. prior to the implantation, the patient was listed in not good condition. during the implantation, they were not able to enter the wire into the left ventricle, and after 1 hour, the cardiac team made the decision to abort the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14248 | 1220648-2024-23503 | 20724726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724726&pc=OZD | 12/02/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 12/02/2022 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022085141 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella device for mechanical circulatory support. during support, the patient experienced bleeding in the chest, a wound vac was placed, and 5 units of red blood cells were administered. additionally, the patient's internal pacer stopped capturing the later third, causing ventricular tachycardia arrest. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14249 | 1220648-2024-23497 | 20724739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724739&pc=OZD | 11/24/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 11/24/2021 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022118186 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the pump was impossible to place in the left ventricle due to the vascular condition of the patient. the procedure was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14250 | 1220648-2024-23494 | 20724749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724749&pc=OZD | 10/24/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548416 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | Death |  |  |  |  | CONTINUOUS RENAL REPLACEMENT THERAPY | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a patient with cardiogenic shock and cardiomyopathy. the pump was placed in the operating room as escalation from the intra-aortic balloon pump and cardiopulmonary resuscitation. the pump was placed and there were concerns of low hemoglobin and hematocrit/hypovolemia and hemolysis. the patient was treated with continuous renal replacement therapy and blood product was infused. the pump remained on until the family/team made the choice to withdraw care and the patient expired. | [5 empty] | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the data log were reviewed and there were many suction alarms that occurred. many unknown position alarms observed pre the reported event days. no motor current spike or clear positioning issue. logs indicated placement signal being noisy during the case. the root cause of the hemolysis issue most likely patient condition related. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23494. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14251 | 1220648-2024-23501 | 20724757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724757&pc=OZD | 10/26/2024 | 12/16/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553100 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the patient had limb ischemia. a distal perfusion catheter was placed and support continued. the patient had continued limb ischemia, and the pump was eventually explanted. upon explant, a clot was noted distally. a thrombectomy was performed. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-23501 was submitted.
14252 | 1220648-2024-23500 | 20724850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724850&pc=OZD | 11/17/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 11/17/2022 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022119401 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced bleeding and heparin was withheld. additionally, the patient developed a clot, and it was evacuated. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14253 | 1220648-2024-23509 | 20724913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724913&pc=PYX | 10/26/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025560981 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Unspecified Vascular Problem | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the current impella rp was weaned and explanted with wire access maintained through the cannula. a new sheath was inserted, and patient was monitored. numbers began to deteriorate so the team opted to place a new impella rp even weighing anticoagulation issues with argatroban and patient being heparin-induced thrombocytopenia positive. new rp was inserted with no issues. pa line and 0.027 wire were in the right pulmonary artery and pump ended up in an ideal position. flows were 3.9 l/m at p8, and the patient¿s pressures improved immediately. when prepping her for transport to the intensive care unit (icu), frank blood was noticed in the endotracheal tube. the physician noticed significant clot and tried for 1 hour to remove them. the patient deteriorated rapidly. unable to do cpr due to unknown bleeding in the lungs. team did not see any path for further escalation and brought the patient to the icu. the patient expired in the icu. | [5 empty] | The investigation for the thrombus and vascular damage has been completed. no product was returned for analysis. clinical details were not sufficient to determine the root cause. the root cause of vascular damage and thrombus could not be determined as no product was returned and limited clinical information was provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿.
14254 | 1220648-2024-23505 | 20724945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724945&pc=OZD | 01/24/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/24/2022 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022099326 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, purge flows [pf] dropped below 3ml/hr. bedside nurse-initiated tissue plasminogen activator (tpa) to increase pf and decrease purge pressure. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14255 | 1220648-2024-23506 | 20724977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20724977&pc=OZD | 07/09/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 07/09/2021 | 11/19/2024 |  |  | 12/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037257 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a computed tomography (ct) revealed cerebral hemorrhage and the patient cannot be on heparin. no additional information was provided. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14256 | 1220648-2024-23502 | 20725014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725014&pc=OZD | 12/25/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/25/2021 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075866 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the purge flow was high at 30ml/hr. purge pressure 330. they changed the purge cassette after getting low purge pressure alarms. the nurse denies any leaks and will monitor closely. the patient remained on support. | [5 empty] | D6b explant date unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14257 | 1220648-2024-23504 | 20725042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725042&pc=OZD | 12/27/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 12/27/2021 | 11/19/2024 |  |  | 08/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021009577 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. upon implantation, the cardiac surgeon attempted ld insertion via axillary route utilizing axillary kit with open chest. the physical limitations of aortic arch precluded anastomosis of graft with usual technique. the cardiac surgeon was aware this is an off-label technique: however, the limitations of the operating room did not allow for use of the fluoroscopy to deliver wire for impella 5.0. the device would not advance beyond the graft anastomosis and resulted in the ld catheter severely kinked in two locations proximal to the pump motor. the catheter was removed without having been started. ohs team decided to attempt wired insertion of impella 5.0 without fluoroscopy using trans-esophageal echo guidance. wire was delivered successfully utilizing pigtail and .035 wire/exchange. at that point impella 5.0 was prepared for insertion; the cardiac surgeon misaligned grey connector and applied enough force to bend connection pin within plug. the cardiac surgeon was unconfident of bending pin back to original shape. ultimately, the alternative of inserting an impella cp was pursued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14258 | 1220648-2024-23513 | 20725065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725065&pc=OZD | 01/22/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/22/2022 | 11/19/2024 |  |  | 09/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022133885 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a high purge pressure alarm was noted. dextrose concentration was decreased to g2.5 %. it was noted that purge pressure was still high. echocardiogram confirmed impella was too far in left ventricle [lv]. impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14259 | 1220648-2024-23507 | 20725169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725169&pc=OZD | 10/28/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 10/28/2021 | 11/19/2024 |  |  | 11/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022030173 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the impella pump was explanted due to a bacterial infection. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14260 | 1220648-2024-23508 | 20725196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725196&pc=OZD | 01/09/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/09/2022 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022106298 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during support, multiple bleeding issues were noted. as a result, heparin was held from purge solution. manual pressure was also held to achieve hemostasis. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14261 | 1220648-2024-23510 | 20725236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725236&pc=OZD | 01/10/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/10/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129476 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during procedure, staff noted a large amount of blood loss from the pocket (750ml). patient was transferred to intensive critical unit (icu). physician repositioned the impella pump. a sandbag was placed to access site for some swelling. access site looked good at last check. patient survived the weaning procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14262 | 1220648-2024-23511 | 20725272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725272&pc=OZD | 01/11/2022 | 03/17/2025 |  | 11/19/2024 |  |  | 01/11/2022 | 11/19/2024 |  |  | 12/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022157571 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Complainant reported that after inserting the 18 wires, the pump insertion procedure was tried, but the guidewire came out to the vicinity of the a valve, so it was performed again from the pig insertion procedure. in the meantime, the pump was removed and turned with distilled water. it was later noted that the discharge blockage alarm did not go away even after the pump position was adjusted (apparently the position was adjusted quite shallow), and although auxiliary flow was being obtained, blood pressure was not rising (based on the flow of the conversation, it does not appear that an ultrasound evaluation was performed). it was determined that ar was occurring, and impella was removed. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14263 | 1220648-2024-23517 | 20725290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725290&pc=OZD | 10/29/2024 | 12/19/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539064 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 25-year-old male patient had an attempted impella 5.5 implant however, the artery could not dilate and the patient's vessel anatomy was fragile. multiple wires were attempted. the implant was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, pt vessel anatomy fragile as per the clinical description provided. added h6 impact code 4627 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14264 | 1220648-2024-23521 | 20725321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725321&pc=OZD | 02/24/2022 | 05/21/2025 |  | 11/19/2024 |  |  | 02/24/2022 | 11/19/2024 |  |  | 02/28/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022053039 |  | 005066 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support during support, the patient experienced thrombus and it was removed. it was reported that the patient survived normal explant. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14265 | 1220648-2024-23514 | 20725388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725388&pc=OZD | 10/29/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  | 04/30/2025 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330682 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was further reported that the procedure was aborted and the patient outcome is unknown. | The complainant reported the 62-year-old female patient had an attempted impella 5.5 implant and there was a struggle to insert the pump and it was aborted due to patient anatomy. there was a small dissection noted that required balloon angioplasty. |  |  |  |  | The investigation for the vascular damage and delivery issues has been completed. the pump was not returned for investigation. a data log review was not required for this case run. as per the given clinical information, several unsuccessful attempts were made to insert the pump. the patient was reported to have severe arterial disease. vascular damage was observed during the procedure and was successfully repaired with balloon angioplasty. the cause of the delivery issue was most likely patient condition related to diseased/small vessels, based on given clinal details. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. added b5-patient outcome. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14266 | 1220648-2024-23518 | 20725453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725453&pc=OZD | 10/29/2024 | 12/30/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485924 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella cp was explanted a thrombus was noted in the femoral artery. penumbra was passed down the right femoral artery two times. residual clot noted and a balloon was inflated in the right femoral artery. flow was restored down the right lower extremity. manual pressure was held for 15 minutes after the balloon was deflated. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email.
14267 | 1220648-2024-23520 | 20725463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725463&pc=OZD | 02/10/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/10/2022 | 11/19/2024 |  |  | 09/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-JP |  |  | OZD |  |  | IMPELLA 5.0-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022135038 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced sepsis, the resolution for this issue is unknown. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14268 | 1220648-2024-23515 | 20725494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725494&pc=OZD | 10/29/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a 69-year-old male patient admitted in for cabg, avr, mvr and had the impella cp pump placed to escalate from iabp and ecmo. the cp was placed but there was thrombus developed on the mitral on day 7 and the pump failed and was removed/replaced on day 17. the patient still awaiting left ventricular assist device (lvad). | [5 empty] | The investigation for the thrombus and the pump stop has been completed. the pump nor data logs were returned for review. without the pump or additional clinical details, the root cause of the pump stop or thrombus could not be determined. corrections to the initial mfr report: b1 should have had product problem selected as well. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
14269 | 1220648-2024-23519 | 20725623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725623&pc=OZD | 10/24/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 10/24/2021 | 11/19/2024 |  |  | 09/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013260 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient pulled the impella back over 20cm. the graft was intact, but the decision was ultimately made that the patient would probably benefit from completely removing the impella and not reinserting as it was well out of the axillary artery. the patient was successfully weaned, and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14270 | 1220648-2024-23527 | 20725638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725638&pc=OZD | 03/29/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/29/2022 | 11/19/2024 |  |  | 02/29/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023184270 |  | 05060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. it was reported that implantation was aborted to patient anatomy. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14271 | 1220648-2024-23528 | 20725732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725732&pc=OZD | 04/07/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/07/2022 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022099324 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced bleeding and heparin was withheld. additionally, device exhibited a placement signal not reliable. the resolution for this issue is unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14272 | 1220648-2024-23526 | 20725735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725735&pc=OZD | 10/21/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 10/21/2021 | 11/19/2024 |  |  | 07/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1501268 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement, a hematoma was observed at the access site. it is unknown how the complication was managed, or if hemostasis was achieved. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14273 | 1220648-2024-23522 | 20725756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725756&pc=OZD | 08/11/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/11/2021 | 11/19/2024 |  |  | 08/31/2021 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1443136 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that when the medical staff attempted to remove the impella device, the pump would not come out. the patient experienced ventricular tachycardia and was shocked. it was reported that a thrombus was observed in the graft at explant. weaning was successful, the device was removed, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14274 | 1220648-2024-23530 | 20725788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725788&pc=OZD | 10/23/2024 | 01/09/2025 |  | 11/19/2024 |  |  | 10/24/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025511290 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The patient expired on support of the replacement pump. it is noted the patient expired from multiple organ failure; however, the patient required both ecmo and impella therapies and adequate therapy may have been compromised given the event. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported the patient was on impella cp support. it was reported that the patient was moved to the intensive care unit bed and the patient's posture changed causing the impella to be completely removed from the aortic valve. the extracorporeal membrane oxygenation did not come off, so hemodynamics was maintained. therefore, a new impella was reinserted. |  |  |  |  | The following was inadvertently reported incorrectly in manufacturer device report 1220648-2024-23530 b1 should have been both adverse event and product problem b2 death and date of death added b5 updated with patient outcome. h1 type of reportable event changed to death h6 health effect - impact code updated to reflect patient outcome. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was patient movement related. g1 reporting contact email revised on manufacturer device report 1220648-2024-23530 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14275 | 1220648-2024-23534 | 20725794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725794&pc=OZD | 08/13/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/13/2021 | 11/19/2024 |  |  | 10/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021019970 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced multiple cerebral infarcts. systemic heparin was withheld. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14276 | 1220648-2024-23523 | 20725804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725804&pc=OZD | 08/10/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/10/2021 | 11/19/2024 |  |  | 03/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1471128 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. impella 5.0 was attempted to be implanted into the left femoral artery (lfa). however, the device was unable to pass the superficial femoral artery due to stenosis. a vascular injury occurred, and surgical repair was required. it was decided to implant cp via right femoral artery (rfa), intra-aortic balloon pump (iabp) was removed and pump inserted however unable to regain contractile function. despite best efforts the case was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14277 | 1220648-2024-23537 | 20725844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725844&pc=OZD | 08/18/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/18/2022 | 11/19/2024 |  |  | 05/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023217223 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from thorax drains resulting in restricted anticoagulation treatment. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14278 | 1220648-2024-23525 | 20725856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20725856&pc=OZD | 08/21/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/21/2021 | 11/19/2024 |  |  | 06/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489284 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was observed to have heparin induced thrombocytopenia (hit). the impella purge solution was changed to sodium bicarbonate to manage the complication. it was reported that impella 5.0 was explanted, and a permanent ventricular assist device was placed. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14279 | 1220648-2024-23536 | 20726129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726129&pc=OZD | 08/13/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/13/2022 | 11/19/2024 |  |  | 02/28/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022053049 |  | 005066 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. physician reported impella 5.0 could not be passed through left axillary artery due to aorta resistance. physician was able to insert through left femoral artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14280 | 1220648-2024-23531 | 20726131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726131&pc=OZD | 08/13/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 08/13/2021 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161230 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. systemic heparin was paused for 4 hours due to bleeding and then restarted. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14281 | 1220648-2024-23545 | 20726146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726146&pc=OZD | 03/19/2022 | 05/12/2025 |  | 11/19/2024 |  |  | 03/19/2022 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022073279 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient developed bleeding and a hematoma at the access site. the patient received a blood transfusion to treat the reported problem. it was reported that the patient survived normal explant. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14282 | 1220648-2024-23540 | 20726177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726177&pc=OZD | 10/26/2024 | 01/06/2025 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562893 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support for a biventricular failure. while on support, oozing noted and two units of packed red blood cells were given. it was further reported that the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14283 | 1220648-2024-23548 | 20726433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726433&pc=PYX | 10/26/2024 | 12/23/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025533982 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Hispanic | White | Injury | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported delivery issues with an impella rp flex. rp flex backloaded over 0.018¿ wire but was unable to make the turn toward right ventricular outflow tract (rvot) despite utilization of best practices and multiple attempts. the patient was placed on extra corporeal membrane oxygenation and transferred to the intensive care unit without incidence. | [5 empty] | The investigation for the delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined be patient condition because unable to use sturdy wire to cross pv due to prolapse despite multiple attempts. unable to pass through vasculature and patient aborted. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14284 | 1220648-2024-23541 | 20726452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726452&pc=OZD | 10/16/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 10/16/2022 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022095609 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.0 had positioning issues. impella 5.0 inserted in the right axillary artery. implantation without complication. during support the patient had a red alert impella position incorrect, ecmo flow was not changed, position check recommended and depending on lv filling reduce impella flow slightly. volume management discussed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14285 | 1220648-2024-23542 | 20726494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726494&pc=OZD | 10/25/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 10/25/2022 | 11/19/2024 |  |  | 10/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021012424 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella device for mechanical circulatory support. during insertion, due to poor anatomy, the pump could not be placed from the right or left side. the artery was very deep, with steep angle of entry into vessel despite tunneling. no passage of the vessel beyond the clavicle possible, despite best practices. it was decided to switch from 5.0 to cp axillary. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14286 | 1220648-2024-23550 | 20726571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726571&pc=OZD | 03/13/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/13/2022 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022082500 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. the physician an alarm due to an increase in purge pressure. the physician said that the purge set was replaced. the position was adjusted; however, there was no change. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14287 | 1220648-2024-23547 | 20726638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726638&pc=OZD | 04/25/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/25/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121615 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during normal explant, the pump was removed patient went into ventricular tachycardia and had to be defibrillated one time. the patient went into ventricular tachycardia twice more and potassium was administered. the device was inserted back into the patient for support, it was discovered through echocardiogram that the impella position was too deep. the device was repositioned and there were no more reported runs of ventricular tachycardia. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14288 | 1220648-2024-23516 | 20726766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726766&pc=PYX | 10/29/2024 | 10/28/2025 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025543811 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella rp insertion was aborted after multiple attempts for placement. no issues with access. the wire would pull back out of the right papillary artery (rpa) back into the left ventricle (lv) due to the patient¿s large right ventricle (rv). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. e4 should have been left blank on manufacturer device report 1220648-2024-23516. g1 reporting contact email revised on manufacturer device report 1220648-2024-23516 in accordance with updated procedures. | The device was received for investigation. no change in the investigation results previously reported in manufacturer device report 1220648-2024-23516-1. d9 device returned to manufacturer date added.
14289 | 1220648-2024-23538 | 20726768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726768&pc=OZD | 10/17/2022 | 09/26/2025 |  | 11/19/2024 |  |  | 10/17/2022 | 11/19/2024 |  |  | 02/29/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023206441 |  | 005064-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 20 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported that the patient developed an axillary site hematoma. patient was taken to the or and the hematoma was evacuated. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
14290 | 1220648-2024-23552 | 20726840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726840&pc=OZD | 10/27/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/27/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561587 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and while on support, there were suction alarms. the patient received two units of packed red blood cells, 1.5 liters of albumin and one unit of fresh frozen plasma. the suction alarms resolved. additionally, there was concerns for blood flow. the patient's hand was cool and numb. arterial duplex scan showed no radial or ulnar flow and minimal axillary and brachial flow. the impella was explanted. the patient survived the explant. | [5 empty] | The investigation for the limb ischemia and the low pump flow has been completed. data was not returned for review. pump was returned in damaged condition (catheter was cut off & missing pump plug). visual inspection found no issues on the pump. based on clinical description and product analysis, the root cause of low flow was most likely patient condition as it was resolved by increasing the volume. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14291 | 1220648-2024-23512 | 20726893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726893&pc=OZD | 01/18/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/18/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121633 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Life Threatening | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, patient experienced atrial fibrillation (af) with rapid ventricular response (rvr). no additional information was provided related to the resolution. patient survived the procedure. | [5 empty] | The investigation for the reported atrial fibrillation has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the atrial fibrillation could not be determined. this report is being filed as part of a retrospective review of historical records.
14292 | 1220648-2024-23532 | 20726944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726944&pc=OZD | 10/26/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541917 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a non-st elevated myocardial infarction patient. the pump was placed but there was oozing from the left groin. the patient became increasingly agitated and was very restless/thrashing extremities and attempted to sit up in bed. precedex drip was restarted for the agitation and versed was administered once. due to the continued oozing, the heparin drip was discontinued. two units of red blood cells were transfused as hemoglobin had dropped. the patient was more oriented the next day and moved less in bed and the oozing decreased significantly and hemoglobin had increased. the patient outcome was stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient movement because the patient because agitated and resulted in bleeding. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14293 | 1220648-2024-23524 | 20726965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20726965&pc=OZD | 10/26/2024 | 12/24/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552287 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support an acute myocardial infarction / cardiogenic shock patient. the patient had the pump support ongoing when the pump position was checked due to low flows and they observed the pump not to be in appropriate position and there was clot adhered. the pump was removed after 78 hours. the team achieved hemostasis, and the patient survived. | [5 empty] | The investigation for the thrombus and positioning issues has been completed. the product was returned for investigation. a data log was not provided for this case run. the pump impeller was clean, without any biomaterial on it. no physical damage or abnormalities were observed on the returned product. a biomaterial was found on the inlet cage, which was removed during the investigation. the pump was working per specifications during the test in a bucket of water. as per the clinical, pump had appeared to be at aortic valve while having suction alarm and low pump flow. md adjusted the pump position and a clot was observed on the pump. pump was removed. the cause of the positioning issue was most likely use related, which could also cause the low pump flow. the root cause of the thrombus could not be determined without additional details. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
14294 | 1220648-2024-23533 | 20727022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727022&pc=OZD | 08/09/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/09/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022108621 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a stroke. prior to explant, patient was intubated again and placed back on respirator due to stroke. the patient survived and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14295 | 1220648-2024-23543 | 20727038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727038&pc=OZD | 10/27/2024 | 12/13/2024 |  | 11/19/2024 |  |  | 10/27/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515148 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the impella moved when assisting the patient to bed. the patient had a brief episode of ventricular tachycardia and the patient was shocked and intravenous amiodarone was restarted. the pump moved to the correct position with echocardiogram guidance. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email updated
14296 | 1220648-2024-23529 | 20727158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727158&pc=OZD | 08/19/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/19/2022 | 11/19/2024 |  |  | 03/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023193537 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced high infection parameter and unstable hemodynamics. actions taken are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14297 | 1220648-2024-23557 | 20727274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727274&pc=OZD | 02/26/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/26/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET |  |  | ABIOMED, INC |  |  | US |  | 2023168315 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2022 |  |  |  |  | The user facility reported a 55-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. during the course of support, the impella 5.5 pump was inadvertently pulled back into the aorta overnight and needed repositioning. the impella was removed, and interventional cardiology arrived at re-implant the device. upon reinsertion, the device became stuck under the papillary muscle. attempts to reposition the device failed to free the pump from under the papillary muscle. the staff was advised on the risk of hemolysis and that the flow rates were compromised as a result of the positioning. the doctor did not want to attempt any further and the decision was made to proceed with support with the improper positioning. the staff was unable to re-enable placement monitoring as severely dampened motor current was causing false alarms. the patient reportedly passed away the following day. it cannot be definitively denied that complications with the pump may have contributed to the patient's passing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14298 | 1220648-2024-23565 | 20727280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727280&pc=OZD | 05/12/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 05/12/2022 | 11/19/2024 |  |  | 06/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1493418 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. during impella placement, the patient's right leg, which had prior digit amputations, became numb and painful. it was noted that the extremity was assumed to be ischemic. no additional vascular access was able to be achieved as attempts to check the flow past the peel away sheath was unsuccessful. as a result of the limb ischemia the device was explanted, and it was reported the patent survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14299 | 1220648-2024-23586 | 20727487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727487&pc=OZD | 10/28/2024 | 12/30/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542822 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) has system self-check failed error. during investigation of the aic, it was determined that the event was preceded by an unexpected console shutdown/restart, during which pump (sn# (b)(6) had stopped, and could not be restarted on the same console. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the pump stop has been completed. the pump was not returned for investigation. the data log shows that an ac power disconnected alarm was seen after 568 hours of pump run. meanwhile, the pump stopped during use and was unable to restart console. a different console was used and was able to start the pump running without any issues. data logs were evident that there was no issue related to pump performance during the entire case run on both consoles. no problem was found with the pump as the reported issue was related to the console.
14300 | 1220648-2024-23554 | 20727495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727495&pc=OZD | 02/27/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/27/2022 | 11/19/2024 |  |  | 06/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021015000 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 74-year-old female patient was scheduled for impella 5.0 implant for mechanical circulatory support. during the insertion process, the pump was unable to advance beyond the anastomosis of the artificial vessel and subsequently kinked. the pump was removed and an impella cp was placed for patient support. there was no patient harm resulting from the delivery complication. | [5 empty] | The device was not returned for evaluation. upon conclusion of the investigation, it was determined that the reported delivery issue and kinked cannula were caused by the patient's anatomy. should additional relevant information become available, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
14301 | 1220648-2024-23544 | 20727508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727508&pc=OZD | 03/09/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/09/2022 | 11/19/2024 |  |  | 04/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022078617 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding from an unknown area. the staff administered a wash out and added bicarb to the purge. it was reported that the site was intact and the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14302 | 1220648-2024-23539 | 20727554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727554&pc=OZD | 10/26/2024 | 01/02/2025 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  | 07/31/2025 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376179 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was a problem with high purge pressure and low pump flow. tissue plasminogen activator/lysis was administered successfully. it was further reported that the high purge pressure alarm occurred before the pump stopped. the patient was taken to the operation room and pump was exchanged. the purge had not been changed to lysis. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of high purge pressure and pump stop has been completed. the pump was returned for investigation. the pump was cut from the 20 centimeter mark of the catheter. there was biomaterial found in the outflow cage and in the purge gap. data logs showed that at the time of the failure, there were suction alarms followed by motor current spike and purge pressure high and purge system blocked alarms. the purge pressure was seen rising for the past six hours. this was resolved by adding tissue plasminogen activator to the purge. therefore, the root cause of the high purge pressure issue was biomaterial. the pump investigation results found biomaterial in the purge gap and on the 10th day, data logs showed that there was a rise in purge pressure over the span of 18 hours with a corresponding rise in motor current up to the point of the high prurge pressure alarms and purge pressure blocked alarms. the flows were also saturated. therefore, the root cause of pump stop was biomaterial. d.9 device returned information/date was added. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-23539 was submitted in accordance with updated procedures. h.6 health effect - impact code 4644 was reported incorrectly on supplemental manufacturer device report 1220648-2024-23539-01. | Added-h6 impact code 4644.
14303 | 1220648-2024-23585 | 20727575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727575&pc=OZD | 10/28/2024 | 12/06/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1161396 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had various alarms related to the purge. the aic was replaced and issue resolved. | [5 empty] | H3) the investigation into the reported "aic - purge issue" has been completed since the original report was filed. product was returned for investigation. the console was received and tested with a known-good pump and purge, but issues were initially not reproduced. console inspection and troubleshooting revealed few anomalies, including: glucose residue under the motor shaft gasket (consistent with motor shaft being immobilized giving ¿piston blocked¿ errors reported in complaint), purge pressure sensor out-of-spec (unrelated to complaint) and pbm failure to charge one of batteries (unrelated to complaint). during data analysis, the console logs from the reported day of event are partially consistent with complaint as the console had not been used on the reported day of event (2024-10-28) however logs entries between 2024-10-10 and 2024-10-12 show multiple purge-related alarms during support with pump 525492: purge flow decreased¿, purge pressure high, purge system blocked. pump was transferred to a backup automated impella controller (aic), however its serial number had not been identified therefore we could not confirm that all purge issues were resolved. the most probable root cause for the reported event of "aic - purge issue " was contamination.
14304 | 1220648-2024-23569 | 20727584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727584&pc=OZD | 11/18/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 11/18/2021 | 11/19/2024 |  |  | 08/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022126345 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.0 device for mechanical circulatory support. while on support the pump was reported to have stopped after 22 days of support. the physician tried to restart with p2-p8 several times, but it did not restart. in addition, there was a blood clot attached to the catheter after removal. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The investigation into the pump stop/ thrombus issue has been completed. the impella 5.0 was returned for investigation. the data analysis of the logs showed an initial pump stop occurred on the 22nd day of support. the pump was able to be restarted. upon restart, the motor current (mc) slowly ramps down to normal levels, indicating a possible mechanical obstruction of the motor. the pump was run for 4 more days, when a final high mc pump stop occurs. the pump was returned with a cut in the catheter by the tuohy. the catheter and mc lines were severed, but the purge and nitinol cables were not, and thus the pump could not be tested. the pump was purged and a leak occurred near the cut in the catheter, and no blood ingress was found. the red handle was opened, and no issues were found. the pump was torn down, and no clear evidence of bearing wear was found on the inner bearing or retaining ring. the root cause of the pump stop is most likely due to physical obstruction/thrombus of the impeller or motor, but this could not be confirmed as the pump could not be tested. this report is being filed as part of a retrospective review of historical records.
14305 | 1220648-2024-23535 | 20727686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727686&pc=OZD | 08/13/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 08/13/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161221 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hemolysis was noted on patient labs. it was noted that the complication was successfully managed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14306 | 1220648-2024-23564 | 20727730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727730&pc=OZD | 10/23/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 10/23/2021 | 11/19/2024 |  |  | 05/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022083095 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient implanted with an impella 5.0 device for mechanical circulatory support. while on support there were positive blood cultures, indicating infection. antibiotics were administered. the patient was noted as being afebrile. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14307 | 1220648-2024-23549 | 20727732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727732&pc=OZD | 09/12/2024 | 11/19/2024 |  | 11/19/2024 |  |  | 09/12/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced ventricular fibrillation (vf) arrest overnight with return of spontaneous circulation [rosc] after five minutes. impella was subsequently successfully weaned and explanted. patient survived the procedure. | [5 empty] | The investigation for the reported ventricular fibrillation (vf) has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the vf could not be determined.
14308 | 1220648-2024-23555 | 20727733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727733&pc=OZD | 02/28/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/28/2022 | 11/19/2024 |  |  | 01/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 1491508 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/28/2022 |  |  |  |  | The user facility reported that a patient was implanted with an impella 2.5 pump for mechanical circulatory support. following insertion, it was documented that the pump suffered a spike in motor current and stopped unexpectedly. attempts to restart the pump were not successful. the pump was subsequently replaced with an impella cp device for ongoing patient support. there was no report of patient harm. | [5 empty] | Neither the device nor the log data were returned for evaluation. upon conclusion of the investigation, the reported issue could not be verified, and a cause could not be established. should additional relevant information become available, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
14309 | 1220648-2024-23584 | 20727741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727741&pc=OZD | 10/28/2024 | 01/06/2025 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP P |  |  | ABIOMED, INC. |  |  | US |  | 2025521185 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the high risk coronary intervention (hrpci) procedure the impella cp had a purge high pressure alarm. the cause was unknown. there were no kinks in the purge line. after the percutaneous coronary intervention (pci) the impella was explanted. | [5 empty] | The investigation for the high purge pressure has been completed. the pump was returned for investigation. the returned pump did not have any physical damage. small amount of biomaterial was noted in the purge gap. the pump was primed and ran as expected during testing. high purge pressure did not reproduced on the returned product. as per the given clinical details, a high purge pressure alarm was noted during use. the total case run time was 1.25 hours. the data log showed a spike in motor current before the rise in purge pressure, which resulted in high purge pressure. the pump was explanted. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and returned product investigation. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-23584 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23584 was submitted in accordance with updated procedures. g1 revised manufacturing site fax number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23584. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
14310 | 1220648-2024-23592 | 20727759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727759&pc=OZD | 10/25/2024 | 01/02/2025 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537959 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support being initiated for an acute myocardial infarction/cardiogenic shock patient. the pump failed to appropriately purge so the pump was removed and replaced. no harm or injury was noted. the patient survived the explant. | [5 empty] | The investigation of priming issue has been completed. the pump was returned for investigation. the pump and purge cassette were successfully primed for 15 minutes without any issues. no other issues were observed. there was no issue found during the case. the device functioned/operated to specification. g1 reporting contact email revised on manufacturer device report 1220648-2024-23592 in accordance with updated procedures. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14311 | 1220648-2024-23560 | 20727832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727832&pc=OZD | 10/19/2021 | 05/12/2025 |  | 11/19/2024 |  |  | 10/19/2021 | 11/19/2024 |  |  | 02/28/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022045658 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support the patient experienced bleeding. during the initial insertion the physician shortened the graft too much, due to this he could not fix the wings on the skin. the physician fixed the pump up to the theory borst under the skin. there was bleeding into the sterile sleeve because of the butterfly placed under the skin. no further information was provided as to the treatment or outcome for the reported event and patient. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14312 | 1220648-2024-23567 | 20727879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727879&pc=OZD | 11/10/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 11/10/2022 | 11/19/2024 |  |  | 10/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022028311 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that upon explant of the impella device, a clot was noted on the impella pigtail. no further issues were noted during the weaning and explant of the device. it was noted that the closure was a success. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14313 | 1220648-2024-23553 | 20727906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727906&pc=OZD | 02/25/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/25/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022116052 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient experienced multiple bleeding complications during impella 5.0 support. it was initially documented that the patient developed a hematoma at the intra-aortic balloon pump removal site roughly four days into impella support. it was noted that heparin was in the purge line at the time of the bleed. the patient was given 1 unit of packed red blood cells, heparin levels were reduced, and a sandbag was placed at the site to help treat the condition. oozing at the impella site was also observed at this time. about three days later, the patient experienced a gastrointestinal (gi) bleed. once again, heparin was in the purge line at the time of the bleed. the gi bleed was controlled and support continued. it is unknown of heparin in the purge contributed to the bleeds. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14314 | 1220648-2024-23570 | 20727925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727925&pc=PYX | 10/28/2024 | 12/31/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  | 10/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023283573 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp placement was aborted due to the patient's vascular anatomy. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through the vasculature. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
14315 | 1220648-2024-23571 | 20727936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20727936&pc=OZD | 09/02/2024 | 07/07/2025 |  | 11/19/2024 |  |  | 09/02/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524872 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown race/ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella began to alarm and, and a transthoracic echocardiogram (tte) showed that the pump had pulled into the aorta. the impella was removed and the patient remained on extracorporeal membrane oxygenation. no patient's harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The cause of the pump pulling out into the aorta was not determined due to insufficient clinical details.
14316 | 1220648-2024-23556 | 20728100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728100&pc=OZD | 10/26/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552636 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. following direct graft placement, the patient began bleeding excessively from multiple areas and hindered successful placement. impella purge primed and the surgeon elected to withdraw care and then the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. athe root cause of access site bleeding is determined to be patient condition because of the bleeding from multiple locations. the root cause of delivery issue is determined to be patient condition because the pump was unable to be delivered due to patient anatomy and resulted in excess bleeding. corrections to the initial mfr report:g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14317 | 1220648-2024-23572 | 20728115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728115&pc=OZD | 08/18/2024 | 03/04/2025 |  | 11/19/2024 |  |  | 09/02/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517775 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient's labs were found to be hemolyzed, though no intervention was specified for this issue. it was later found that the patient had also experienced a hematoma at the right upper chest insertion site after the patient tried to pull themselves up in their bed. the hematoma was addressed with pressure and additional dressing, and anticoagulants were temporarily discontinued. the patient recovered with no reported sequelae. | [5 empty] | The investigation for the access site bleed and the hemolysis has been completed. data logs show no motor current spikes and no positioning issues mentioned in clinical. there are suction alarms at the start which resolve. the cause of hemolysis was not determined since product was not returned and there is insufficient clinical information. bleeding reported after patient pulled herself up in bed. the cause of the access site bleeding - major was most likely patient movement. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ corrections to the initial mfr report: g1 mdr reporting contact email | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14318 | 1220648-2024-23590 | 20728129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728129&pc=OZD | 10/28/2024 | 01/02/2025 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508324 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp stopped seven hours after support was initiated. anticoagulation was very difficult to manage. the impella cp was explanted and a new impella cp was placed. the patient survived the explant. | [5 empty] | The investigation for the pump stop has been completed. data logs were reviewed, and the pump stop was confirmed. that day, there is a rise in purge pressure from 500 to roughly 950 mmhg over the course of an hour, but there is no trigger of any purge pressure alarms. alarm 13 triggers during this rise, and the pump isn't able to start back up successfully for the rest of the case. there are no distinct motor current spikes leading up to this trend up of the purge pressure. the only abnormality leading up to the event is earlier in the day where motor current, motor speed, and placement signal all have a clear increase in noise/amplitude without a change in p-level. this could be due to biomaterial or environmental factors. clinical details provided that anticoagulation was very difficult for this case, however, product was not returned for investigation. without returned product and clear signs of biomaterial ingestion versus a complication with the purge lumen, the root cause of the pump stop could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14319 | 1220648-2024-23587 | 20728168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728168&pc=OZD | 10/28/2024 | 12/24/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552306 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, there was a loss of doppler pulses to the distal limbs. the patient was at cardiogenic shock stage e and had endured multisystem organ failure. the patient was made do not resuscitate and eventually expired. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14320 | 1220648-2024-23579 | 20728202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728202&pc=OZD | 10/29/2024 | 12/19/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support ongoing for a high risk coronary intervention (hrpci) 66-year-old male patient. the pump was placed but after 7 hours of support and with patient transferring the pump came out of position. the pump was unable to be re-delivered to the left ventricle (lv) and the pump was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was patient movement related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23579. e4 should have been left blank g1 reporting contact email revised in accordance with updated procedures.
14321 | 1220648-2024-23558 | 20728210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728210&pc=OZD | 02/24/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/24/2022 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1090627 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Failure to Power Up | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/24/2022 |  |  |  |  | The complainant reported that an automated impella controller display would not power on and the console was emitting a continuous beep. the power was cycled but the display remained off. there was no patient involvement at the time of the noted issue. | [5 empty] | The console was returned, and an evaluation was completed to investigate the reported allegation. upon testing the device, it was verified that the display and keyboard would not light up or power on when the aic power button was pushed. the console was opened up and a voltage measurement from the carrier board to the epc was found to be below specifications. this resulted in an insufficient voltage supply to the display and keyboard assembly. the 4in1 assembly was replaced with a known working assembly and the issues resolved. upon conclusion of the investigation, the failure was confirmed, and the cause was traced to a faulty carrier board. should additional relevant information become available, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
14322 | 1220648-2024-23546 | 20728294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728294&pc=OZD | 03/17/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/17/2022 | 11/19/2024 |  |  | 01/31/2024 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023174264 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding in the lung and required a transfusion of red blood cells. heparin was withheld. it was reported that the patient survived normal explant of the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14323 | 1220648-2024-23581 | 20728451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728451&pc=OZD | 10/28/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533224 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella 5.5. impella backloaded onto the wire and was unable to pass the clavicle/sternal junction despite repeated attempts. the wire was exchanged for a stiff angled glide wire and was still unable to advance pump. the decision was made to abort impella 5.5 placement and place an impella cp in axillary graft due to anatomy. the patient¿s outcome is unknown. | [5 empty] | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23581 in accordance with updated procedures. | The impella device was not received from the customer and therefore, ¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14324 | 1220648-2024-23559 | 20728554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728554&pc=OZD | 02/23/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/23/2022 | 11/19/2024 |  |  | 06/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC |  |  | US |  | 2021008993 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/23/2022 |  |  |  |  | The user facility reported a 77-year-old male patient was scheduled for impella 5.0 implant for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance through the bifurcation at the aorta and subclavian artery. after attempts to manipulate positioning failed, the pump was removed, and a kink occurred near the motor. an impella cp pump was subsequently placed for patient support. there was no patient harm resulting from the delivery issue. | [5 empty] | The impella device was not returned for evaluation. an investigation into the noted allegation was completed. based on the provided information, it was determined that the delivery issue was caused by the patient's anatomy/condition. should additional relevant information become available, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14325 | 1220648-2024-23591 | 20728768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728768&pc=OZD | 04/05/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/05/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023160027 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited a decrease of purge flow and an increase in of purge pressure. the pump was explanted and replaced. no report of harm or injury to the patient. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, biomaterial was observed on the outflow cage and kinks in the catheter were observed. therefore, the root cause of the high purge pressure was due to a kinked purge line. no other damage was found to the pump. this report is being filed as part of a retrospective review of historical records.
14326 | 1220648-2024-23561 | 20728853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728853&pc=OZD | 10/27/2024 | 12/16/2024 |  | 11/19/2024 |  |  | 10/27/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548303 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a patient admitted in for st elevated myocardial infarction. the pump was placed after the diagnosis of multi-vessel disease, coronary artery disease, and ventricular tachycardia with hypotension. the levophed pressor was increased and the pump-accessed limb became ischemic. the ischemia was seen by loss of pulses and the staff placed a femorofemoral bypass. the pump remained on until the patient expired after 46 hours of support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14327 | 1220648-2024-23562 | 20728897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728897&pc=OZD | 10/26/2024 | 12/16/2024 |  | 11/19/2024 |  |  | 10/26/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519825 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to st elevated myocardial infarction. while on support, there was a lack of pedal pulses to leg. the decision was made to remove impella cp and replace with intra-aortic balloon pump. the procedural outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into limb ischemia ¿ reversible has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23562.
14328 | 1220648-2024-23573 | 20728926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728926&pc=OZD | 10/28/2024 | 12/16/2024 |  | 11/19/2024 |  |  | 10/28/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550455 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump supporting a patient admitted in after arresting in the field. the patient was suffering from an st elevated myocardial infarction, pulmonary edema, coronary artery disease, and was on continuous renal replacement therapy. the team observed bilateral limb ischemia and the team placed an antegrade sheath. the family withdrew care and the patient expired. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report (b)(4). | B.2 the exact date of death is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14329 | 1220648-2024-23583 | 20728975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728975&pc=OZD | 10/29/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 69-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella cp was exchanged due to a pump failure. the patient had a significant drop in blood pressure with a need for chemical resuscitation. after the pump was exchanged, the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The manufacturer device report number for impella cp with serial number (b)(6) was submitted in error as there were no reportable allegations for this pump. the reported complaint of "pump stop" along with thrombus has been captured and reported under manufacturer device report number 1220648-2024-23515 for an impella cp with serial number (b)(6).
14330 | 1220648-2024-23595 | 20728991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20728991&pc=OZD | 10/25/2024 | 12/13/2024 |  | 11/19/2024 |  |  | 10/25/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | [11 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a first impella cp was inserted but the cp was having a sensor problem that was reported in complaint (b)(4). the medical doctor exchanged to a new impella cp. the patient continued to code and expired in the procedural area. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23595. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14331 | 1220648-2024-23566 | 20729075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729075&pc=OZD | 10/27/2024 | 12/12/2024 |  | 11/19/2024 |  |  | 10/27/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546039 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted into a tertiary center from the community hospital with a st elevated myocardial infarction. the pump was placed for angioplasty and was supporting for four hours. the pump was then removed due to a hematoma at the access site. the explant and manual hold helped the hematoma/bruising to resolve. no other intervention was reported. the patient¿s outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely patient movement since hematoma was discovered after the patient was transferred. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23566.
14332 | 1220648-2024-23575 | 20729144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729144&pc=OZD | 04/04/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/04/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023187923 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. the complainant reported that the pump would not start, the device was replaced with a new impella. no report of patient harm or injury. | [5 empty] | The device was returned for evaluation. data logs were not sent. significant kinks were found throughout the length of the cannula and the catheter. no other damage or abnormalities were found. the pump started and ran with no issues. therefore, the cause of the pump stop was most likely use related. this report is being filed as part of a retrospective review of historical records.
14333 | 1220648-2024-23576 | 20729180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729180&pc=OZD | 10/29/2024 | 12/24/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520065 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown age and gender in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, distal pulses were lost in distal extremity. impella was removed due to limb ischemia. | [5 empty] | The investigation limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23576. g1 reporting contact email revised in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14334 | 1220648-2024-23568 | 20729217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729217&pc=OZD | 11/11/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 11/11/2022 | 11/19/2024 |  |  | 11/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232172 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Ventricular Fibrillation | [10 empty] | Other | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that during impella placement, the patient experienced episodes of ventricular tachycardia and ventricular fibrillation. it was stated that the patient was shocked, and rhythm was restored. the patient remained on impella support for the intervention, and was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14335 | 1220648-2024-23577 | 20729393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729393&pc=OZD | 10/29/2024 | 12/30/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had access site bleeding. four units of blood products were given to the patient, pressure was held and pump was explanted. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-23577. g1 reporting contact email revised on manufacturer device report 1220648-2024-23577 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14336 | 1220648-2024-23596 | 20729459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729459&pc=OZD | 04/02/2022 | 04/16/2025 |  | 11/19/2024 |  |  | 04/02/2022 | 11/19/2024 |  |  | 06/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 1481548 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 2.5 for mechanical circulatory support. during support, it was discovered that the pump was misplaced in the ascending aorta, it was impossible to reimplant, and the physician decided to explant the device and support the patient with inotropes. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Correction information for the initial mfr 1220648-2024-23596 was provided in sections d4 for udi. | Correction to follow up report 1 of medical device report 1220648-2024-23596. b1 has been corrected to select adverse event and product problem. h1 has been corrected to reportability type death. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14337 | 1220648-2024-23593 | 20729568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729568&pc=OZD | 12/01/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/04/2022 | 11/19/2024 |  |  | 12/31/2020 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-JP |  |  | OZD |  |  | IMPELLA 2.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 1391206 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Heart Problem | [12 empty] | Other | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella for mechanical circulatory support. six months after the impella device was explanted, the patient was observed by transthoracic echocardiogram and the staff discovered an aortic valve injury that they attributed to use of the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the heart valve damage issue has been completed. the cause of the aortic valve injury and its relation to the impella was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
14338 | 1220648-2024-23551 | 20729682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729682&pc=OZD | 10/25/2024 | 12/24/2024 |  | 11/19/2024 |  |  | 10/25/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557212 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was taken to the cath lab and was hypotensive and it was decided to intubate patient prior to impella cp procedure. prior to insertion the staff had to unplug the impella to recalibrate the sensor. the medical doctor wanted a new impella in case sensor was compromised. additionally, the patient coded and required cardiopulmonary resuscitation upon insertion of the impella. the cp was exchanged for a new one. the patient continued on support which will be captured in complaint 24-39115-1 and another manufacturer device report will be submitted. | [5 empty] | The investigation placement signal issue and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the root cause of the optical signal issue was most likely due to an air gap. as the necessary information was not provided, the root cause of the vt/vf was not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23551 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14339 | 1220648-2024-23580 | 20729685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729685&pc=OZD | 10/29/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 47-year-old female patient with caridomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was noted at the access site. the hematoma was managed by managing activated clotting time (act), giving packed red blood cells and applying pressure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14340 | 1220648-2024-23594 | 20729781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729781&pc=OZD | 04/03/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/03/2022 | 11/19/2024 |  |  | 12/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2023161625 |  | 005082-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the device exhibited a placement signal and then small drops of blood leaking from the impella red plug was observed. the staff tapped gently on red impella plug and bleeding stopped, along with improvement of placement signal, and flows. additionally, the purge cassette was also leaking. it was reported that the motor current was pulsatile and the patient stable. the patient was normally weaned and survived the procedure. | [5 empty] | The impella device was returned for evaluation. data logs showed that the placement signal suddenly became erratic, accompanied by alarm 120. at this time, the motor current had low pulsatility and the pump flow was clipped. soon after, the "purge volume low" alarm occurred. after the alarm resolved, all pump metrics began to stabilize. the pump was returned with the catheter cut off near the motor housing. no holes were observed in the segment of the catheter was that was returned. there was a significant amount of blood found in the red pump plug. however, the cause of the blood ingress into the pump handle was not determined since the pump was not returned intact. the cause of the placement signal issue was most likely an electrical short due to the blood ingress into the pump plug, attributed to the patient condition, however the cause is unable to be determined. the purge cassette was cut open and leak reproduction testing was performed; a cartridge cap leak was observed. this was attributed to the manufacturing process. therefore, the cause of the purge cassette leak was a cartridge cap bond failure. this report is being filed as part of a retrospective review of historical records.
14341 | 1220648-2024-23578 | 20729797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729797&pc=OZD | 10/22/2024 | 12/18/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025566135 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Sepsis | [12 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 48-year-old male patient was on impella 5.5 due to non-ischemic cardiomyopathy. while on support, it was noted that purge pressure was fluctuating. tissue plasminogen activator (tpa) was infused in purge overnight for elevated pressures at 1000. additionally, the patient developed sepsis/bacteremia. the decision was made to explant the impella due to sepsis. | [5 empty] | The investigation into infection/sepsis and high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure was most likely biomaterial. because this evidence has not been provided, the root cause of the infection cannot be determined the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23578. b1 should have been both adverse event and product problem b3 date of event e4 should have ben left blank g1 reporting email address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14342 | 1220648-2024-23588 | 20729810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729810&pc=OZD | 03/22/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/22/2022 | 11/19/2024 |  |  | 07/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129453 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during support, there was bleeding during suturing on graft. surgical techniques were used and heparin withheld to manage the reported observation. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14343 | 1220648-2024-23589 | 20729824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20729824&pc=OZD | 04/06/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 04/06/2022 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125045 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventative maintenance the automated impella controller's (aic) display would turn black after the bias was loaded. the user replaced the 4 and 1 board, this resolved the issue. no patient involvement. | [5 empty] | The console (aic) was not returned for evaluation; however this issue was resolved at the customer site. the root cause of the black screen was due to a malfunctioning 4-in-1 board. this report is being filed as part of a retrospective review of historical records.
14344 | 1220648-2024-23582 | 20730005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730005&pc=OZD | 10/29/2024 | 12/26/2024 |  | 11/19/2024 |  |  | 10/29/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536896 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. an attempt was made to insert an impella cp from the femoral artery. however, the bend in the replaced graft remained, and it got stuck there, preventing the impella from advancing, and cp insertion was abandoned. | [5 empty] | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the bend in the replaced graft remained, and it got stuck there, preventing the impella from moving forward. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings and cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23582 g1 reporting contact email revised in accordance with updated procedures. h10 instructions for use missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
14345 | 1220648-2024-23602 | 20730115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730115&pc=OZD | 02/08/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/08/2022 | 11/19/2024 |  |  | 03/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022085138 |  | 005082-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. it was reported that there was successful coronary artery bypass grafting x 2 with the impella ld implant; however, heparin was withheld due to bleeding. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14346 | 1220648-2024-23574 | 20730130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730130&pc=OZD | 10/21/2024 | 12/13/2024 |  | 11/19/2024 |  |  | 10/21/2024 | 11/19/2024 |  |  |  |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551213 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient with cardiogenic shock from an acute myocardial infarction was implanted an impella cp device for mechanical circulatory support and developed bleeding at the insertion site. post implant same day the patient had moderate amount of bloody drainage at insertion site that was successfully management with manual pressure and pressure dressing with sandbag. the next day a hematoma was noted. hemostasis was obtained with a mechanical compression device proximal to insertion site, and one unit of blood was administered. vascular surgery was consulted and was able to control the site. follow up the next day noted the patient was stable; hematoma was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-23574.
14347 | 1220648-2024-23563 | 20730242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730242&pc=OZD | 04/05/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 04/05/2021 | 11/19/2024 |  |  | 11/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043202 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported after the impella 2.5 was successfully weaned and explanted, there was difficulty tightening the perclose device to close the impella access site. the access site was oozing after perclose deployment, aso angioseal was also deployed to achieve hemostasis. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was most likely use related, as after explant there was difficulty tightening percloses as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
14348 | 1220648-2024-23597 | 20730254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730254&pc=OZD | 07/13/2020 | 10/07/2025 |  | 11/19/2024 |  |  | 03/30/2022 | 11/19/2024 |  |  | 10/31/2021 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1450376 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma. this was treated with a blood transfusion. additionally, it was reported that the patient experienced sudden cardiac death, this occurred 12 days after the device was explanted, there is no allegation against the device. however, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-23597. g1 reporting contact email revised on manufacturer device report 1220648-2024-23597 in accordance with updated procedures. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-23597. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-23597.
14349 | 1220648-2024-23598 | 20730256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730256&pc=OZD | 03/27/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/27/2022 | 11/19/2024 |  |  | 12/31/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023166449 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, a thrombus was identified, the resolution for this issue is unknown. additionally, the sidearm of the device was leaking. the resolution for these issues is unknown. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14350 | 1220648-2024-23605 | 20730360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730360&pc=OZD | 02/03/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/03/2022 | 11/19/2024 |  |  | 09/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013252 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma and was bleeding at right axillary incision site. issue was addressed with packed red blood cells and manual compression. it was also noted that staff was aware of thrombocytopenia. action taken is unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14351 | 1220648-2024-23599 | 20730429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730429&pc=OZD | 03/23/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 03/23/2022 | 11/19/2024 |  |  | 06/30/2023 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022101658 |  | 005066 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the device exhibited high purge pressure. it was reported that the purge flow rate reached 2.0 milliliters per hour and the physician implanted an emergency left ventricular assist device and explanted the 5.0. it was reported that the patient survived explant. | [5 empty] | The device was returned for evaluation. data logs show a slow rise in purge pressure of the course of 4 hours, eventually leading to a high purge pressure alarm. product evaluation showed evidence of biomaterial in the purge gap. the pump ran to specification in the lab, with purge pressures around 700 mmhg. based on the log analysis and the evidence of biomaterial in the purge gap, the root cause of the high purge pressure was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
14352 | 1220648-2024-23601 | 20730515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730515&pc=OZD | 07/15/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 07/15/2021 | 11/19/2024 |  |  | 03/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1475275 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support it was reported that the patient had an episode of atrial fibrillation while walking. levo and epi hung briefly. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14353 | 1220648-2024-23600 | 20730563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730563&pc=OZD | 06/25/2021 | 11/19/2024 |  | 11/19/2024 |  |  | 06/25/2021 | 11/19/2024 |  |  | 09/30/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1504064 |  | 005082 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. while on support the purged was switched to d5 due to bleeding concerns. it was also noted that the patient was having symptoms of heparin-induced thrombocytopenia. heparin was removed from purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14354 | 1220648-2024-23603 | 20730585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730585&pc=OZD | 01/30/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 01/30/2022 | 11/19/2024 |  |  | 01/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1465640 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site oozing. doctors applied new dressing to the access site, and managed to bring anticoagulation down. several days later, oozing at the access site was reported by a nurse. he explained that doctors redressed the access site to achieve hemostasis. several days later, patient went to the cardiac cath lab for ventricular septal defect repair. during the procedure, a patch was placed. while the patch was released, the pressure from the right ventricle broke the tissue causing the patch to be free floating in the left ventricle with embolization. the patient was rushed to or where they were bleeding profusely, then sent to the unit as they were too unstable to continue procedure. impella was unable to be re-positioned away from the mitral valve due to clots in the graft. patient¿s urine had dropped to less than 30 cc¿s over the next two hours. staff also noted that patient had dialysis line being placed for continuous renal replacement therapy. patient was having pink colored urine that turned merlot over night. patient¿s urine had dropped to less than 30 cc¿s within the last two hours. also, patient had surgical complications due vsd repair. patient went to ccl where md tried to place a patch. patch had thromboses and a clot was dislodged into the lv. patient was rushed to the or. graft had clots in it, so pump was unable to be re-positioned away from the mitral valve. blood products given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14355 | 1220648-2024-23604 | 20730587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20730587&pc=OZD | 02/02/2022 | 11/19/2024 |  | 11/19/2024 |  |  | 02/02/2022 | 11/19/2024 |  |  | 08/31/2022 |  | 11/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504554 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support the patient had complications from intra-aortic balloon pump. patient went back to surgery to repair femoral artery. patient was noted as improving post repair and drops were weaned; bleeding stopped and heparin maintained in purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14356 | 1220648-2024-23320 | 20705350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705350&pc=OZD | 06/04/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 06/04/2022 | 11/18/2024 |  |  | 03/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2023199522 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Pulmonary Edema | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported the patient was escalated to the impella 5.5 with a right ventricular assist device from the impella cp and extra-corporeal membrane oxygenation. reportedly the impella came out of position and the patient had a flash pulmonary edema. the patient's care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14357 | 1220648-2024-23322 | 20705646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705646&pc=OZD | 07/16/2024 | 06/27/2025 |  | 11/18/2024 |  |  | 10/21/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488356 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "definite" relationship to the impella cp device used: ¿subject developed hematuria. interventionalist to bedside to attempt repositioning cp impella but unsuccessful. impella removed at bedside at 1530. heparin stopped at this time per md. subject taken back to cath lab at 1600 to remove impella sheath and close site. labs for hemolysis ordered. hemolysis ld total was 2,070 on (b)(6) 2024 at 17:55). hgb 14.7 on (b)(6) 2024 at 13:42). echo completed the next morning on (b)(6) 2024: ef 35-40%, when compared to previous echo on (b)(6) 2024, the lv function appears improved. clear amber urine noted on (b)(6) 2024 at 1826. hemolysis ld was not repeated.¿ the patient recovered with no sequelae. additionally, the patient endured hypotension with a "definite" relationship to the impella device used: after impella was removed, the patient developed hypotension requiring pressors. blood pressures in low 80¿s/50¿s, map 57. md notified. subject started on dopamine for a goal of map >65. subject developed shortness of breath and chest pain after dopamine was started. md notified: stopped dopamine infusion and obtained ekg. symptoms resolved shortly after dopamine was turned off. started on levophed to keep map >65. provider reviewed ekg: no changes noted. maintained map >65 while on levophed. was transiently on levophed following a staged impella supported pci procedure on (b)(6) 2024 (levophed had been discontinued at 1220 after procedure completed but briefly restarted from 1646-1745 due to maps 58-62). levophed discontinued on (b)(6) 2024 at 1745. denies chest pain or sob. bp¿s 92-114/57-81 & maps remain >65.¿. the patient recovered with no sequelae. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | E4 should have been left blank on the initial manufacturer device report number 1220648-2024-23322 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-23322 in accordance with updated procedures. g2 report source; study was missing from manufacturer device report 1220648-2024-23322. | The investigation of hemolysis and vt/vf/at/af has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in improper position corresponded with the time hemolysis was reported per clinical description that triggered alarms impella position in aorta. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the cause of the hypotension was unable to be determined due to insufficient clinical details.
14358 | 1220648-2024-23323 | 20705663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705663&pc=OZD | 05/28/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 05/28/2022 | 11/18/2024 |  |  | 02/29/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196105 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intracranial Hemorrhage | Intraoperative Cardiac Injury | [11 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella for mechanical circulatory support. during support, it was reported that the patient left ventricle was perforated and the physician performed a sternotomy with left ventricle repair. additionally, the patient had a cerebral hemorrhage, all anticoagulation was stopped and heparin removed from purge fluid. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14359 | 1220648-2024-23321 | 20705768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705768&pc=OZD | 06/23/2022 | 10/02/2025 |  | 11/18/2024 |  |  | 06/23/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022154247 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | Asian | Death | Yes | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Respiratory Failure | Thrombocytopenia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The procedural outcome noted that the patient expired. cause of death is respiratory failure. the user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma. to help resolve this issue the patient was administered a blood transfusion. additionally, the patient experienced thrombocytopenia. the resolution for this issue is unknown. however, it was noted that the patient recovered. the device exhibited a purge leak and the pumped stop. the physician decided to explant the device. the procedural outcome noted that the patient expired. cause of death is respiratory failure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the pump stop and purge leak issues has been completed. the impella device was not received from the customer. therefore, the root cause for the could not be determined. | This report is being filed as part of a retrospective review of historical records. access site bleeding-major and thrombocytopenia: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23321 in accordance with updated procedures.
14360 | 1220648-2024-23325 | 20705976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705976&pc=OZD | 01/13/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/13/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022125695 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the doctor attempted to implant the patient with impella 5.0 device for mechanical circulatory support. the doctor made an attempt from the right and left side, but implant was not possible due to patient¿s anatomy (anomaly at the subclavia). no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14361 | 1220648-2024-23324 | 20705988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20705988&pc=OZD | 05/29/2022 | 11/30/2024 |  | 11/18/2024 |  |  | 05/29/2022 | 11/18/2024 |  |  | 03/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2023193511 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella placement, weaning from venoarterial extracorporeal membrane oxygenation (va ecmo) could not be performed due to spontaneous bleeding from the patient¿s lung and nose. the device was explanted, and the procedural outcome was the patient expired on support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | Additional information noted that the procedure outcome was that patient expired due to cardiovascular insufficiency. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Correction: b2, b5, and h1 was updated with correct information which was reported incorrectly in the initial report.
14362 | 1220648-2024-23327 | 20706105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706105&pc=OZD | 05/23/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/23/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129269 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device was repositioned and then exhibited suction issues indicating low flow. it was noted that the device was in a good position. the resolution for this issue is unknown. it was reported that the patient survived normal explant. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device nor the data logs were not received from the customer. therefore, the root cause of low pump flow cannot be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
14363 | 1220648-2024-23328 | 20706159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706159&pc=OZD | 11/11/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 11/11/2021 | 11/18/2024 |  |  | 02/28/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060424 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. the patient was successfully weaned from the impella and the device was removed, however, there was access site bleeding upon removal. manual pressure was held and a small hematoma was present at the site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14364 | 1220648-2024-23326 | 20706238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706238&pc=OZD | 05/25/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/25/2022 | 11/18/2024 |  |  | 12/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-JP |  |  | OZD |  |  | IMPELLA 5.0-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023162648 |  | 005066-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Heart Failure/Congestive Heart Failure | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella for mechanical circulatory support. during support, the patient experienced limb ischemia, the resolution for this issue is unknown. the procedural outcome was death, the cause of death was heart failure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14365 | 1220648-2024-23329 | 20706295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706295&pc=OZD | 05/20/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/20/2022 | 11/18/2024 |  |  | 06/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022102900 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. the impella was inserted through the guide wire, however it was not possible to move forward near the anastomotic site of the graft, and it became impossible to insert it more, so the procedure was aborted. no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. therefore, the root cause of the delivery issues cannot be determined.
14366 | 1220648-2024-23331 | 20706370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706370&pc=OZD | 05/17/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 05/17/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121665 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding from an unknown location and heparin was withheld. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14367 | 1220648-2024-23330 | 20706445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706445&pc=OZD | 05/17/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 05/17/2022 | 11/18/2024 |  |  | 04/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075943 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, it was reported that the patient was oozing from multiple sites. the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14368 | 1220648-2024-23333 | 20706472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706472&pc=OZD | 05/12/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/12/2022 | 11/18/2024 |  |  | 08/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503737 |  | 005062-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that during delivery the physician used an angled pigtail to get access to the left ventricle with some difficulty accessing across aortic valve. the device would not make a right angle into ascending aorta. the device was replaced with another impella 5.0, this device was implanted successfully. it was noted that the anatomy of the brachiocephalic/aorta angle was very torturous. no further information is available. | [5 empty] | The investigation into the delivery issue has been completed. the impella 5.0 was returned for investigation. no visual damage was found during inspection of the returned pump. the patient's condition most likely contributed to the inability to insert the pump. the cause of the issue was patient condition related since the ascending aorta was difficult to navigate
14369 | 1220648-2024-23334 | 20706524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706524&pc=OZD | 05/09/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 05/09/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022146952 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. it was reported that the patient experienced a gastrointestinal bleed any medical intervention for this issue is unknown. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14370 | 1220648-2024-23332 | 20706534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706534&pc=OZD | 05/13/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/13/2022 | 11/18/2024 |  |  | 02/28/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA LD AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023188887 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the staff took the patient to the operating room for a washout due to signs of tamponade. they successfully removed a significant clot. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus could not be determined.
14371 | 1220648-2024-23335 | 20706692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706692&pc=OZD | 11/25/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 11/25/2021 | 11/18/2024 |  |  | 09/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021009599 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient had a gastrointestinal bleed. it was determined that the bleed was from vascular damage to the mitral valve. actions taken to resolve and patient outcome are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14372 | 1220648-2024-23341 | 20706713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706713&pc=OZD | 12/13/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 12/13/2021 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1482671 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited a lowering trend in the purge flow. the staff shifted the p level, however this issue frequently returned. no report of patient harm or injury. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14373 | 1220648-2024-23338 | 20706723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706723&pc=OZD | 12/03/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 12/03/2021 | 11/18/2024 |  |  | 05/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022088769 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced limb ischemia. actions taken are unknown. it was noted that the procedure was successful. the patient was bridged to a left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14374 | 1220648-2024-23336 | 20706756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706756&pc=OZD | 05/04/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 05/04/2022 | 11/18/2024 |  |  | 12/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023166395 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure and the physician decided to explant the device, it was reported that the patient was stable and survived explant. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, there was a piston-cylinder leak. no other damage was found to the pump.
14375 | 1220648-2024-23340 | 20706880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706880&pc=OZD | 11/26/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 11/26/2021 | 11/18/2024 |  |  | 12/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037351 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient was shocked multiple times for ventricular tachycardia. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14376 | 1220648-2024-23342 | 20706910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706910&pc=OZD | 04/28/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 04/28/2022 | 11/18/2024 |  |  | 12/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044600 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. the physician attempted to insert the device; however, he was unable to cross the pump into the ascending aorta. support was aborted. no further information is available. | [5 empty] | The device was not returned for investigation, however clinical details were sufficient to determined root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
14377 | 1220648-2024-23337 | 20706956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20706956&pc=OZD | 05/01/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 05/01/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 2022148758 |  | 005082-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella for mechanical circulatory support. during support, the pump stopped and was able to be restarted. the pump was implanted for more than 14 days. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14378 | 1220648-2024-23344 | 20707045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707045&pc=OZD | 04/28/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/28/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022121691 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the patient developed thrombocytopenia, and the staff administered argatroban. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14379 | 1220648-2024-23339 | 20707060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707060&pc=OZD | 04/29/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 04/29/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022122803 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.5 for mechanical circulatory support. during delivery, the device would not cross through the innominate artery. the physician decided to remove the device and replaced it with an impella cp. no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy.
14380 | 1220648-2024-23346 | 20707079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707079&pc=OZD | 04/27/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/27/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023187922 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced access site bleeding and heparin was withheld. no further information is available. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14381 | 1220648-2024-23345 | 20707154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707154&pc=OZD | 01/20/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/20/2022 | 11/18/2024 |  |  | 06/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023161460 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a clot was found in aorta alongside impella device. the patient had been off heparin previously due to concerns for heparin induced thrombocytopenia [hit] and partial thromboplastin time [ptt] values had decreased. patient was taken to operating room for impella removal and endovascular thrombectomy. it was also reported that the patient had ep consult for afib, possible cardioversion. at that time, p level decreased to p6, slowly weaning off impella support. impella was successfully removed with septal closure device placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14382 | 1220648-2024-23355 | 20707169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707169&pc=OZD | 10/29/2024 | 12/18/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562017 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 was unable to cross the 64-year-old male patient¿s new aortic valve. several unsuccessful attempts were made. the impella procedure was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. clinical details mention patient had multiple complications post bypass with the av not opening very well. the root cause of the delivery issues was most likely patient condition related as evidenced by clinical details corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: delivery issues. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
14383 | 1220648-2024-23348 | 20707185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707185&pc=OZD | 01/28/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/28/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129308 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from gastrointestinal area and esophagus; therefore, heparin was removed. recommendation was made to add heparin in purge. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14384 | 1220648-2024-23358 | 20707193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707193&pc=OZD | 10/29/2024 | 01/02/2025 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518786 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that upon removing the impella cp from the 74-year-old male patient, the 0.35 wire got caught in the peel-away sheath. the physician continued with cp removal. the wire did not break, nor the peel-away split. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿. | The investigation for the removal issues has been completed. the pump was not returned for investigation. the data log review was not required for this case. as per the given clinical details, the 0.35 wire was caught in the peel-away sheath. the removal issue failure mode was changed to a delivery issue since the peel-away sheath and guide wire were utilized during the delivery of the pump. the cause of the delivery issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email. h5 changed to yes.
14385 | 1220648-2024-23359 | 20707207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707207&pc=OZD | 04/14/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 04/14/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123949 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding. the resolution for this issue is unknown. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14386 | 1220648-2024-23349 | 20707210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707210&pc=OZD | 10/24/2024 | 01/02/2025 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563457 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | Yes | Material Puncture/Hole | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support when there was bleeding at the repositioning lock between the blue and white areas. the pump was explanted and the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Product was returned for investigation. data log review was not required for this case run. as per the clinical details, bleeding was reported from the butterfly and white repo hub. no damage was found on the pump or the reposition sheath. the reposition sheath was pressure tested at 350 mmhg, and no leak was observed from the repo unit. the product damage failure mode was changed to access site bleeding, as per the investigated clinical details. the cause of the access site bleeding issue was most likely use related, based on returned product investigation. h6 health effect - clinical code and medical device problem code updated to reflect investigation results g1 reporting contact email revised on manufacturer device report 1220648-2024-23349 in accordance with updated procedures.
14387 | 1220648-2024-23343 | 20707212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707212&pc=OZD | 01/28/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/28/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022108618 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had to go back to operating room for an axillary hematoma. the procedure outcome was noted to be successful. | [5 empty] | Explant date is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14388 | 1220648-2024-23351 | 20707256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707256&pc=OZD | 10/29/2024 | 12/26/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH |  |  | ABIOMED, INC. |  |  | US |  | 2025515603 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during impella 5.5 support. following the initial implant, the pump flipped positioning and attempts to torque the device away from the mitral valve were unsuccessful. the pump was removed to restart the implant process and significant blood loss was observed via the drains. the chest was reopened as the patient became hemodynamically unstable and a perforation was discovered. blood products were given and the perforation was patched to resolve the condition. the pump was then reinserted successfully. shortly after, a significant amount of blood was seen in the sterile sleeve. a defect on the hemostatic valve within the repo sheath was suspected. an additional ligature was added to the catheter to prevent further bleeding into the sterile sleeve. support continued for nine days with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the product damage, ventricle perforation, positioning issues has been completed. product not returned for investigation. the cause of the ventricle perforation was most likely use related as the pump was pushed back and forth repeatedly during repositioning. the positioning issue was most likely patient condition related as the patient had pad/pvd and resistance was felt crossing the av. clinical details states that there was no sterile sleeve damage reported and only bleeding into the sleeve requiring an additional ligature. the cause of the access site bleeding major was not determined since no product was returned for investigation and lack of clinical details. corrections to the initial mfr report: g1 mdr reporting contact email updated. g1 mfr site name and address updated.
14389 | 1220648-2024-23353 | 20707269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707269&pc=OZD | 04/22/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/22/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | CANADA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022149780 |  | 005064 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced hemolysis. the device was repositioned and the hemolysis started to resolve. the patient received blood for the low hemoglobin related to the hemolysis. it was reported that the patient survived explant. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
14390 | 1220648-2024-23347 | 20707284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707284&pc=OZD | 04/25/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/25/2022 | 11/18/2024 |  |  | 01/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023178207 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. during delivery, the pump was unable to pass though the patient vasculature and attempts were aborted. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted.
14391 | 1220648-2024-23356 | 20707294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707294&pc=PYX | 10/29/2024 | 12/16/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025543811 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella rp flex implant and it was aborted. the patient had an enlarged right ventricle (rv). access was achieved with the wire, however during the placement of the flex, it became difficult due to the rv size of the patient. multiple wires were attempted without success. the implant was aborted. there was no reported patient harm. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since patient had an enlarged rv. the following have been reported incorrectly or missing from manufacturer device report. d6b missing e4 should have been left blank g1 reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
14392 | 1220648-2024-23352 | 20707343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707343&pc=OZD | 02/24/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 02/24/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022150907 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced pericardial bleeding. anticoagulation was restricted. the patient outcome was noted as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14393 | 1220648-2024-23357 | 20707361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707361&pc=OZD | 03/09/2022 | 03/07/2025 |  | 11/18/2024 |  |  | 03/09/2022 | 11/18/2024 |  |  | 03/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063511 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced arrhythmias. the resolution for this issue is unknown, it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14394 | 1220648-2024-23350 | 20707489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707489&pc=OZD | 04/25/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/25/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136272 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. during delivery, the pump was unable to pass through the patient's vasculature. the physician decided to remove the device and replace it with an impella cp. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14395 | 1220648-2024-23366 | 20707494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707494&pc=OZD | 05/18/2022 | 05/20/2025 |  | 11/18/2024 |  |  | 05/18/2022 | 11/18/2024 |  |  | 05/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022088342 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | Unspecified Kidney or Urinary Problem | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that the impella was pulled into the aorta. without left ventricle support, the patient's creatine stated to rise, requiring dialysis. it was noted that the improper position of the catheter was not noticed for several days. the implanting physician was off for the weekend and the covering physician didn¿t notice the change in patient's condition. it was reported that the device was normally explanted and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14396 | 1220648-2024-23363 | 20707512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707512&pc=OZD | 04/07/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 04/07/2022 | 11/18/2024 |  |  | 09/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022131575 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure. the staff increased the purge solution heparin concentration and restarted intravenous heparin. the patient survived normal explant. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, biomaterial was found to be built up in the purge gap. however, the root cause of the high purge pressure was determined to be biomaterial buildup due to patient condition. no other damage was found to the pump.
14397 | 1220648-2024-23371 | 20707529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707529&pc=OZD | 12/30/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 12/30/2021 | 11/18/2024 |  |  | 11/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022113971 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 5.0 for hemodynamic support. it was reported that the patient was diagnosed with pancytopenia and experienced major bleedings. the heparin rate was lowered in purge solution to 10ui/ml. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14398 | 1220648-2024-23367 | 20707563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707563&pc=OZD | 10/24/2024 | 12/16/2024 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed for support of a patient who was admitted and arrested three times with many cardiac comorbidities. it was reported that the patient had no pulse in either limb. the patient received a total of five liters of cyrstalloid, four units of packed red blood cells, 500 albumins, and bicarbonate was given. axillary aline was placed. the following day, the patient was not expected to survive and was made do not resuscitate order. it was further reported that the patient expired while on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23367. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14399 | 1220648-2024-23360 | 20707581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707581&pc=OZD | 04/11/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 04/11/2022 | 11/18/2024 |  |  | 10/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022145141 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge and the patient experienced hemolysis. no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device logs were not received from the customer. however, no abnormalities were found with the returned pump. the high pressure was not reproduced; therefore, the root cause of the high purge pressure could not be determined. the relationship between impella and hemolysis is unknown as no details were provided.
14400 | 1220648-2024-23361 | 20707614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707614&pc=OZD | 10/24/2024 | 12/01/2025 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211685 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was prepping for impella support when the automated impella controller had a controller error and failed to recognize the purge cassette. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14401 | 1220648-2024-23372 | 20707637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707637&pc=OZD | 06/24/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 06/24/2022 | 11/18/2024 |  |  | 03/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023195855 |  | 385603 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was going to be on an impella 5.0 for mechanical circulatory support. the device was not able to be implanted due to anatomy issues, the device was unable to pass through the patient's vascular and attempts were aborted. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14402 | 1220648-2024-23354 | 20707705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707705&pc=OZD | 10/29/2024 | 12/30/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533742 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Material Separation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 61-year-old male patient was implanted with an impella cp device for mechanical circulatory support. upon explanting pump, resistance was met and with further tension. the pump came apart at the area of the connection of the motor housing and cannula. the separated portion of pump was contained in the peel away sheath and was obtained when peel away sheath removed. manual pressure held for hemostasis, distal neuro-vascular intact. there was no harm reported. | [5 empty] | The failure mode of "removal issues" was removed and investigated under "difficulty inserting/removing pump through introducer" as the removal issue occurred when removing the pump through the introducer. the root cause of the difficulty inserting/removing pump through introducer was most likely due to a sheath kink. the root cause of the product damage was most likely use related as the delo joint was intact and resistance was encountered when removing pump through introducer sheath kink. an additional failure mode of "product damage (detached inflow/outflow - peripheral vessel) was added as the pump was returned damaged with outflow cage separated. the root cause of the product damage was most likely use related as the delo joint was intact and resistance was encountered when removing pump through introducer sheath kink. the root cause of the difficulty inserting/removing pump through introducer was most likely due to a sheath kink. e4 should have been left blank on manufacturer device report. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
14403 | 1220648-2024-23370 | 20707778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707778&pc=OZD | 06/13/2022 | 10/07/2025 |  | 11/18/2024 |  |  | 06/13/2022 | 11/18/2024 |  |  | 01/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023170876 |  | 005064 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient developed a left ventricle thrombus that blocked the impella purge system. the patient was unstable and required an intra-aortic balloon pump. it was reported that the impella was explanted and the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14404 | 1220648-2024-23382 | 20707793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707793&pc=OZD | 08/09/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/09/2021 | 11/18/2024 |  |  | 08/31/2021 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1438105 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, that there was a pump stop after 30 minutes of use. the pump could not be restarted. the pump was removed, and intervention was continued without support. there was no harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned, and the issue was confirmed. the field data log was reviewed which confirmed multiple impella stopped ¿ motor current high alarms during case start. visual inspection of the returned pump revealed a fractured red lumen within the outflow cage, obstructing the impeller, and extending up into the pigtail. per the results of the clinical review and investigation, the root cause was determined to be a fractured red lumen.
14405 | 1220648-2024-23364 | 20707799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707799&pc=OZD | 03/14/2022 | 03/07/2025 |  | 11/18/2024 |  |  | 03/14/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022113963 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. it was reported that the patient was bleeding, and the staff dressed the area. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14406 | 1220648-2024-23373 | 20707822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707822&pc=OZD | 10/24/2024 | 12/23/2024 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531559 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the peel away sheath was left in place after the insertion of impella cp. the registered nurse stated that the patient experienced significant bleeding from the site overnight continuing through the morning. the patient received three units of packed red blood cells overnight. the physician placed a stitch at insertion site to manage bleeding. however, the bleeding persisted despite transfusion, stitch placement, and discontinuation of systemic anticoagulation. a sandbag was placed but was removed as angle was unable to be maintained due to the sandbag and dressing. the following day, no significant oozing was seen after angiomax was restarted. the patient survived the explant and continued on extracorporeal membrane oxygenation. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of access site bleeding was determined to be use issue because the peel away sheath left in place after discussion regarding using repo sheath causing the bleed. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23373 was submitted in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14407 | 1220648-2024-23375 | 20707851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707851&pc=OZD | 07/29/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 07/29/2022 | 11/18/2024 |  |  | 01/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023180489 |  | 005064 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | Failure to Sense | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited a high-pressure purge alarm which was attributed to kinking of the purge cassette line. the purge cassette and bag were replaced. additionally, there was a placement signal alarm, that was reported to be resolved without a specified resolution. additionally, the patient was diagnosed as heparin induced thrombocytopenia (positive), the resolution for this issue is unknown. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14408 | 1220648-2024-23362 | 20707888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707888&pc=OZD | 04/11/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 04/11/2022 | 11/18/2024 |  |  | 12/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023160079 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced bleeding from the kidney. the resolution for this issue is unknown, it was reported that the patient survived explant. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14409 | 1220648-2024-23365 | 20707963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20707963&pc=OZD | 01/09/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/09/2022 | 11/18/2024 |  |  | 04/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022077219 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [10 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced tachycardia and atrial fibrillation. additionally, it was noted that the patient was in poor condition due to sepsis. the patient also had to be thoracotomized due to bleeding. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14410 | 1220648-2024-23383 | 20708050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708050&pc=OZD | 08/08/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/08/2021 | 11/18/2024 |  |  | 04/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022075922 |  | 005062 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. while on support, placement signal (ps) was not reliable, and introducer sheath was leaking. pump was replaced with another pump. no patient harm was reported. | [5 empty] | The investigation for the reported placement signal issue has been completed. the product was returned, and the placement signal issue was confirmed. the cause of the placement signal issue was a sensor short. the cause of this short could not be determined.
14411 | 1220648-2024-23368 | 20708052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708052&pc=OZD | 06/01/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 06/01/2022 | 11/18/2024 |  |  | 10/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023199398 |  | 005066 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure and low pump flow. the purge pressure was being controlled with medication to reduce the purge rate. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14412 | 1220648-2024-23369 | 20708066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708066&pc=OZD | 01/04/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/04/2022 | 11/18/2024 |  |  | 02/28/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054635 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was noted that during access when 13fr dilator was removed from peel away sheath, there was profuse bleeding coming from the diaphragm of the sheath. the 13fr peel away was taken out and replaced with 14fr peel away sheath from the insertion kit. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14413 | 1220648-2024-23378 | 20708076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708076&pc=OZD | 01/11/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/11/2022 | 11/18/2024 |  |  | 11/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021014087 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient of unknown ethnicity was being placed on impella 5.0 for hemodynamic support. it was reported that the since day fifteen of support, lowering trend in the purge flow, solved with pressure level shifting, but problem returning frequently. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14414 | 1220648-2024-23374 | 20708148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708148&pc=OZD | 06/29/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 06/29/2022 | 11/18/2024 |  |  | 03/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063495 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced a hematoma at the right axillary access site, the resolution for this issue is unknown. additionally, the patient was diagnosed as heparin induced thrombocytopenia (positive) and to help resolve this issue bicarbonate was added to the purge. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14415 | 1220648-2024-23385 | 20708161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708161&pc=OZD | 08/07/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/07/2021 | 11/18/2024 |  |  | 11/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024711 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. it was noted that the impella pump was implanted off center through the peel-away sheath. a five hundred milliliter blood loss was observed in the patient. albumin was provided and the patient¿s hemoglobin was to be drawn in unit. weaning was successful, the device was removed, and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14416 | 1220648-2024-23377 | 20708186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708186&pc=OZD | 06/30/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 06/30/2022 | 11/18/2024 |  |  | 01/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023169583 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced access site bleeding. the resolution for this issue is unknown. however, it was noted that this issue had no significant effect on patient hemodynamics. the patient survived normal explant of the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14417 | 1220648-2024-23380 | 20708202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708202&pc=OZD | 10/22/2024 | 12/19/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558668 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed ischemia during impella cp support. the pump was subsequently removed as a result of the condition. the patient remained stable. the patient survived the explant. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14418 | 1220648-2024-23376 | 20708296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708296&pc=OZD | 01/12/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 01/12/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129448 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that impella site had been bleeding with hematoma at access site. pressure dressing was applied, and systemic heparin stopped. additionally, it was noted that the patient remained on dextrose 5% in water due to heparin induced thrombocytopenia (hit). patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14419 | 1220648-2024-23379 | 20708325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708325&pc=OZD | 08/11/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/11/2021 | 11/18/2024 |  |  | 09/30/2021 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA LD |  |  | OZD |  |  | IMPELLA LD |  |  | ABIOMED, INC. |  |  | US |  | 1446202 |  | 005082-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella ld device for mechanical circulatory support. while on support, bleeding was noted from multiple sites, purge flow slightly decreased increasing purge pressure, suction, and dumping of blood into chest tubes was noted. packed red blood cells and blood products were administered for bleeding. suction resolved by turning pump down to p8 setting. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14420 | 1220648-2024-23384 | 20708498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708498&pc=OZD | 08/08/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/08/2021 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218783 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported the automatic impella controller (aic) had a controller failure alarm. after multiple attempts to restart the aic another controller was used with no report of harm to the patient. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the impella device has been returned for evaluation. the cause of the controller failure alarm was a defective impellatronic board. the cause of the defective impellatronic board is unknown.
14421 | 1220648-2024-23386 | 20708805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20708805&pc=OZD | 10/22/2024 | 12/23/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560839 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient went into torsades and was successfully cardioverted via one shock. amiodarone and magnesium were started. additionally, clot was removed from left atrium followed by aortic valve replacement. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into vf/vt/af/at and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23386 accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14422 | 1220648-2024-23389 | 20709207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20709207&pc=OXD | 07/06/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 07/06/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OXD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023207698 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. it was reported that the device exhibited high-pressure, the physician stated that he did not want to try and resolve this issue since the patient was experiencing bleeding. the patient experienced bleeding from an unknown location and the resolution for this issue is unknown. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14423 | 1220648-2024-23398 | 20709409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20709409&pc=OZD | 03/09/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 03/09/2022 | 11/18/2024 |  |  | 07/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022028016 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device experienced pump stop and thrombus was suspected to have gotten in the gap in the motor. the pump stop occurred beyond the pump's indication of use/duration. upon the stop the team made the choice to remove but not replace the pump as the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14424 | 1220648-2024-23391 | 20709756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20709756&pc=OZD | 10/22/2024 | 12/20/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552967 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a high risk percutaneous coronary intervention (hrpci) patient. the pump was not explanted at the conclusion of the hrpci and was to be on support in the intensive care unit. the site began to bleed after the 14 french was removed and repositioning sheath was placed. to resolve the bleeding, the team had to adjust the perclose and apply figure eight suturing. the patient survived the event. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anticoagulation management because the ptt values are very high and heparin was stopped and direct pressure applied to groin site. g1 reporting contact email revised on manufacturer device report 1220648-2024-23391 in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
14425 | 1220648-2024-23387 | 20709977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20709977&pc=OZD | 07/01/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 07/01/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023207749 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. after multiple attempts to place the impella, they physician was unable to insert the device due to the patient vasculature. the device was removed and replaced with an impella cp. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14426 | 1220648-2024-23397 | 20710000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710000&pc=OZD | 02/21/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 02/21/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022113978 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported that pump was explanted due to bleeding and bad placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14427 | 1220648-2024-23406 | 20710145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710145&pc=OZD | 10/22/2024 | 12/17/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553708 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, pump stop occurred. pump was explanted. the patient survived the event. | [5 empty] | The investigation into pump stop has been completed. the impella device was not returned for evaluation. the log review showed gradual increase in purge pressure, which eventually led to "purge pressure high" alarms. the motor current began to increase after the purge pressure rise, which eventually led to a pump stop with "impella stopped - motor current high" alarms. biomaterial build-up likely occurred with the purge pressure building up, which then caused the motor current to gradually rise due to inadequate purge flow. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23406. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
14428 | 1220648-2024-23390 | 20710232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710232&pc=OZD | 08/31/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 08/31/2022 | 11/18/2024 |  |  | 05/31/2024 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023220062 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that the patient experienced bleeding from various surgical site, it was noted that bleeding at the impella site was minimal. the patient received 2 units of red blood cells, fresh frozen plasma, and pressure bandaging to help resolve this issue. additionally, plavix medication was discontinued and heparin reduced. it was reported that the patient survived normal explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14429 | 1220648-2024-23408 | 20710256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710256&pc=OZD | 09/08/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 09/08/2022 | 11/18/2024 |  |  | 10/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023178287 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, pump stoppage occurred. the pump was restarted after the second attempt. a purge block alarm with purge pressure in the 1000's was also noted. the nurse checked line for kinks and found the purge pressures came down to 800sl. the pump is scheduled for early removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14430 | 1220648-2024-23392 | 20710361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710361&pc=OZD | 08/30/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 08/30/2022 | 11/18/2024 |  |  | 10/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021019963 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was going to be on an impella 5.0 for mechanical circulatory support. it was reported that implantation was aborted because the device could not pass though the patient vasculature. the device was successfully inserted, however upon removal of wire it was notice that the impella was starting to pull back across the aortic valve. the impella brought back and out of the graft and the graft was cut down further to help with insertion. the second attempt was very similar with insertion of impella into anastomosis was very tight. once successfully in the left ventricle, the impella catheter once again was pulled back across the aortic valve. at this time impella support was aborted and patient was placed on central extra-corporeal membrane oxygenation for support. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14431 | 1220648-2024-23401 | 20710366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710366&pc=OZD | 04/29/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 04/29/2021 | 11/18/2024 |  |  | 01/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022042834 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Injury | Yes | Yes | Break | Fluid/Blood Leak | [5 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, it was noted that the filter was broken in the purge system and leakage was observed. additionally, patient was positive for heparin induced thrombocytopenia. argatroban was administered to replace heparin. weaning was successful and the device was removed. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. neither the product nor the data logs were returned for investigation. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely purge sidearm filter damage due to alcohol interaction. this report is being filed as part of a retrospective review of historical records.
14432 | 1220648-2024-23394 | 20710403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710403&pc=OZD | 02/17/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 02/17/2022 | 11/18/2024 |  |  | 12/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023162530 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. delivery issues were noted. the device was switched to impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14433 | 1220648-2024-23399 | 20710414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710414&pc=OZD | 10/22/2024 | 12/18/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553700 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. during the delivery process, it was noted that the #6 french diagnostic catheter was unable to advance through the innominate due to significant calcification. after discussions regarding additional options, the decision was made to abort the implant. the patient was transferred back to the cardiovascular intensive care unit without impella support. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the patient had calcification. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23399 was submitted. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14434 | 1220648-2024-23405 | 20710500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710500&pc=OZD | 07/05/2024 | 12/16/2024 |  | 11/18/2024 |  |  | 10/21/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025491907 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Respiratory Failure | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support an acute myocardial infarction/cardiogenic shock patient admitted in through the emergency room with plans for high risk percutaneous coronary intervention. the support was for 80 hours and allowed for the percutaneous coronary intervention and was successfully weaned and explanted. months later, abiomed¿s long-term outcome and quality indicator (loqi) impella registry database reported an adverse event with an undetermined relationship to the impella use. the patient had suffered acute hypoxemic respiratory failure secondary to congestive heart failure. the patient was medically managed, inclusive of bilevel positive airway pressure, and recovered. | [5 empty] | The investigation for the acute hypoxemic respiratory failure secondary to the congestive heart failure has been completed. the pump nor the data logs were returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the acute hypoxemic respiratory failure secondary to the congestive heart failure could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email. h6 clinical code 1721 changed to 4446 and 2484. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14435 | 1220648-2024-23403 | 20710622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710622&pc=OZD | 10/22/2024 | 12/16/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551034 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that it was unable to doppler pulses for a patient on impella cp support. there was no arterial flow to the right leg. the physician placed an anti-grade sheath next to the impella and connected to the right radial sheath. flow was re-established and the leg had color. it was further reported that the family withdrew care and the patient expired soon after therapies stopped. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14436 | 1220648-2024-23400 | 20710725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710725&pc=OZD | 10/22/2024 | 12/13/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501003 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed ischemia in their right lower extremity during impella cp support. the patient's leg was described to be cold with no distal pulse. it was noted that a distal perfusion catheter had been placed during implant, but the catheter became clotted off during the course of support. the decision was made to wean and explant the device to resolve the condition. the patient was medically managed following explant and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23400. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
14437 | 1220648-2024-23396 | 20710996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20710996&pc=OZD | 10/22/2024 | 12/31/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527107 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella cp. the procedure was aborted due to patient anatomy and there was no reported patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23396 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14438 | 1220648-2024-23393 | 20711192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20711192&pc=OZD | 09/22/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 09/22/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022136242 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced atrial fibrillation, and the staff started intravenous therapy and administered amnio bolus. additionally, the patient experienced a gastrointestinal bleed and heparin was withheld. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14439 | 1220648-2024-23395 | 20711286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20711286&pc=OZD | 02/27/2022 | 05/20/2025 |  | 11/18/2024 |  |  | 03/02/2022 | 11/18/2024 |  |  | 12/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023166401 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that roughly nine days into support, the patient was diagnosed with heparin induced thrombocytopenia. the purge solution was switched from heparin to sodium bicarb in response to the condition. due to the reduction/removal of heparin, the staff was advised to monitor the purge flow and purge pressure for any changes. when high purge pressure was eventually encountered, a replacement pump was scheduled for implant. upon removal of the faulty pump, the patient remained hemodynamically stable, and the decision was made to withhold the replacement. no report of patient harm or injury. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump displayed high purge pressure issues. actions taken to resolve the issue are unknown. the patient continued on support and the patient outcome was noted to be stable. |  |  |  |  | Additional information added to b1, b2, b5, h1, and h6. the investigation of the reported allegations remains ongoing. upon conclusion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14440 | 1220648-2024-23402 | 20711760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20711760&pc=OZD | 03/22/2022 | 04/22/2025 |  | 11/18/2024 |  |  | 03/22/2022 | 11/18/2024 |  |  | 09/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022133724 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited high purge pressure, the resolution for this issue is unknown. additionally, the patient experienced thrombocytopenia, their partial thromboplastin time was 48 and the next day 56. the resolution for this issue is unknown. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. correction to initial mdr 1220648-2024-23402 in section b5, it was reported "additionally, the patient experienced thrombocytopenia, their partial thromboplastin time was 48 and the next day 56.". this is an incorrect statement, and partial thrombin time result is not a indicative of thrombocytopenia. the patient did not have any reported thrombocytopenia nor a decrease in platelet count. the h6 code was incorrectly reported as 4431 thrombocytopenia, and has been corrected to 4582 no clinical signs, symptoms or conditions. | The impella device has not been received from the customer, therefore, evaluation of the device has not been possible. upon conclusion of the investigation, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14441 | 1220648-2024-23388 | 20711924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20711924&pc=OZD | 10/22/2024 | 12/23/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557430 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient¿s right hand was dusky, cool to the touch and pain was present. the following day, the pain was better controlled. right upper extremity remained dusky, but sensation improved. patient had intermittent loss of doppler signals in bilateral lower extremity too with correlated pulse pressure narrow and occasional impella position unknown alarm. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
14442 | 1220648-2024-23407 | 20712006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712006&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/18/2024 |  |  | 10/21/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025529932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 being placed over the .018 guidewire to support a patient admitted in for the surgical repair of the ventricular septal defect (vsd). the wire and pump interacted at the complicated delivery and there was an observation made of what was presumed to be a left ventricle (lv) perforation from the .018 wire. when the wire and pump interacted, there was a kink within the guidewire. the team surgically repaired and then later questioned again the etiology of the perforation, as it could have been a bleed from the vsd and not the lv/wire interaction. the patient later expired when care was withdrawn. | [5 empty] | The investigation for the ventricle perforation and delivery issues has been completed. no product was returned for analysis. the root cause of the ventricle perforation was most likely patient condition related as evidenced by mention of patient's friable tissues. the root cause of the delivery issues was not determined as no product or images were returned for analysis. b.2 required intervention was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23407 and was added. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23407 was submitted. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14443 | 1220648-2024-23409 | 20712351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712351&pc=OZD | 10/22/2024 | 12/12/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547663 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding and limb ischemia. the groin had continuous oozing. the patient had significant bilateral groin ecchymosis. the groin was re-dressed and one unit of packed red blood cells was administered due to hemoglobin dropping. additionally, flow status verified by distal runoff via the re-access port revealed present but extremely sluggish flow. after distal runoff performed, physician decided to explant impella. distal pulse were dopplerable post explant and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the access site bleed and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14444 | 1220648-2024-23414 | 20712454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712454&pc=OZD | 09/14/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 09/14/2021 | 11/18/2024 |  |  | 06/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1484645 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. the impella 2.5 was placed into left ventricle without any issues, pump starts and stopped suddenly, because of high motor current. attempt to restart failed. explantation of the impella and cleaning outside the body. start of impella in sterile field into water. this was successful. re implant of impella and start again with no issues, finish procedure with the device without any further problems. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned, and the issue was not confirmed. the field data logs were reviewed, and impella stopped ¿ motor current high alarms were confirmed. the returned pump was visually inspected, and no abnormalities were noted. the pump was plugged into a console and ran at p-9 with no issues. based on the provided information that the pump stopped, was cleaned, and was able to be restarted, the root cause of the observed pump stop was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
14445 | 1220648-2024-23417 | 20712749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712749&pc=OZD | 10/24/2024 | 12/27/2024 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468641 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed but there was an access site complication. the site had bleeding which prompted the team to intervene. the team gave blood, platelets, and protamine to reverse anticoagulation. a jackson-pratt drain was placed. the pump remained on for support for over nine days. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14446 | 1220648-2024-23431 | 20712752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712752&pc=OZD | 09/13/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/13/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483514 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, patient experienced gastrointestinal [gi] bleeding. as a result, heparin was withheld, and left ventricle assist device (lvad) support was being planned. patient survived the procedure. | [5 empty] | The investigation for the reported vascular damage - peripheral vessel has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the vascular damage - peripheral vessel could not be determined.
14447 | 1220648-2024-23426 | 20712814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712814&pc=OZD | 10/29/2024 | 12/20/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517682 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump to support a 43-year-old female patient admitted in with a history of st-segment elevation myocardial infarction (stemi), heart failure (h)f, and an automated implantable cardioverter defibrillator (aicd). the patient was in need of percutaneous cardiac insertion (pci) for coronary microvascular disease (mvd) coronary artery disease (cad) and the impella cp was placed. the pump supported for over 8 hours and was explanted due to a groin site bleed and hemolysis concerns. the pump was explanted and blood products were given. | [5 empty] | The investigation for the hemolysis and access site bleed has been completed. product and data logs were not returned for review. based on clinical description, the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14448 | 1220648-2024-23428 | 20712835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712835&pc=OZD | 10/29/2024 | 12/27/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523303 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 75-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impellas support and as soon as pump was started, notification immediately appeared saying impella stopped. an attempt was made to restart impella per protocol, but impella would not restart. impella was replaced. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped: ¿if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. | The investigation for the pump stop has been completed. the pump was returned for investigation. a catheter kink was noted at the kink protector, which is sign of applied excessive force. no other physical damage was observed on the returned product. the pump failed electrical testing on one motor cable line. further investigation revealed that the red motor cable line was opened inside the red handle. the data log shows the pump was prepared and placed in the patient. upon start, the impella stopped 119 alarm was seen. as per the clinical details, the pump was not able to start after placement. the cause of the pump did not start issue was an open red electrical motor line inside red handle, which is likely due to excessive force applied during the procedure. corrections to the initial mfr report:g1 mdr reporting contact email.
14449 | 1220648-2024-23416 | 20712844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20712844&pc=OZD | 09/24/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 09/24/2021 | 11/18/2024 |  |  | 09/30/2021 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1463393 |  | 005042 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the pump stopped after over 2 hours of support. tried to restart the impella several times, replaced console and eventually the impella was removed. the patient decompensated and needed medical support for the procedure. the patient's pci was cancelled and rescheduled for when impella support possible. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned, and the issue was confirmed. the field log was reviewed, and a sudden pump stop alarm 119 (electrical failure) occurred 2 hours into the run. there were no purge alarms prior to the event. the returned pump was visually inspected, and no abnormalities were observed. the pump was connected to a console and the pump stop was reproduced. electrical resistances were measured, and all were found to be above specification. the red handle was opened and purge solution poured out of the handle. the purge line was pressurized via syringe with water connected to the sidearm and a leak was reproduced with droplets from the catheter side of the strain relief. per the results of the clinical review and investigation, the root cause was determined to be an electrical short due to purge ingress into the red handle from the catheter. there was likely excessive force applied to the catheter during pci resulting in the purge line leak. this report is being filed as part of a retrospective review of historical records.
14450 | 1220648-2024-23412 | 20713264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713264&pc=OZD | 08/10/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/10/2021 | 11/18/2024 |  |  | 11/30/2021 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1471862 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during the implantation, there was a failure to deliver the device due to the patient¿s anatomy. the pump was removed without any harm to the patient and the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
14451 | 1220648-2024-23419 | 20713302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713302&pc=OZD | 10/24/2024 | 12/13/2024 |  | 11/18/2024 |  |  | 10/24/2024 | 11/18/2024 |  |  | 08/31/2025 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391654 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed an infection at their access site during impella 5.5 support. it was documented that an abscess was found at the impella site. antibiotics were utilized to treat the infection and support with the implanted pump remained ongoing. the patient remains on impella support. | [5 empty] | The date of explant was added. the investigation for the infection has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the infection could not be determined. revised reporting contact email since the initial manufacturer device report number 1220648-2024-23419 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
14452 | 1220648-2024-23415 | 20713611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713611&pc=OZD | 05/30/2024 | 12/13/2024 |  | 11/18/2024 |  |  | 10/23/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486155 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured atrial fibrillation with a possible relationship to the impella device used: the participant developed atrial fibrillation. rate was controlled with intravenous amiodarone. cardioversion was done to restore normal sinus rhythm. amiodarone continued for months. the patient recovered with no sequelae. the pump was weaned successfully and the patient survived the explant. | [5 empty] | The investigation for the reported atrial fibrillation (af) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the af could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14453 | 1220648-2024-23424 | 20713654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713654&pc=OZD | 10/29/2024 | 12/27/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545787 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a 56-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump was explanted after just 25 minutes due to optical signal loss. the signal loss was troubleshot first by automated impella controller (aic) replacement. patient on support with the 2nd impella cp without harm or injury from interruption. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. product was not returned. the data logs were reviewed which showed "case start optical_sensor_error_controller" events and "case start optical_sensor_error_catheter" events. the 5s parameters were out of specification with snr being zero. the imc logs indicated an eeprom reading issue with "placement signal not reliable" alarms on both consoles (ic4406 and ic4407), which confirms this to be a pump issue. this could have been caused by a damaged optical fiber or other fiber connected to the eeprom, but it cannot be verified without the returned product. immediately upon connection to aic, placement signal alarms began to trigger indicating an out-of-box failure. the cause of the optical signal issue was most likely a damaged optical fiber line but the exact location and nature of the damage is unknown since the product was not returned. corrections to the initial mfr report: g1 mdr reporting contact email
14454 | 1220648-2024-23420 | 20713807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713807&pc=OZD | 02/16/2022 | 03/10/2025 |  | 11/18/2024 |  |  | 02/16/2022 | 11/18/2024 |  |  | 08/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121491 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump flow was low during impella 5.5 support. the staff added tpa per protocol to the pump flow bag to troubleshoot the issue. support continued with the implanted pump, but the family later decided to withdraw care and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. upon conclusion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14455 | 1220648-2024-23421 | 20713817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713817&pc=OZD | 10/29/2024 | 11/22/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526800 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that an impella cp pump experienced sudden high purge pressure coinciding with low purge flow during patient support. a purge-lysis protocol was sent to the team as requested, however a purge system blocked alarm appeared shortly after. eventually a pump stop was documented and the pump was explanted. there was no patient harm resulting from the failures. |  |  |  |  | The additional information is reflected in b2, b3, g6, h2. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
14456 | 1220648-2024-23432 | 20713861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713861&pc=OZD | 09/13/2024 | 05/08/2025 |  | 11/18/2024 |  |  | 09/13/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484099 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, flow decreased in the impella 5.5 due to possible clot ingestion. no additional information was provided. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23432. g1 reporting contact email revised on manufacturer device report 1220648-2024-23432 in accordance with updated procedures.
14457 | 1220648-2024-23411 | 20713941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20713941&pc=OZD | 10/18/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 10/18/2022 | 11/18/2024 |  |  | 11/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022024674 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced hemolysis while on the impella 2.5 with a mixed venous in the 40-50%'s. the device was explanted and replaced with an impella cp. after being placed on the impella cp, the hemolysis was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14458 | 1220648-2024-23404 | 20714061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714061&pc=OZD | 02/15/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 02/15/2022 | 11/18/2024 |  |  | 02/28/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044173 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support there was high purge pressure during use of the impella device. the pump was removed and replaced without any harm to the patient. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14459 | 1220648-2024-23429 | 20714363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714363&pc=PYX | 09/13/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/13/2024 | 11/18/2024 |  |  | 08/31/2025 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024397327 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with right heart failure was implanted with an impella rp flex device for mechanical circulatory support. during implant, impella did not fully advance even after multiple wire changes and buddy wire usage. impella pump was then moved to the right internal jugular (rij) but was not able to fully advance despite buddy wire usage. implant procedure was eventually aborted with no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy.
14460 | 1220648-2024-23439 | 20714491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714491&pc=OZD | 09/26/2024 | 05/20/2025 |  | 11/18/2024 |  |  | 09/26/2024 | 11/18/2024 |  |  | 03/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063513 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported, a patient was implanted with an impella device. for mechanical circulatory support. the device had issues of a defective sheath in box. and was unable to be used. no patient harm was reported. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14461 | 1220648-2024-23418 | 20714499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714499&pc=OZD | 02/18/2022 | 03/17/2025 |  | 11/18/2024 |  |  | 02/18/2022 | 11/18/2024 |  |  | 06/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1493418 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted for high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. during impella removal, the perclose sutures broke. multiple additional perclose were utilized without success. a balloon tamponade and a covered stent were subsequently placed to achieve hemostasis. a hematoma also developed in the left axillary site, but was managed successfully with the stent and manual pressure. the patient survived the procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella 2.5 device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14462 | 1220648-2024-23410 | 20714506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714506&pc=OZD | 10/22/2024 | 12/17/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532999 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support due to st elevated myocardial infarction. while on support, bleeding was noted at the groin. anticoagulation on hold due to groin site bleeding. femstop was placed. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14463 | 1220648-2024-23427 | 20714602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714602&pc=OZD | 09/13/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/13/2024 | 11/18/2024 |  |  | 07/31/2025 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379205 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. during implant, delivery issues were encountered. the cutdown was performed in standard fashion. during patient support, the attempted arterial access presented without flow for over an hour. attempt to implant impella 5.5 was unsuccessful. procedure was ultimately aborted with no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be patient condition because flow no flow was achieved even after long time.
14464 | 1220648-2024-23413 | 20714722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714722&pc=PYX | 10/21/2024 | 12/23/2024 |  | 11/18/2024 |  |  | 10/22/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025560585 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support. while on support, the impella rp flex experienced inability to flow >2.9 liters per minute regardless of p-level. additionally, the patient developed hemolysis. the impella rp flex was explanted due to hemolysis. upon explant of the impella rp flex, biomaterial/suspected old white clot was visualized in the inlet. the patient survived the explant. | [5 empty] | The investigation for the hemolysis, low pump flow, and thrombus has been completed. data logs were investigated, and it was confirmed that pump flows were below specification throughout the entire case. there are no purge trends or suction alarms. the placement signal (ps) is trending up gradually throughout the case. while running at p-6 without change - the motor current is unstable. both the trend of the ps and motor current (mc) abnormality indicate something was disrupting flow. clinical details report that neither adding blood volume nor repositioning the pump resolved the complication, but a clot was seen on the pump's inlet and there were a several old indwelling lines. product was not returned for investigation. based on clinical details provided and data log review, the root cause of the low pump flows was determined to most likely be biomaterial. data logs were investigated and there were no suction alarms or purge trends. as mentioned above, the placement signal was gradually trending upward throughout the case and the motor current was abnormally variable while running at p-6. clinical details report that giving blood volume nor repositioning the pump did not resolve the complication. that being said, thrombus was found at the inlet of the pump upon explant and there were several old indwelling lines. the variability in the motor current, lack of purge pressure spike, and slight trend upward of the placement signal over the course of the case, aligns most with gradual biomaterial build up. the root cause of hemolysis was determined to most likely be biomaterial according to clinical details and data log review. the root cause of the thrombus could not be determined without additional details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14465 | 1220648-2024-23423 | 20714873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714873&pc=OZD | 10/29/2024 | 01/02/2025 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528595 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 72-year-old male patient had an impella cp implant and after the implant, there were positioning issues and low pump flow. the patient had a small frame with a steep aortic arch in which the angle to the heart caused the cannula to kink just below the outlet cage. the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow and positioning issues has been completed. the impella cp pump was returned and was visually inspected. there was a cannula kink near the outlet observed. no broken blades were found. no biomaterial was observed. no other abnormality was found. the root cause of low pump flow and positioning issues were patient condition related due to the patient's anatomy who is small in size with steep aortic arch in which the angle to the heart caused the cannula to kink just below the outlet cage. d.9 device return date/information was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-23423 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23423 was submitted in accordance with updated procedures.
14466 | 1220648-2024-23441 | 20714890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714890&pc=OZD | 07/20/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 07/20/2021 | 11/18/2024 |  |  | 06/30/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1495906 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 5.0 device for mechanical circulatory support. during implantation, the grant was sewn on, wire crossed into the left ventricle (lv), however despite multiple attempts, the device could not make the final turn into the innominate. the 5.0 was aborted and the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14467 | 1220648-2024-23446 | 20714895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20714895&pc=OZD | 09/05/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/05/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528991 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was oozing at the insertion site. this was addressed with manual pressure, holding heparin, and platelet infusion. the patient ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was most likely patient condition since the patient was coagulopathic. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14468 | 1220648-2024-23436 | 20715221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20715221&pc=OZD | 10/29/2024 | 12/23/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531460 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an impella cp implant and after the implant, the patient blood pressure dropped with negative left ventricle diastolic pressure noted. it was discovered there was bleeding at the right external iliac. patient¿s groin and abdomen were noted to be distended. the pump was explanted and replaced. stents were used to control the bleeding. four units of blood products were given to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Due to lacking clinical details and product not being returned for investigation, the root cause of the vascular damage could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23436.
14469 | 1220648-2024-23437 | 20715402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20715402&pc=OZD | 10/29/2024 | 01/31/2025 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557600 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a 72-year-old female patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump had persistent low flows due to patient anatomy. there was a small patient with a complication form the angle of the femoral approach. team attempted to resolve but flows remain low and pump remained on for support for 27 hours. patient expired on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the low pump flow has been completed. no data logs returned for investigation. the pump returned cut. no biomaterial was around the impeller. a kink in the cannula was observed. low flows were reported on the first day of support. the cause of the low flows was most likely a kinked cannula. d.9 device return date/information was added. b.1 product problem was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23437 and was added. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-23437 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number was submitted in accordance with updated procedures. revised manufacturing site name and address as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23437. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
14470 | 1220648-2024-23442 | 20715448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20715448&pc=OZD | 07/22/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 07/22/2021 | 11/18/2024 |  |  | 01/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022031987 |  | 005064-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported an outflow blocked alarm was received from the impella device. it was requested that the impella be pulled out two centimeters however, the medical staff refused. a pinhole was observed in the pressure reservoir. it was reported that dextrose with ten percent water was utilized to maintain purge pressure. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The investigation for the reported high purge pressure/product damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the purge leak was sidearm reservoir damage. past investigations of similar complaints indicate that the root cause of the pin hole leak was most likely improper insertion of a needle into the pressure reservoir. this most likely punctured the membrane and caused the sidearm leak. this report is being filed as part of a retrospective review of historical records.
14471 | 1220648-2024-23425 | 20715521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20715521&pc=OZD | 10/29/2024 | 09/26/2025 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525341 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Malposition of Device | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  | Arrhythmia | Hematoma | [11 empty] | Required Intervention | Life Threatening |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | It was also reported on (b)(6) 2024, the pump experienced positioning issues during which the device was oriented towards the mitral valve and interacting with the mitral apparatus. repositioning attempts failed resulting in reports of ectopy/arrhythmia. support continued with the compromised positioning and the patient experienced intermittent runs of nonsustained ventricular tachycardia. the patient has not yet been weaned or explanted. | The complainant reported a 52-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient had a hematoma at the access site. heparin was withheld, a drain was used and compression was held. support continued. | The complainant also reported that the pump was experiencing placement signal issues. the pump's placement signal was low. position was confirmed and lows were confirmed to be adequate. support continued and staff monitored. |  |  |  | Additional information added to b5 and h6. the impella device is still in use, and thus an evaluation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added. b2 outcome life threatening was erroneously selected on the initial manufacturer device report number 1220648-2024-23425-1. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-23425 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-23425. g6 type of report; 30 day was inadvertently omitted on the initial manufacturer device report number 1220648-2024-23425-1. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were not returned as well. without the pump and data logs, the failure cannot be further investigated. therefore, the root cause of the optical signal issue was not determined since product and data logs were not returned. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issues. h6 updated to include optical signal coding. corrections that were made from the initial. manufacturer device report number 1220648-2024-23425. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
14472 | 1220648-2024-23443 | 20715584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20715584&pc=OZD | 08/29/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/29/2021 | 11/18/2024 |  |  | 04/30/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022083137 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during placement, it was reported that the impella was noted to be under the papillary muscle. the device was repositioned. it was reported that the link kinked, however no visual holes were visualized. the patient was transferred to the operating room for a cholecystectomy. bleeding complications occurred, and blood products were required. the patient experienced ventricular tachycardia, and the medical staff made an unsuccessful attempt to reposition the pump. it was reported that after twenty-two days of support, low pump flows on the impella 5.0 prompted removal and replacement of the device. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported issues has been completed. the impella device has not been returned for evaluation. the cause of the reported issue was not determined as no product or data was returned. this report is being filed as part of a retrospective review of historical records.
14473 | 1220648-2024-23440 | 20716310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20716310&pc=OZD | 07/16/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 07/16/2021 | 11/18/2024 |  |  | 07/31/2022 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1498950 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was persistent oozing at the left groin. graft site bleeding required daily blood transfusions. the medical staff planned to utilize ¿v-a ecpella¿ overnight to allow the patient¿s ventricular arrhythmias to subside. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14474 | 1220648-2024-23435 | 20716382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20716382&pc=OZD | 07/23/2022 | 05/28/2025 |  | 11/18/2024 |  |  | 07/23/2022 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1466562 |  | 005060 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported while the patient was in the operating room on impella 5.0 support, the automated impella controller experienced purge issues with no resolution specified. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Console logs or rlm journals from the day of reported event were not available for analysis. testing of the unit (aic and rlm) performed by aachen fs did not reveal any anomalies. since the occurrence of reported event in 2022, the console (and its rlm) were subject to routine annual maintenance and upgrades, and no deficiencies of the impellaconnect system were recorded. repair: perform preventive maintenance of the console.
14475 | 1220648-2024-23422 | 20716425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20716425&pc=OZD | 10/29/2024 | 06/26/2025 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510223 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant had an impella 5.5 pump support placed for an acute myocardial infarction / cardiogenic shock (ami/cgs) patient. patient additionally had intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation (ecmo). on the day of implant the ps was lost, as there was no left ventricle (lv) signal visible. pump remains on for support without harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the optical signal issue and its relation to the impella device was not determined since the product, data logs, and sufficient clinical information were not provided. the additional information is reflected in g3, g6, h1, h6, and h10. | The product was returned and there were no placement signal alarms that were reproduced. the 5s parameters were within specification. there were no data logs returned or images from the case. due to the lack of clinical information provided and data logs, the root cause of the optical signal issue cannot be determined. d9 device returned to manufacturer date was inadvertently omitted on manufacturer device report number 1220648-2024-23422-1. h6 codes were erroneously entered on manufacturer device report number 1220648-2024-23422.-1. h10 narrative was erroneously entered on manufacturer device report number 1220648-2024-23422.-1.
14476 | 1220648-2024-23430 | 20716677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20716677&pc=OZD | 10/29/2024 | 12/31/2024 |  | 11/18/2024 |  |  | 10/29/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after impella cp started to ramp up, flows decreased. the volume stayed less than 0.5 l per minute with no change even after attempting a better position. the impella cp was exchanged for another one. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the data logs showed after pump started, motor current spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. the product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. e4 should have been left blank on manufacturer device report 1220648-2024-23430. g1 reporting contact email revised on manufacturer device report 1220648-2024-23430 in accordance with updated procedures. h6 code 4642 was reported incorrectly on manufacturer device report 1220648-2024-23430. new code was added to health effect - impact code.
14477 | 1220648-2024-23438 | 20716813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20716813&pc=OZD | 06/20/2023 | 11/18/2024 |  | 11/18/2024 |  |  | 06/20/2023 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | CONTROL,,PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351867 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient's nurse stated that the automated impella controller requested a cardiac input every hour and sometimes twice an hour. there was no reported interventions performed to resolve the issue. it was noted that the patient was escalated to an impella 5.5 from the impella cp. the patient's procedure outcome was noted to be expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported alarm issue has been completed. the console was returned for investigation but the field events were not reproduced. the patient's instability was a key factor in the cardiac output messages that was not recreated during the investigation. the console was fully functional through out all testing. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
14478 | 1220648-2024-23449 | 20717173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717173&pc=OZD | 09/04/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/04/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493266 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during placement, the pump was unable to be advanced, so it was removed and replaced with a new pump. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue is determined to be use issue because the second pump was successfully inserted. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14479 | 1220648-2024-23434 | 20717265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717265&pc=OZD | 02/15/2022 | 11/18/2024 |  | 11/18/2024 |  |  | 02/18/2022 | 11/18/2024 |  |  | 07/31/2023 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC |  |  | US |  | 2022114079 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 64-year-old male patient was implanted with an impella 5.0 pump for mechanical circulatory support. roughly ten days into support, the pump began to experience high purge pressure. lyse therapy was completed successfully in response to the condition and the staff was advised to call support if additional assistance was required. support continued with the implanted pump following resolution of the issue. | [5 empty] | The impella device was not returned from the customer, therefore, evaluation of the device could not be completed. an investigation into the allegation was completed. based on the provided information, it was determined that the presence of biomaterial likely cause the increase in pressure. should additional relevant information become available, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
14480 | 1220648-2024-23447 | 20717329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717329&pc=OZD | 09/05/2024 | 02/11/2025 |  | 11/18/2024 |  |  | 09/05/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. after placement, the pump stopped and could not be restarted. the pump was explanted, and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. the product and data logs were not returned for investigation. the cause of the pump stop could not be determined without data logs and returned product. g.1 revised reporting contact email since the initial manufacturer device report number in accordance with updated procedures.
14481 | 1220648-2024-23444 | 20717454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717454&pc=OZD | 08/26/2021 | 11/18/2024 |  | 11/18/2024 |  |  | 08/26/2021 | 11/18/2024 |  |  | 12/31/2021 |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1463041 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Pneumonia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient developed an infection / pneumonia . it was reported that the impella device was removed due to the complication. weaning was successful, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14482 | 1220648-2024-23450 | 20717476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20717476&pc=OZD | 09/04/2024 | 11/18/2024 |  | 11/18/2024 |  |  | 09/04/2024 | 11/18/2024 |  |  |  |  | 11/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532424 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's right and left groins were found to be oozing. the left groin was addressed with mattress sutures x4, and the right groin received a femstop. the family withdrew care and patient expired, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue most likely patient condition due to multiple bleed sites as per the clinical description provided. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14483 | 1220648-2024-23311 | 20702816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702816&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after sheath exchange, the patient developed a hematoma. manual pressure was held. the patient was then transferred and the hematoma was noted to have significantly grown. the decision was made to explant the impella. a balloon was inflated over the insertion site. extravasation still occurred post closure. stenting of the insertion site with a covered stent was then performed and post dilation with a balloon. the patient required two units of blood. the patient survived the explant and is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14484 | 1220648-2024-23312 | 20702886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702886&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. following insertion, it was documented that the pump was unable to advance through the distal aorta. imaging of the area noted calcification and a lesion in the distal aorta. the pump was removed and a 14 french x 25 centimeter introducer was inserted to bypass the lesion. once again, the staff was unable to advance the sheath past the lesion and the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into delivery issues and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue and difficulty removing introducer was most likely patient condition since resistance was noted on the distal aorta due to the patients calcified vessels. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23312 was submitted. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14485 | 1220648-2024-23313 | 20702890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702890&pc=OZD | 10/23/2024 | 12/19/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025514716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced oozing from their access site during impella cp support. a new dressing was placed and reinforced with a sandbag to manage the condition. two units of packed red blood cells were transfused to counteract the blood loss. support was continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D.6b explant date was added. the investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issue could not be determined. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23313 was submitted in accordance with updated procedures.
14486 | 1220648-2024-23314 | 20702902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702902&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508839 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. it was documented that following insertion, the pump would not advance beyond the innominate due to the patient's anatomy. the implant was aborted and an impella cp was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since insertion was aborted due to patient's anatomy. corrections to the initial mfr report: g1 mdr reporting contact email.
14487 | 1220648-2024-23315 | 20702910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702910&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and the patient had ventricular tachycardia/ventricular fibrillation. the patient was defibrillated. the pump was explanted the following day and the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23315 was submitted in accordance with update procedures. h.6 health effect - impact code 4627 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23315 and was added.
14488 | 1220648-2024-23316 | 20702916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702916&pc=OZD | 10/23/2024 | 12/17/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557189 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump being placed for a patient admitted in and found to have coronary artery disease. the pump was unable to be delivered despite all measures of iliac intervention, wire and sheath exchange. the impella cp placement was aborted and there was no reported patient harm. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since patient had significant iliac tortuosity as well as calcification. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23316. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14489 | 1220648-2024-23317 | 20702922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702922&pc=OZD | 10/23/2024 | 12/24/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557260 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was inserted via direct aortic conduit graft but struggled to achieve optimal flows due to the patient¿s anatomy. a decision was made to try and insert via right axillary conduit graft to achieve optimal pump placement but was unsuccessful. after many attempts, the decision was ultimately made to abort impella placement. the patient survived the explant. | [5 empty] | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed low flow with pump was in improper position corresponded with clinical description that triggered impella position in aorta & suction alarms. no other issues were found on the logs. product was not returned for review. the cause of the low pump flow issue was patient condition related since the patient's anatomy prevented adequate positioning in the lv. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14490 | 1220648-2024-23318 | 20702926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702926&pc=OZD | 10/23/2024 | 12/19/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529133 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp insertion, the impella cp was unable to be advanced through the aortic stent and impella cp placement was aborted. the patient¿s original aortic stent was highly calcified and narrowed. there was no reported patient harm and an intra-aortic balloon pump was placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since insertion was aborted due to patient's anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23318 accordance with updated procedures.
14491 | 1220648-2024-23319 | 20702932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20702932&pc=OZD | 10/23/2024 | 12/24/2024 |  | 11/17/2024 |  |  | 10/23/2024 | 11/17/2024 |  |  |  |  | 11/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534734 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump being placed for support for a patient with an acute myocardial infarction/cardiogenic shock. the pump was being placed but native anatomy prevented the delivery. there was no patient harm. | [5 empty] | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23319 g1 reporting contact email revised in accordance with updated procedures. g4 pma/510(k)number missing | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14492 | 1220648-2024-23298 | 20701288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701288&pc=OZD | 10/25/2024 | 12/18/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that one day following implant, an automated impella controller (aic) displayed a placement signal unreliable alarm. both the aorta and left ventricle signals were lost as a result of the failure. the aic was replaced to resolve the issue. there was an initial period of delay of therapy. the pump was attached to the replacement console and hemodynamic support was provided to the patient. the patient's condition was poor and subsequently demised (expired) due to anoxic brain injury which is unrelated to the use the automated impella controller. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into optical signal issue has been completed. the data and device analysis confirmed the reported failure mode. the root cause was the oscillating contrast and unreliable placement signal were due to a bad lamp. g1 reporting contact email revised from the initial submission of manufacturer device report in accordance with updated procedures.
14493 | 1220648-2024-23299 | 20701289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701289&pc=OZD | 10/25/2024 | 12/20/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556172 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp occluding flow requiring to be removed to restore flow. upon removal of the cp, clot down arm on impella side was found requiring cut down for removal. the patient¿s outcome is unknown. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the thrombus and limb ischemia has been completed. product was returned and evaluated. no abnormalities were found with the pump/repositioning sheath/introducer. logs are not applicable. without additional clinical details, the root cause of the thrombus and limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14494 | 1220648-2024-23300 | 20701292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701292&pc=OZD | 10/25/2024 | 01/02/2025 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519928 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. on the same day as implant, it was reported that the pump developed a flat motor current coinciding with a low flow rate. suction alarms sounded despite proper positioning being verified. volume and blood products were administered in response to the issue and the staff was advised to replace the pump. despite the recommendation, support with the implanted pump continued. three days later, the patient came back from a planned operation that resulted in cardiopulmonary resuscitation and was noted to be experiencing some oozing from the impella site. blood products were again given to treat the condition along with unspecified surgical intervention and five hours of coagulation management. the pump remained in use following the intervention. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow and the access site bleed has been completed. the pump was not returned for evaluation. data logs were reviewed finding many suctions occurred. many impella positioning unknown alarms were observed. no motor current (mc) spike were observed. the cause of the low pump flow issue could not be determined since the complaint device was not returned for analysis and lack of clinical details. the cause of the access site bleeding - major issue could not be determined since the complaint device not returned for analysis and lack of clinical details. revised reporting contact email since the initial manufacturer device report number 1220648-2024-23300 was submitted in accordance with updated procedures.
14495 | 1220648-2024-23301 | 20701325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701325&pc=OZD | 10/25/2024 | 09/22/2025 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562040 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella 5.5 support and during support there were placement signal not reliable alarms. there were no change in patient hemodynamics or motor current. support continued. additionally, the patient had a hematoma at the access site. the intervention for the hematoma is unknown. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the access site bleed has been completed. data logs are not relevant to the complication. returned product was investigated and there were no abnormalities with the repositioning unit that would suggest a cause for the complication. however, the sterile sleeve was noted to have been ripped/torn off of the hub with clear force. this was not reported as a complication. clinical details provide limited information regarding what may have caused the bleed or how it was resolved. based on limited clinical details and no abnormalities with returned product related to the complication, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs were investigated, and it was confirmed that the placement signal not reliable (psnr) alarm triggered on (b)(6) 2025. at that time, the signal-to-noise ratio (snr) dropped to zero, contrast began oscillating between 3200/-3200, lamp remained at 100%, and gain/light both dropped. the first alarm was active for 24 hours, resolved for 3 hours, then triggered again for 66 hours, and then resolved for the remainder of the case. returned product was investigated and 5s parameters were all within specification, the psnr alarm did not reproduce when the pump was connected to a console, and the pump passed the laser continuity test. upon reviewing the data logs and the pump, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-23301. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
14496 | 1220648-2024-23302 | 20701329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701329&pc=OZD | 10/25/2024 | 12/23/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561183 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had been attempting to place a 5.5 pump to support a patient coming in with aortic stenosis and having a transcatheter aortic valve replacement. the pump failed to deliver despite all measures and troubleshooting. the team removed the 5.5 and plans for a cp to be delivered. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, device was unable to pass through subclavian despite numerous attempts/interventions as per the clinical description provided. g1 reporting contact email revised on manufacturer device report 1220648-2024-23302 in accordance with updated procedures.
14497 | 1220648-2024-23303 | 20701330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701330&pc=OZD | 10/25/2024 | 12/23/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534688 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the impella was found to be kinked between the radiopaque marker and the motor. it was only able to achieve flows of 1.5 l/min throughout the case. the patient survived the event. | [5 empty] | The investigation for the product damage has been completed. the impella was not returned for evaluation. however, clinical details were provided which state that the impella flow was low and not achieved (1.5 l/min) throughout the case and cannula kink observed after insertion. as a result, the root cause of the low pump flow issue most likely was cannula kink. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
14498 | 1220648-2024-23304 | 20701332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701332&pc=PYX | 10/25/2024 | 09/25/2025 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  | 01/31/2026 |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025462111 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that multiple complications were encountered during impella rp flex support. following implant, the patient experienced bleeding from a chest closure resultant of a concurrent planned procedure. the patient was given 3000 units of heparin immediately prior to the bleed during impella rp implant. protamine was given and a heparin drip was stopped to manage the condition. the staff opted to maintain support with the implanted pump while utilizing other means to monitor its performance. the patient survived the event. | The following day, it was reported that the placement signal became unreliable. troubleshooting was completed, but the issue remained. the staff opted to maintain support with the implanted pump while utilizing other means to monitor its performance. |  |  |  |  | The investigation into vascular damage - peripheral vessel has been completed. the root cause of the vascular damage - peripheral vessel was not determined as limited clinical information was provided. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23304. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the failure mode 'optical signal issue' was changed to 'placement signal issue' as failure is related to the dp sensor. the returned product was inspected and there were no physical abnormalities. the sensor was not damaged. the 5s parameters were within normal range. the pump was run in the flow loop test and the drift was reproduced. the data logs show placement signal drifting down and pump flows drifting up before placement signal and cvp went out of range. the logs confirm alarms #1051 - dp signal out of range and #1052 - epm sensor failure. later during the case there was also alarm #1053 - optical signal out of range. there were no motor current spikes and no position related alarms. the root cause of the placement signal issue was determined to be due to sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5. updated to include placement signal issue description. d10. concomitant medical products and therapy dates. h6. updated to capture placement signal issue and investigation.
14499 | 1220648-2024-23305 | 20701352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701352&pc=OZD | 10/25/2024 | 12/23/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490350 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the physician was concerned by signs of hemolysis including hematuria and labs hemolyzing. the decision was made by the physician team to replace the pump. the pump was explanted and inspected and there was no visible clot present on pump. the pump was replaced. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis has been completed. visual inspection found biomaterial inside the purge gap. no other issues were found on the rest of the pump. the root cause of hemolysis was most likely biomaterial ingestion. g1 reporting contact email revised on manufacturer device report 1220648-2024-23305 in accordance with updated procedures.
14500 | 1220648-2024-23306 | 20701358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701358&pc=OZD | 10/25/2024 | 12/13/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556282 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was inserted and high risk percutaneous coronary intervention was done. after a few hours in the intensive care unit, no pulse was found any longer and pain to the right leg was noted. the patient was taken to the catheterization lab to have the cp removed. upon removal thrombus was noted and patient was taken to the vascular operating room to have thrombectomy. the patient now has positive pulses and good perfusion down the leg. the patient survived the event. | [5 empty] | The investigation into limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23306. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
14501 | 1220648-2024-23307 | 20701360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701360&pc=OZD | 10/25/2024 | 06/26/2025 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025524542 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported there was a small leak from the clear plastic side arm of the impella 5.5. the staff had trouble locating the exact location of the leak. there was only around a drop per hour. support continued and staff monitored. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: cleaning : ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. the pump was successfully purged. after 30 minutes a small drop of purge was observed leaking. the leak was observed from the yellow luer connection from cassette to the sidearm. the leak was very small and was from damage to the yellow luer connection. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc. d9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report 1220648-2024-23307 in accordance with updated procedures.
14502 | 1220648-2024-23308 | 20701369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701369&pc=OZD | 10/25/2024 | 12/23/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557265 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had severe stenosis at innominate takeoff and severe calcification of aorta. the impella 5.5 placement was aborted. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of severe stenosis at innominate and severe calcification of aorta. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14503 | 1220648-2024-23309 | 20701376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701376&pc=OZD | 10/25/2024 | 12/27/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  | 07/31/2025 |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441357 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella cp implant and it was aborted due to anatomy. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23309 in accordance with updated procedures.
14504 | 1220648-2024-23310 | 20701384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701384&pc=OZD | 10/25/2024 | 12/26/2024 |  | 11/16/2024 |  |  | 10/25/2024 | 11/16/2024 |  |  |  |  | 11/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient had runs of ventricular tachycardia (vt). amio was started. there was also oozing noted at the groin. quick clot was applied, perclose strings were tightened, hub sutured and hooked up to pressure bag. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23310 in accordance with updated procedures.
14505 | 1220648-2024-23193 | 20693233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693233&pc=OZD | 10/23/2024 | 12/09/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210879 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient expired while on support. per medical safety officer review use of the automated impella controller cannot conclusively be associated with the outcome. | The complainant reported that the automated impella controller (aic) did not recognize or detect the purge disc. the aic was replaced. no patient harm has been reported. |  |  |  |  | The investigation into automated impella controller (aic) - purge disc/flag issues has been completed. after reviewing the logs reported issue that purge disc couldn't be detected was confirmed. during the visual inspection the purge disc flag was found to be broken. the root cause of the issue was a broken purge disc flag. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-23193. b5: included patient outcome and mso statement. d2: common device name. g1: reporting contact email. | The console has been returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14506 | 1220648-2024-23041 | 20693446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693446&pc=OZD | 04/22/2021 | 03/25/2025 |  | 11/15/2024 |  |  | 04/22/2021 | 11/15/2024 |  |  | 09/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021019214 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced limb ischemia. it was reported that at case conclusion the decision was made to remove impella for patient recovery and due to concerns with limb ischemia. the patient was successfully weaned and the impella 2.5 was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14507 | 1220648-2024-23042 | 20693447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693447&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/15/2024 |  |  | 10/21/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516554 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella 5.5 pump for the support of a cardiogenic shock patient admitted into the hospital without known history. it was reported that while on support, the patient had a small brain bleed and heparin was removed. it was further reported that the patient expired while on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1: mdr reporting contact email.
14508 | 1220648-2024-23194 | 20693450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693450&pc=OZD | 06/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/05/2021 | 11/15/2024 |  |  | 08/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1453869 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was placed without issue. it was stated that upon insertion, suction was observed. it was reported that the patient was sent to the intensive care unit and expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14509 | 1220648-2024-23195 | 20693472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693472&pc=OZD | 06/23/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/23/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022040079 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement, a computed tomography (ct) scan was performed. it was observed that the patient experienced a hemorrhagic stroke. to manage the complication, heparin was removed from the purge system. the device was explanted, care was withdrawn, and the procedural outcome was the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14510 | 1220648-2024-23196 | 20693505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693505&pc=OZD | 11/22/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 11/22/2021 | 11/15/2024 |  |  | 06/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022105049 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the right leg of the patient became more mottled with no pulses in the lower right extremity and urine output was decreasing with lactate rising. it was reported that the intensivist was at bedside. the clinical staff started slowly weaning levo and vasopressin with no change in pressure on placement signal. it was reported that the after 17.33 hours of support that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14511 | 1220648-2024-23197 | 20693548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693548&pc=OZD | 11/28/2022 | 10/13/2025 |  | 11/15/2024 |  |  | 11/28/2022 | 11/15/2024 |  |  | 11/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232186 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. after support, the device was weaned and explanted. the sheath was removed, and manta was deployed but failed, and manual pressure was needed to achieve hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14512 | 1220648-2024-23198 | 20693559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693559&pc=OZD | 08/30/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 08/30/2022 | 11/15/2024 |  |  | 08/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1502241 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the pump was successfully weaned and removed. dry closure technique was used from contralateral access and 7.0 balloon. one perclose failed, so angioseal was placed to achieve hemostasis. the patient remained stable | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14513 | 1220648-2024-23199 | 20693586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693586&pc=OZD | 10/25/2024 | 12/27/2024 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556819 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella had a pump stop. several attempts were made for the pump to restart but were unsuccessful. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the reported pump stop has been completed. pump was not returned for evaluation. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14514 | 1220648-2024-23200 | 20693626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693626&pc=OZD | 12/16/2021 | 10/13/2025 |  | 11/15/2024 |  |  | 12/16/2021 | 11/15/2024 |  |  | 06/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130980 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. while on support the patient had heparin induced thrombocytopenia with heparin being removed. additional information was provided that noted the patient expired. | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was reported that the patient had heparin induced thrombocytopenia with heparin being removed. it was reported that after 238.50 hours of care the patient expired on support. the investigation results shows that the event could not be determined as insufficient clinical details were provided. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: corrected mfg. email address.
14515 | 1220648-2024-23202 | 20693672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693672&pc=OZD | 02/17/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 02/17/2022 | 11/15/2024 |  |  | 06/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022148685 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. during impella placement, the patient's leg became ischemic when the impella sheath was inserted. an external percutaneous femoral bypass was performed to manage the complication. additional information was provided that noted the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14516 | 1220648-2024-23201 | 20693683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693683&pc=OZD | 09/19/2024 | 11/15/2024 |  | 11/15/2024 |  |  | 09/19/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563292 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported bleed involving an impella cp pump. left femoral artery (lfa) was accessed using best practices. during patient implantation of the pump, the medical doctor had trouble advancing, due to calcium, plus tortuous anatomy. long and short sheath were used and then the physician used a fast catheter to pre dilate. they were able to insert the 14 french short sheath in. the patients hemoglobin dropped to 6.1, along with blood pressure¿s in the 50¿s. a bleed was suspected. four units of blood were rapidly transfused. the impella was removed and balloon tamponade was applied to the lfa. hemostasis was achieved through that along with a perclose and manual pressure. the procedural outcome was the patient survived surgery. the patient survived the explant. | [5 empty] | The investigation has been completed. no product was returned. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details provided. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
14517 | 1220648-2024-23203 | 20693697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693697&pc=OZD | 04/01/2022 | 03/07/2025 |  | 11/15/2024 |  |  | 04/01/2022 | 11/15/2024 |  |  | 10/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023160520 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. the complainant reported that during impella placement, the patient's leg was observed to be mottled, and the suction was at pump speed level 8. the device was explanted. additional information was provided that noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14518 | 1220648-2024-23204 | 20693717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693717&pc=OZD | 06/22/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 06/22/2022 | 11/15/2024 |  |  | 06/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022099497 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 2.5 device for mechanical circulatory support. the complainant reported that during impella placement, the pump speed level was reduced to zero as the device was noted to be in the aorta. no further information was provided. additional information was provided that noted the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14519 | 1220648-2024-23205 | 20693740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693740&pc=OZD | 06/24/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 06/24/2022 | 11/15/2024 |  |  | 10/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173355 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support it was noted that pump was unable to run above p-2/p-3 (flows ~1.3l/min max) without suction. additional information was provided that noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14520 | 1220648-2024-23206 | 20693760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693760&pc=OZD | 06/28/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 06/28/2022 | 11/15/2024 |  |  | 11/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188623 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. while on support the patient was taken to interventional radiology for a suspected abdominal bleed. plans were to start heparine once patient returned back to the floor. additional information was provided that noted the family withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14521 | 1220648-2024-23208 | 20693767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693767&pc=OZD | 08/03/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 08/03/2022 | 11/15/2024 |  |  | 06/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139404 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella 2.5 device for mechanical circulatory support due to an aortic valvuloplasty. while on support the patient underwent 10 minutes of cardiopulmonary resuscitation. after multiple bradycardic events the patient progressed to asystole. additional information was provided that noted the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14522 | 1220648-2024-23207 | 20693769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693769&pc=OZD | 10/24/2024 | 12/16/2024 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133102 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support when the automated impella controller (aic) was showing a lot of ectopy combined with intermittent normal waveforms and metrics. the aic was also connecting and disconnecting from wifi repeatedly. the aic was replaced and the metrics changed and looked better. support continued. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the optical issues has been completed. data logs were reviewed finding there are two failure modes alleged in this complaint. the intermittent abnormal placement signal was found in the logs as there were multiple instances of placement signal not reliable during the case. the real time (rt) logs showed that placement signal oscillated between the correct placement signal and -3276.8mmhg. additionally, there was more persistent barcoding in the rtlogs during the placement signal not reliable alarms. additionally, during the case there was an unexpected reboot by the aic with no indication in the logs for the cause. the console was returned for evaluation. upon functional inspection, the placement signal failure mode was confirmed. the optical bench was found to have a very low snr of 728 and the ccd array was distorted which would have contributed to the unreliable placement signal. the unexpected reboot was not reproduced during testing and it is unknown what caused it. the cause of the placement signal issues was a defective optical bench with very low snr and a distorted ccd array. the cause of the unexpected reboot could not be determined as it was not reproduced and there were no indicators in the logs. the cause of the signs of usd card corruption could not be determined from the logs but could have been caused by the frequent rlm reboots due to low rssi. corrections to the initial mfr report: g1 mdr reporting contact email updated h5 changed to no h8 changed to reuse
14523 | 1220648-2024-23209 | 20693851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693851&pc=OZD | 10/25/2024 | 12/27/2024 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490018 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the right lower extremity felt cold and pale. the sheath was inserted in reverse and the blood flow was returned, and the color improved, and the cold feeling remained slightly compared to the left. reflux by reverse insertion sheath resolved ischemia. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14524 | 1220648-2024-23210 | 20693975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20693975&pc=OZD | 10/25/2024 | 12/23/2024 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764459 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) was having a preventative maintenance and it was found that the screen was flickering. this issue did not occur during patient support. | [5 empty] | The investigation for the console (aic) display issues has been completed. data logs were reviewed but found there was no explicit evidence in the logs to confirm the reported failure mode - flickering screen. the aic was returned for evaluation. during boot-up, the console started flickering, and it occurred intermittently while running the pump. the flickering pattern was horizontal. no damage was found on the console hardware related to the reported failure mode. the wires on the 4-in-1 connectors x26 and x51 were appropriately connected. the root cause of the intermittent flickering screen was not determined. the horizontal flickering could most likely be caused by display malfunction and/or 4-in-1 malfunction. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
14525 | 1220648-2024-23211 | 20694020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20694020&pc=OZD | 01/31/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 01/31/2022 | 11/15/2024 |  |  | 01/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1464493 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. during impella placement, the pump could not be delivered to the femoral artery due to stenosis. the device was removed, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14526 | 1220648-2024-23212 | 20694212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20694212&pc=OZD | 10/24/2024 | 01/03/2025 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558210 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp pump placed to support a patient with acute myocardial infarction / cardiogenic shock. the pump was placed but there was an access site bleed that prompted sheath replacement and sutures. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by replacing the sheath and redoing the forward suture.
14527 | 1220648-2024-23213 | 20694250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20694250&pc=OZD | 10/25/2024 | 12/12/2024 |  | 11/15/2024 |  |  | 10/25/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547760 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. upon insertion of pci sheath slight ooze/bleed was noted at hemostatic valve. an attempt was made to reposition sheath with slight control although bleeding continued throughout case. slight ooze continued at impella access site and angioseal was placed as well as manual pressure for 20 minutes. femstop also placed on patient as precaution to prevent further bleeding. patient hemodynamically stable at end of case. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since non-abiomed product was used and it is not recommended per ifu. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
14528 | 1220648-2024-23214 | 20694291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20694291&pc=OZD | 06/29/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/29/2021 | 11/15/2024 |  |  | 04/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022077501 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement, the patient experienced septic shock. it was reported that, as escalation of therapy was not an option, care was withdrawn. the procedural outcome was the patient expired an hour after explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14529 | 1220648-2024-23215 | 20695340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695340&pc=OZD | 10/12/2024 | 01/06/2025 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533061 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp smart assist support, the 75-year-old male patient had a cerebral hemorrhage. the patient expired due to the cerebral hemorrhage. the impella functioned normally. the cause of death was determined to be multiple organ failure caused by cerebral hemorrhage. the relationship between the impella and the cerebral hemorrhage and cause of death is unknown. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-23215 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23215 was submitted in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14530 | 1220648-2024-23216 | 20695342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695342&pc=OZD | 10/29/2024 | 12/30/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534164 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump placed to support 60-year-olld male patient who was in cardiogenic shock (cgs) and acute myocardial infarction (ami). the pump was supporting as care escalation from an intra-aortic balloon pump (iabp). the pump was seen to have sterile sleeve damage on day 13 of the support. the team cleaned it and wrapped in tegaderm. on day 14 the patient was seen to be bleeding and so the team reduced the ooze by use of surgicel. patient still awaiting transplant. pump remains on for support at this time. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of access site bleeding and product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23216. a5 ethnicity. d6b missing. e4 should have been left blank. g1 reporting contact email revised in accordance with updated procedures.
14531 | 1220648-2024-23217 | 20695543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695543&pc=OZD | 04/03/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/03/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022030189 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, oozing was observed at the right auxiliary pocket from recent eliquis. site oozing had stopped due to dressing change. customer was advised that some heparin can be added to the purge since bleeding had stopped. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14532 | 1220648-2024-23218 | 20695553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695553&pc=OZD | 04/08/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/08/2021 | 11/15/2024 |  |  | 01/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022048237 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during support, care was stopped as the patient had ischemic cerebral vascular accident (cva) due to carotid dissection. however, it is unknown if the dissection was there prior to impella, or if impella contributed to the vascular damage. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14533 | 1220648-2024-23219 | 20695568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695568&pc=OZD | 04/06/2021 | 03/17/2025 |  | 11/15/2024 |  |  | 04/06/2021 | 11/15/2024 |  |  | 08/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503737 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, a low purge pressure alarm was observed that could not be resolved. the pump speed level decreased due to suction alarms. the medical staff was advised to zero the sensor to confirm the suction alarm was authentic, and to then consider reloading the pump. weaning was successful, the device was explanted, and the patient survived the procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14534 | 1220648-2024-23220 | 20695611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695611&pc=OZD | 04/11/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/11/2021 | 11/15/2024 |  |  | 08/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1447360 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with off-pump coronary artery bypass (opcab) was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, heparin was stopped due to bleeding at the extracorporeal membrane oxygenation (ecmo) insertion sites. weaning was successful, the device was explanted, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14535 | 1220648-2024-23221 | 20695677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695677&pc=PYX | 10/23/2024 | 12/19/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  | 07/31/2024 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551294 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and during support, there was a sudden drop in flows with an associated motor spike. the impella was increased from p6 to p7 in order to maintain flows of 1.5 liters per minute or greater. transesophageal echocardiogram confirmed the presence of a small clot at the outlet causing a partial occlusion of the cannula. the pump was successfully weaned. the patient remained stable and survived the explant. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Impella rp flex returned for evaluation. upon visual inspection, biomaterial was present on the impeller blade. data logs were reviewed and lt logs at time of sudden drop in flows shows a motor current spike with a step up in motor current and placement signal. purge flow drops slightly with purge pressure becoming slightly more pulsatile. clinical details noted that patient experienced sudden drop in flows and a motor current spike was seen. the root cause of the low pump flow was due to biomaterial ingestion. the failure mode of "thrombus" was investigated under "low pump flow" as the root cause of the low pump flow was due to biomaterial. g1 reporting contact email revised on manufacturer device report 1220648-2024-23221 in accordance with updated procedures.
14536 | 1220648-2024-23222 | 20695762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695762&pc=PYX | 10/22/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/22/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529768 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2024 |  |  |  |  | The complainant reported that an impella rp pump experienced a pump stop coinciding with a high motor current alarm during patient support. the pump was successfully restarted after three failed attempts. however, support levels had to remain at or below p-2 or another stoppage would occur. the pump was explanted and it was discovered that fibrous biomaterial was within the inlet of the pump. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. | The investigation into temporary pump stop and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23222 b1 should have been both product problem and adverse event. d4 model number. g1 reporting contact email revised from the initial submission in accordance with updated procedures. g2 report source health professional missing. g6 type of report missing.
14537 | 1220648-2024-23223 | 20695793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695793&pc=OZD | 04/08/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/08/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022030186 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. insertion of the impella 5.0 was attempted, however the pump could not pass the innominate artery of the aorta. implant process was aborted and an impella cp was placed to manage support. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14538 | 1220648-2024-23224 | 20695921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695921&pc=OZD | 04/15/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/15/2021 | 11/15/2024 |  |  | 04/30/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1430017 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient was diaphoretic and developed tachycardia. foley was discontinued and care team increased mili to .3 and adjusted the p levels. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14539 | 1220648-2024-23225 | 20695956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695956&pc=OZD | 04/23/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/23/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037271 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, patient had bleeding issues related to tracheostomy. there was no heparin in the purge solution and d5 was being run due to bleeding issues. patient received blood products and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14540 | 1220648-2024-23226 | 20695973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695973&pc=OZD | 10/23/2024 | 12/13/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1346474 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller (aic) fail with a red controller alarm at boot up. the aic was removed and replaced. this occurred during preparation. therefore, no patient was involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into automated impella controller (aic) - controller failure (red alarm) has been completed. during the device analysis the controller was booted up, instantly getting controller failure alarm for defective purge pressure sensor that resolved and recurred. the console was immediately shut down by the user. the console was returned for investigation but, despite power cycling 200 times, the failure was unable to be reproduced both with or without purge cassette. the voltage across the purge pressure sensor was found to be marginally passing which is why the alarm was triggering and resolving in the field just the one time. the cause of the controller failure was most likely a marginally passing pud. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23226.
14541 | 1220648-2024-23227 | 20695983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20695983&pc=OZD | 04/18/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/18/2021 | 11/15/2024 |  |  | 06/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1489284 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient was stable but continued to ooze from the access site. 3 units of packed red blood cells were given to the patient. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14542 | 1220648-2024-23228 | 20696028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696028&pc=OZD | 05/04/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/04/2021 | 11/15/2024 |  |  | 09/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504534 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Life Threatening | Required Intervention | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, impella access site was infected. infection/sepsis was treated with vacuum- assisted closure (vac). patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14543 | 1220648-2024-23229 | 20696144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696144&pc=OZD | 04/28/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/28/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022033041 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient had 17 minutes run of ventricular tachycardia (vt). one (1) electric shock was received, and the patient was placed on amiodarone and magnesium drips. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14544 | 1220648-2024-23230 | 20696162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696162&pc=OZD | 04/22/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/22/2021 | 11/15/2024 |  |  | 04/30/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1426710 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in pre-op placement was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient started bleeding from the axillary site. manual pressure was held, and patient was taken to surgery due to impella sheath site bleeding. pocket was cauterized to stop bleeding. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14545 | 1220648-2024-23231 | 20696183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696183&pc=OZD | 04/30/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 04/30/2021 | 11/15/2024 |  |  | 04/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1476519 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 5.0 device for mechanical circulatory support. during insertion attempt, the patient had very torturous anatomy with mild to moderate calcification present. the physician was unable to advance the 5.0 pump over aortic arch. the impella 5.0 was removed and replaced with an impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14546 | 1220648-2024-23232 | 20696243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696243&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547788 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 sterile sleeve disconnected from the distal locking end. the staff cleaned the area and resecured using tape and tegaderm. additionally, a hematoma with some oozing was noted at the impella access site. sandbag and systemic bivalirudin was on hold. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into access site bleeding and product damage (sterile sleeve damage) has been completed. data logs are not relevant to the issues and the product was not returned for investigation. clinical details provided limited information regarding the issue and do not suggest what could have caused the bleed 12 days after the start of support. since the product was not returned and sufficient clinical details were not provided, the root cause of the access site bleed could not be determined. clinical details suggested that the sterile sleeve detached from the hub quite easily. however, as the product was not returned for investigation, the root cause of the product damage could not be determined. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23232 was submitted.
14547 | 1220648-2024-23233 | 20696324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696324&pc=OZD | 10/29/2024 | 12/24/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560149 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the 74-year-old female patient developed ischemia in their lower leg during impella cp support. the patient was scheduled for an external bypass to return flow to the leg. support with the implanted pump remained ongoing. the pump was successfully explanted after two days of support. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14548 | 1220648-2024-23234 | 20696340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696340&pc=OZD | 10/29/2024 | 12/19/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523182 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support, the 59-year-old male patient had hemolysis and renal failure. the patient started on dialysis due to the renal failure. the impella cp was weaned and another impella device was used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis and the renal failure has been completed. no product was returned. log review showed positioning alarms within first two hours of case. clinical details note that the position was checked with echo and the pump was repositioned after the patient was transferred to icu. no motor current spikes or suction were present in the logs prior to the report of hemolysis/renal failure. poor lv function was also mentioned in clinical details at this point, and the patient struggled to wean off cp later that day. the root cause of the hemolysis was most likely patient condition related. the root cause of the renal failure was not determined since insufficient clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14549 | 1220648-2024-23235 | 20696568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20696568&pc=OZD | 10/23/2024 | 12/12/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530860 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a dressing change, the patient began to bleed uncontrollably from the impella cp access site. manual pressure was held for three hours but the condition remained. the patient was later taken down to the catheter lab where the decision was made to remove the pump. the bleeding had stopped at the time of explant. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14550 | 1220648-2024-23236 | 20697164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20697164&pc=OZD | 10/29/2024 | 12/31/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505127 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support, the 58-year-old male patient developed a hematoma at insertion site prompting the team to stop anticoagulation. the pump was successfully explanted after 11 days of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation for the reported access site bleed has been completed. the pump was investigated and catheter was damaged near the cannula. the catheter damage does not affect the failure mode investigation. however, clinical details were not sufficient. therefore, the root cause of the access site bleed could not be determined. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email.
14551 | 1220648-2024-23238 | 20697444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20697444&pc=OZD | 05/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/05/2021 | 11/15/2024 |  |  | 01/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022035380 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited outflow blocked alarms. an echocardiogram was performed and revealed that the impella was far out of the left ventricle. the device was repositioned, and it was noted that the patient was stable and survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14552 | 1220648-2024-23237 | 20697478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20697478&pc=OZD | 10/29/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 58-year-old female patient became combative and dislodged the pump from their left ventricle. the pump was subsequently removed following the displacement. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement. corrections to the initial mfr report: g1 mdr reporting contact email.
14553 | 1220648-2024-23239 | 20697686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20697686&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515170 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support, there was flattened motor current and impella flows. echocardiogram was done and the position was verified. the impella was confirmed to be in the mid ventricle. several attempts were made to reposition. however, the issue remained. blood products were given to the patient. the pump was exchanged and biomaterial was observed. the patient remained stable and the patient survived the explant. | [5 empty] | The investigation into low pump flow and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23239 | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿.
14554 | 1220648-2024-23240 | 20697770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20697770&pc=OZD | 05/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/05/2021 | 11/15/2024 |  |  | 02/28/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1467661 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, a high purge pressure (pp) alarm occurred. purge had not been trending upward; this was a sudden development. they ran heparin, checked for kinks in the catheter and none were found. the cassette was replaced. later, high pp alarms still occurred. running d5/12.5. it had been trending down. they tried the decel method and issue resolved. ran systemic heparin also. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14555 | 1220648-2024-23241 | 20698027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698027&pc=OZD | 05/03/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/03/2021 | 11/15/2024 |  |  | 05/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1434505 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported an inability to advance 5.0 past anastomosis and insertion was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14556 | 1220648-2024-23243 | 20698103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698103&pc=OZD | 10/23/2024 | 12/10/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1246566 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) arrived at the customer site damaged following annual preventative maintenance. it was noted that the box the aic arrived in was undamaged. there was no patient involvement. | [5 empty] | The investigation for the console (aic) damage has been completed. data logs were not informative for this investigation. console was returned for evaluation and was able to perform as expected but there were two distinctive cracks in the top right. the cause of the cracked housing could not be determined. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14557 | 1220648-2024-23242 | 20698193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698193&pc=OZD | 05/06/2021 | 05/28/2025 |  | 11/15/2024 |  |  | 05/06/2021 | 11/15/2024 |  |  | 08/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1503737 |  | 005062 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient developed hematuria. it was stated that the hematuria was potentially from a bleed as blood remained in the patient¿s fecal matter. weaning was successful, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14558 | 1220648-2024-23247 | 20698361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698361&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511302 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during insertion under transesophageal echocardiogram (tee), the impella 5.5 was attempted to be placed wireless. perforation of left ventricle wall happened, and it was patched. the impella 5.5 was attempted to be place again under tee and the patch did not hold. the patch was again repaired, and the decision was made to proceed and the patient was weaned off of the bypass without the impella. the impella 5.5 case was aborted and the patient remained stable. | [5 empty] | The investigation for the ventricle perforation has been completed. the product was not returned for investigation. data logs were reviewed and there were no positioning alarms in the available logs. clinical details reported that the pump was inserted without a guidewire and that the left ventricle was perforated during insertion. based on clinical details, the root cause of the ventricle perforation was determined to most likely be improper technique inserting without a guidewire. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23247 was submitted in accorance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
14559 | 1220648-2024-23246 | 20698365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698365&pc=OZD | 05/12/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/12/2021 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022044595 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, there was a low purge pressure shortly after starting the pump, there was no leaks from purge reservoir/filter, yellow luer lock was tightened, a new purge cassette was changed, the aic (automated impella controller) console was changed, and the issue resolved when purge was changed from d5% to d20%. additionally, the patient began to bleed profusely from all lines and was brought to the surgery suite to suture up lines. bicarb replacement for heparin due to bleeding. the decision was made to leave impella in for a few more days. | [5 empty] | The investigation for the reported low purge pressure issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause, which was determined to be related to the use of purge fluid with low dextrose concentration. this report is being filed as part of a retrospective review of historical records.
14560 | 1220648-2024-23248 | 20698409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698409&pc=OZD | 05/16/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/16/2021 | 11/15/2024 |  |  | 09/30/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1453344 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for hemodynamic support. while on support, was bleeding from ¿everywhere¿ and switched purge to bicarb. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14561 | 1220648-2024-23249 | 20698426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698426&pc=OZD | 10/23/2024 | 11/30/2025 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1710720 |  | 1000432 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during removal there was red notifications of impella stopped controller failure and purge system open occurred on the automated impella controller. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14562 | 1220648-2024-23244 | 20698445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698445&pc=OZD | 05/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/05/2021 | 11/15/2024 |  |  | 09/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1504585 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, an unspecified bleed was noted, and heparin was held until the bleed diminished. fresh frozen plasma and packed red blood cells were provided to the patient overnight to manage the complication and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14563 | 1220648-2024-23251 | 20698512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698512&pc=OZD | 10/29/2024 | 08/14/2025 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1752855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller error alarm. the aic was removed. there was no patient involvement. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.1 added information as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-23251. a.2 age should have been left blank on the initial manufacturer device report number 1220648-2024-23251 as no patient was involved. a.3 revised as selection was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. d.1 brand name and d.2 common device name were revised as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. d.4 model number and serial number were revised as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. e.2 and e.3 revised as selections were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-23251 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 revised mdr reporting contact email as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. h.6 codes 3221 and 2199 were entered incorrectly on the initial manufacturer device report number 1220648-2024-23251. code 2645 was added.
14564 | 1220648-2024-23253 | 20698538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698538&pc=OZD | 05/18/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/18/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022050926 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was going to be on an impella 5.0 for mechanical circulatory support. the physician was unable to advance the impella into the left ventricle due to the patient's vasculature. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14565 | 1220648-2024-23255 | 20698543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698543&pc=PYX | 10/29/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025560979 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella rp flex was placed to support a 58-year-old female patient with multiple cardiac and systemic comorbidities. the patient became bradycardic and proceeded to pulseless electrical activity (pea) arrest, prior to the placement of the rp. the pump was placed but had low flows and high purge pressure/purge blocked. the family made choice to withdraw care and patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smartassist system: section: using the automated impella controller with the impella rp with smartassist system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into low pump flow and high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow and high purge pressure is most likely due to biomaterial g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23255.
14566 | 1220648-2024-23257 | 20698564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698564&pc=OZD | 05/23/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/23/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022060110 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy needing an impella 5.0 device for mechanical circulatory support. while on support the purge was changed from d5 with heparin to sodium bicarbonate due to access site bleed. in addition, hemolysis was noted. the device could not be repositioned due to the inability to free the pigtail. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14567 | 1220648-2024-23245 | 20698568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698568&pc=OZD | 05/12/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/12/2021 | 11/15/2024 |  |  | 07/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the cardiac surgeon instructed to stop the impella due to the risk of the chordae tendineae being sucked in by the impella during dislocation during coronary artery bypass graft. the patent was weaned and the device was explanted successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14568 | 1220648-2024-23256 | 20698572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698572&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/15/2024 |  |  | 10/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558895 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was placed for low ejection fraction bentall. on day two of support, the impella got alarms for air in the purge that eventually turned into high purge pressure system blocked alarm. pump flow was still stable. the bag and cassette were changed out twice with no resolution. tissue plasminogen activator was administered through the purge with a new cassette the following day. all evening and into the next morning, the purge pressure remained high. the pump saturation was now evident on waveforms. the impella 5.5 was weaned to p2 and explanted. the patient remained stable and the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure and saturated flow has been completed. no product was returned. log review showed no issues or trends in purge pressure or flow up when "purge system open" alarms occurred followed by "air in purge system" and "purge system blocked" alarms. the pressure had dropped toward zero at this point. about thirty minutes after the first alarm, the pressure came back but immediately increased above 1000 mmhg. after multiple bag/cassette changes, the purge pressure did not resolve. the motor current and pump flow began increasing later that day. the pump was explanted after flow saturation occurred. the purge system alarms did not follow any trend or events in purge pressure, and the purge pressure had dropped to zero at this point. the pump cannot function properly without adequate purge flow. the root cause of the high purge pressure was most likely use-related. the saturated flow occurred later in the day after the high purge pressure began. it is likely that the inadequate purge flow allowed for biomaterial to buildup around the purge gap and lead to saturated flow. the root cause of the saturated flow was most likely biomaterial. added h6 impact code 4644 corrections to the initial mfr report: g1 mdr reporting contact email.
14569 | 1220648-2024-23254 | 20698597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698597&pc=OZD | 05/19/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/19/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063204 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support heparin was withheld due to drainage from the jackson-pratt drain. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14570 | 1220648-2024-23252 | 20698601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698601&pc=OZD | 10/29/2024 | 12/26/2024 |  | 11/15/2024 |  |  | 10/29/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482901 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 49-year-old male patient was attempted to be implanted with impella 5.5. attempts were made for the right axillary 5.5 insertion and they were unable to advance the pump due to patient anatomy. there was no harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since insertion was aborted due to patient's anatomy. corrections to the initial mfr report: g1 mdr reporting contact email.
14571 | 1220648-2024-23250 | 20698727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698727&pc=OZD | 05/17/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/17/2021 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022042557 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Renal Failure | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 5.0 for mechanical circulatory support. it was reported that the device was difficult to position due to the patient's reconstructed aorta and tilted heart. additionally, during support the patient experienced kidney and renal failure. the resolution is unknown. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14572 | 1220648-2024-23258 | 20698730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698730&pc=OZD | 06/03/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/03/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022064680 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Unspecified Infection | Renal Failure | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device became mal positioned during a walk, the resolution for this issue is unknown. additionally, the patient developed increased white count and was treated for infection. the patient also experienced renal failure and was treated with continuous renal replacement therapy. also, the patient had an esophageal tear, this was treated by clipping the tear and the bleeding was resolved. an upper/lower gastrointestinal bleed was reported with no resolution specified. it was reported that the patient survived normal explant and the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14573 | 1220648-2024-23263 | 20698786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698786&pc=OZD | 10/13/2021 | 11/23/2024 |  | 11/15/2024 |  |  | 10/13/2021 | 11/15/2024 |  |  | 04/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022083145 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an impella 5.0 for mechanical circulatory support. during support, the purge flows dropped to 1.5ml/hr and purge pressure increased to over 1000. the decision was made to change the purge solution to bicarb and the start/stop technique was used successfully. it was reported that the device was normally weaned, and the procedure was successful. | It was additionally reported that there were positional alarms during the support. no further information was provided. it was reported that the device was normally weaned, and the procedure was successful. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14574 | 1220648-2024-23262 | 20698803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698803&pc=OZD | 06/03/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/03/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022060184 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient was bleeding from an unknown location and bicarb was added to the purge. heparin was discontinued. it was reported that the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14575 | 1220648-2024-23261 | 20698833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698833&pc=OZD | 06/07/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/07/2021 | 11/15/2024 |  |  | 11/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022033040 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | American Indian Or Alaskan Native, White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the device exhibited ongoing high-pressure alarms. the staff administered tissue plasminogen activator in the purge to control the purge flow. the purge flow was normalized. additionally, a thrombus was identified on the pigtail. it was reported that the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14576 | 1220648-2024-23264 | 20698856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698856&pc=OZD | 10/13/2021 | 03/17/2025 |  | 11/15/2024 |  |  | 10/13/2021 | 11/15/2024 |  |  | 04/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022075955 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support a small hematoma was discovered. the resolution for this issue is unknown, however the patient survived the procedure. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14577 | 1220648-2024-23259 | 20698887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698887&pc=OZD | 05/31/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/31/2021 | 11/15/2024 |  |  | 01/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0-EU |  |  | OZD |  |  | IMPELLA 5.0-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022043287 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 5.0 for mechanical circulatory support. it was reported that the device moved into the aorta, repositioning attempts were unsuccessful. the physician decided to explant the device. it was reported that the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14578 | 1220648-2024-23260 | 20698890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20698890&pc=OZD | 05/28/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 05/28/2021 | 11/15/2024 |  |  | 10/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024081 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 5.0 for mechanical circulatory support. it was reported that the implant was unsuccessful due to the patient's narrowed artery distal to anastomoses. the device could not pass through the patient vascular. it was noted that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14579 | 1220648-2024-23265 | 20699046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699046&pc=OZD | 12/09/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 12/09/2021 | 11/15/2024 |  |  | 07/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022129474 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in pre-op placement was implanted with an impella 5.0 device for mechanical circulatory support. impella 5.0 was inserted in the right axillary artery was successfully implanted without issues. during support, the patient had vascular damage. as a result, systemic heparin was turned off for 24hrs. impella was subsequently successfully weaned and explanted. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14580 | 1220648-2024-23266 | 20699122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699122&pc=OZD | 09/27/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 09/27/2022 | 11/15/2024 |  |  | 02/28/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022092427 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. during impella placement, a blockage in the purge line was suspected. weaning was successful, the device was removed, and the patient survived the explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: this report is being filed as part of a retrospective review of historical records.
14581 | 1220648-2024-23268 | 20699170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699170&pc=OZD | 01/05/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 01/05/2022 | 11/15/2024 |  |  | 04/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022077203 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient had a stroke. no additional information was provided related to the resolution. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14582 | 1220648-2024-23269 | 20699192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699192&pc=OZD | 01/04/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 01/04/2022 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022028721 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, the patient had an access site bleed at the sheath site that required a suture to control the bleeding. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14583 | 1220648-2024-23270 | 20699209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699209&pc=OZD | 01/06/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 01/06/2022 | 11/15/2024 |  |  | 11/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022148715 |  | 005066-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support, patient developed thrombocytopenia. no additional information was provided. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14584 | 1220648-2024-23271 | 20699393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20699393&pc=OZD | 01/06/2022 | 11/15/2024 |  | 11/15/2024 |  |  | 01/06/2022 | 11/15/2024 |  |  | 06/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021015001 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.0 device for mechanical circulatory support. complainant reported that after returning to the intensive critical unit (icu) after surgery, the tip of the catheter came out of the left ventricle (lv). when the catheter position was adjusted it could not be reinserted, therefore, it was removed. the patient also had a meandering lower limb, and the catheter came out when it was adjusted. no reinsertion was performed. patient's condition was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14585 | 1220648-2024-23272 | 20700412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700412&pc=PYX | 11/13/2024 | 06/27/2025 |  | 11/15/2024 |  |  | 11/13/2024 | 11/15/2024 |  |  | 11/30/2025 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024425626 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemothorax | Pneumothorax | Great Vessel Perforation | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that heavy pressure had to be applied when going in with 23 french sheath. the physician had difficulty passing and a skin nick done to assist in sheath insertion however the sheath eventually passed through. the patient began to decompensate requiring high doses of inotropes, vasopressors and eventually requiring cardiopulmonary resuscitation (cpr). a right chest tube was placed when a pneumothorax was noted on a chest x-ray. a chest tube was placed and diagnosed hemothorax. cpr continued and multiple units of red blood cells for suspected superior vena cava perforation from dilator insertion. it was further reported that the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of access site bleeding and ventricle perforation has been completed. access site bleeding- the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. ventricular perforation- the data logs show that there were suction alarms but no positioning alarms were seen. due to lack of sufficient clinical details provided and no product returned, the root cause of the ventricle perforation issue could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23272 in accordance with updated procedures.
14586 | 1220648-2024-23273 | 20700760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700760&pc=OZD | 06/16/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/16/2021 | 11/15/2024 |  |  | 09/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2021013237 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced alot of bleeding in the chest tubes and received 7 units of packed red blood cells. the patient underwent exploratory surgery for active bleeding. the left internal mammary artery was found to be bleeding at anastomoses and was repaired. it was reported that the device was normally explanted, and the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14587 | 1220648-2024-23274 | 20700765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700765&pc=OZD | 06/09/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/09/2021 | 11/15/2024 |  |  | 10/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022022713 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an impella 5.0 for mechanical circulatory support. the device was implanted, and the physician was trying to reposition the device, the device was pulled out and was it was difficult to recross the valve. the new valve has a cage around it and decision was made after several attempts to not recross and risk damaging valve. the device was explanted, and support was continued with an intra-aortic balloon pump. it was reported that the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14588 | 1220648-2024-23275 | 20700774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700774&pc=OZD | 06/09/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/09/2021 | 11/15/2024 |  |  | 03/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022063500 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was going to be on an impella 5.0 for mechanical circulatory support. it was reported that the device was not able to be implanted due to the patient's vasculature. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14589 | 1220648-2024-23276 | 20700779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700779&pc=OZD | 06/14/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/14/2021 | 11/15/2024 |  |  | 06/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1488546 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella placement, bleeding was noted at the insertion site. it was reported that heparin was held to manage the complication. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14590 | 1220648-2024-23277 | 20700846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700846&pc=OZD | 06/16/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/16/2021 | 11/15/2024 |  |  | 10/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1457632 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella placement, the patient was returned to the operating room with bleeding from the chest tube. it was reported that various products were provided to the patient. heparin was removed from the purge to manage the complication. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14591 | 1220648-2024-23278 | 20700847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700847&pc=OZD | 06/16/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/16/2021 | 11/15/2024 |  |  | 01/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022042820 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Injury | Yes | No | Deformation Due to Compressive Stress | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during an attempt at impella implant, insertion was difficult due to the placement of the guidewire and a kinked cannula. it was reported that the patient experienced an arrhythmia and insertion of a pacemaker was discussed. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14592 | 1220648-2024-23279 | 20700848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700848&pc=OZD | 06/21/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/21/2021 | 11/15/2024 |  |  | 11/30/2021 |  | 11/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022030206 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.0 device for mechanical circulatory support. it was reported that the patient had significant bleeding and there were concerns for apical injury. there was general ooziness cited as the source for the bleeding. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14593 | 1220648-2024-23280 | 20700850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700850&pc=OZD | 06/24/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 06/24/2021 | 11/15/2024 |  |  | 10/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1460455 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 5.0 device for mechanical circulatory support. it was reported that the impella device would not cross into the left ventricle. multiple unsuccessful attempts were reported, and the physician decided to abort the implant of the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14594 | 1220648-2024-23281 | 20700851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700851&pc=OZD | 07/01/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/01/2021 | 11/15/2024 |  |  | 01/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1464339 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient was significantly bleeding from extracorporeal membrane oxygenation (ecmo) decannulation site. blood transfusion was done to resolve the issue. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14595 | 1220648-2024-23282 | 20700872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700872&pc=OZD | 07/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/05/2021 | 11/15/2024 |  |  | 07/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1430738 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, a decrease in purge flow rate and an increase in purge pressure was noted. additionally, it was reported that the patient had a blood clot formation in the lower limbs and veins. decision was made to replace the pump. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14596 | 1220648-2024-23283 | 20700878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700878&pc=OZD | 07/02/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/02/2021 | 11/15/2024 |  |  | 04/30/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022065986 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. it was reported while on impella 5.0 support, there was a flow blockage. the impella 5.0 was successfully exchanged with a new impella 5.0 the patient was hemodynamically stable, awake, and mobile following the pump exchange. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14597 | 1220648-2024-23284 | 20700901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700901&pc=OZD | 07/05/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/05/2021 | 11/15/2024 |  |  | 02/28/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1480827 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. during support, the patient had a decrease in the purge flow rate. no additional information was provided related to the resolution. impella was successfully weaned and explanted and patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14598 | 1220648-2024-23285 | 20700909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700909&pc=OZD | 07/22/2021 | 04/16/2025 |  | 11/15/2024 |  |  | 07/22/2021 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022042548 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Vascular Dissection | Thrombocytopenia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella 5.0 device for mechanical circulatory support due to cardiomyopathy. while on support the patient experienced some blood loss requiring heparin to be discontinued. red blood cells were transfused to manage the complication. the patient¿s partial thromboplastin time (ptt) was elevated. it was reported that the purge solution was changed to sodium bicarbonate. in addition, a transvenous pacemaker was implanted due to temporary bradycardia following the impella 5.0 insertion. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14599 | 1220648-2024-23286 | 20700911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700911&pc=OZD | 07/24/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/24/2021 | 11/15/2024 |  |  | 10/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022024093 |  | 005062 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported that a hematoma was observed at the right axillary impella insertion site. manual pressure was held, and the site was redressed. it was reported that the patient¿s tissue appeared to be soft and bruising that extended from the lower right pectoral area to the upper right bicep was noted. the patient remained on impella support, was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14600 | 1220648-2024-23287 | 20700931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700931&pc=OZD | 07/30/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/30/2021 | 11/15/2024 |  |  | 10/31/2021 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1450376 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hematoma | [11 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. while on support a hematoma was observed and blood was transfused to the patient to manage the complication. in addition, due to a fever, the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14601 | 1220648-2024-23288 | 20700932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700932&pc=OZD | 07/28/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 07/28/2021 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022035363 |  | 005060 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 5.0 device for mechanical circulatory support. the complainant reported that a high purge pressure system blocked alarm was observed on the impella device. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14602 | 1220648-2024-23289 | 20700934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700934&pc=OZD | 08/05/2021 | 05/16/2025 |  | 11/15/2024 |  |  | 08/05/2021 | 11/15/2024 |  |  | 09/30/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 1504533 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. during support, the patient experienced hemolysis, and the decision was made to explant the device. no further information is available. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14603 | 1220648-2024-23290 | 20700935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700935&pc=OZD | 08/09/2021 | 11/15/2024 |  | 11/15/2024 |  |  | 08/09/2021 | 11/15/2024 |  |  | 12/31/2022 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2022037354 |  | 005062-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella 5.0 for mechanical circulatory support. it was reported during impella 5.0 support, it was noted from impella connect that the pump had an alarm for purge pressure high and the alarm was disabled. the medical team did not require assistance with resolving the issue. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14604 | 1220648-2024-23291 | 20700938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700938&pc=OZD | 08/05/2021 | 09/26/2025 |  | 11/15/2024 |  |  | 08/05/2021 | 11/15/2024 |  |  | 05/31/2023 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022086135 |  | 005060 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 5.0 for mechanical circulatory support. it was reported that despite advice against the using trans-esophageal echocardiogram (tee) exclusively, the wired insertion was performed and failed due to poor observation of the guidewire. wireless insertion proceeded; implantation was successful, however the impella was reported to be in a poor position. the procedure was successful, without report of any additional issues. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issues has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
14605 | 1220648-2024-23292 | 20700998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20700998&pc=OZD | 10/01/2024 | 09/23/2025 |  | 11/15/2024 |  |  | 10/01/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535050 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient of unknown race/ethnicity undergoing heart transplant rejection was implanted with an impella cp device for mechanical circulatory support. during support, the pump suddenly stopped after a high motor current, though it began again after several restart attempts. the patient remained on impella support, and there was no reported harm to the patient. | The complainant also reported that there were incorrect measures after the pump stop. the left ventricle signal and pump flow were not calculating correctly. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. the pump nor the data logs were returned for evaluation. clinical details were not sufficient to determine the root cause. therefore, the root cause of the pump stop could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was not determined since product was not returned and data logs were not returned as well. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-23292. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
14606 | 1220648-2024-23293 | 20701016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701016&pc=OZD | 09/26/2024 | 02/13/2025 |  | 11/15/2024 |  |  | 09/23/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510361 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | Other | Life Threatening | Death | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp placed for support after the 5.5 failed to deliver. the cp was supporting and had loss of optical signal. no harm or injury noted and console was replaced. this was seen on day 2 of the support. the team replaced the automated impella controller (aic) and transferred the pump to a new aic. later on day 3 of the support the patient experienced ventricular tachycardia (vt) that prompted shock x3. the pump additionally had position adjusted within the left ventricle. later on day 6 the patient had more vt runs x6. the patient this time was not shocked. the patient later expired on day 7 after care was withdrawn. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: november 12, 2024. | The investigation of ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23293 was submitted in accordance with updated procedures. g.3 date received by manufacturer was reported incorrectly on the initial manufacturer device report number 1220648-2024-23293. the date should of been 26-sep-2024. h.6 code 1721 was reported incorrectly on the initial manufacturer device report number 1220648-2024-23293. code 4617 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-23293 and was added.
14607 | 1220648-2024-23294 | 20701043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701043&pc=OZD | 09/26/2024 | 12/16/2024 |  | 11/15/2024 |  |  | 09/30/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536779 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Air Embolism | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an unknown primary indication was implanted with an impella cp device for mechanical circulatory support. during support, icu team saw air in the left ventricle (lv), patient is on impella cp support. additionally. the impella had a drop in flow and an increase in purge pressure. the purge line was clamped by mistake while the staff was handling the ecmo line. tpa was added to the purge, but after several days the purge pressure and purge flow was still too high. it was further reported that the impella had a pump stop. the patient expired while on impella support, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into high purge pressure, embolism, and pump stop has been completed. the impella device or data logs were not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-23294. b3 date of event: g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ embolism: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
14608 | 1220648-2024-23295 | 20701057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701057&pc=OZD | 09/30/2024 | 11/15/2024 |  | 11/15/2024 |  |  | 09/30/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2025547101 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient receiving an elective coronary artery bypass graft and mitral valve was implanted with an impella cp for mechanical circulatory support. while manipulating the patient's heart to remove air from the left ventricle (lv), the physician found that the impella had perforated the lv. the impella device was pulled back and a hole was sewn using a patch and pledgeted stitches. heart was de-aired for a second time and device position was optimized. ultimately, care was withdrawn and the patient expired, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported ventricle perforation has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical description, the root cause of the ventricle perforation was a result of the md's manipulation of the heart. as noted in the impella ifu: impella cp with smartassist system and impella 5.5 with smartassist: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14609 | 1220648-2024-23296 | 20701060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701060&pc=OZD | 09/21/2024 | 11/15/2024 |  | 11/15/2024 |  |  | 09/21/2024 | 11/15/2024 |  |  | 05/31/2025 |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024357390 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma and received six units of packed red blood cells due to blood loss at the impella site. additionally, the impella was explanted , the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding could not be determined. as noted in the impella ifu: access site bleeding: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14610 | 1220648-2024-23297 | 20701065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20701065&pc=OZD | 09/20/2024 | 11/25/2024 |  | 11/15/2024 |  |  | 09/20/2024 | 11/15/2024 |  |  |  |  | 11/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545541 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, high purge pressure and low purge flow alarms were triggered. no patient harm reported. | [5 empty] | On manufacturer device report 1220648-2024-23297 the date of 20-sep-2024 was inadvertently listed in g3 the correct date is 04-nov-2024. | The investigation for the reported high purge pressure has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the high purge pressure could not be determined.
14611 | 1220648-2024-23111 | 20682652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20682652&pc=OZD | 08/11/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 08/11/2021 | 11/14/2024 |  |  | 02/28/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022068049 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that at the completion of the case, the patient was sent to vascular surgery to remove the 13x13 sheath after multiple perclose devices failed from circumferential calcium. the patient was successfully weaned and the impella device and sheath were explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14612 | 1220648-2024-23112 | 20682676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20682676&pc=OZD | 10/17/2024 | 12/16/2024 |  | 11/14/2024 |  |  | 10/17/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482953 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, dorsalis pedis pulse was not present and posterior tibial pulse was faint in the right foot. the nurse reported that the patient needed a thrombectomy. additionally, the patient developed torsades therefore defibrillation and magnesium were given. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported ventricular tachycardia (vt) and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt and limb ischemia could not be determined.
14613 | 1220648-2024-23113 | 20682925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20682925&pc=OZD | 09/14/2023 | 09/16/2025 |  | 11/14/2024 |  |  | 10/17/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded the publication entitled "early clinical outcomes of patients with stress-induced cardiomyopathy receiving acute mechanical support in the us" with primary author from yale documented a review of 902 patients, of which 189 had impella support. impella was associated with an increased adjusted risk of 30-day readmission compared to intra-aortic balloon pump (iabp) (aor, 2.53; 95% ci, 1.16- 5.51). patients with extracorporeal membrane oxygenation (ecmo)or impella had a higher incidence of renal replacement therapy and vascular/bleeding complications compared to those who received iabp. conclusions: in this nationwide analysis using an administrative database, patients who received ecmo and impella showed higher rates of in-hospital mortality, renal replacement therapy, and vascular/bleeding complications compared to those who received iabp. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. jang, s.-j., malaguez webber, f., alam, m. m., bae, j. y., aggarwal, a., thomas, a., zarich, s. w., ahmad, t., miller, p. e., & davila, c. d. (2023). early clinical outcomes of patients with stress-induced cardiomyopathy receiving acute mechanical support in the us. journal of the society for cardiovascular angiography & interventions, 2(6), 101185. https://doi.org/10.1016/j.jscai.2023.101185. | The investigation of renal failure and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23113 in accordance with updated procedures.
14614 | 1220648-2024-23114 | 20682966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20682966&pc=OZD | 10/18/2024 | 11/14/2024 |  | 11/14/2024 |  |  | 10/18/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1648111 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomed department reported that there a broken controller, the holding bar on the back of the unit broke and the console is cracked on the side. the damage was observed prior to use therefore no patient involvement. | [5 empty] | The investigation into the automated impella controller has been completed. the impella controller was returned for investigation. the aic logs did not contain any data relevant to this complaint. the reported damage was confirmed. the aic support bracket was broken and there was damage found on the side of the rear housing. the support bracket broke at the weld point. no other damages were observed aside from what was reported. the aic functioned as normal. the root cause of this broken support bracket was unable to be determined because there was not enough information provided as to how the support bracket broke.
14615 | 1220648-2024-23115 | 20683022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683022&pc=OZD | 10/21/2024 | 01/02/2025 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had absent distal pulses. external femorofemoral bypass to supply blood to distal extremity was done. it was further reported that the patient on impella cp support had some runs of ventricular tachycardia overnight. echocardiography in the morning showed impella cannula was deep in the ventricle. ejection fraction was 30%. the physician repositioned the impella cp. additionally, when the impella was explanted post angiogram from the right radial access a large thrombus burden was in the right common iliac, profunda, and mid superficial femoral artery. this required intervention using balloon angioplasty and thrombectomy. the patient remained stable and survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death." | The investigation of limb ischemia, thrombus, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23115 in accordance with updated procedures.
14616 | 1220648-2024-23116 | 20683074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683074&pc=OZD | 10/14/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 10/14/2022 | 11/14/2024 |  |  | 05/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 EU |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022086326 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 2.5 for mechanical circulatory support. during support, the pump stopped. it was reported that the procedure was completed with another device, no harm or injury to patient reported. | [5 empty] | The investigation into the pump stop issue has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, and biomaterial and a green foreign substance was found in the outflow. the green substance and dried dextrose were cleaned off and the immediate pump stop continued to reproduce in the laboratory in a basin of water. the cause of the issue was biomaterial as a result of the patient condition. this report is being filed as part of a retrospective review of historical records.
14617 | 1220648-2024-23117 | 20683075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683075&pc=OZD | 10/18/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/18/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501626 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support when the patient was having scheduled surgery for coronary artery bypass grafting. the impella was pulled out of the left ventricle (lv) and placed in surgical mode. upon completion of the aortic valve replacement, the impella cannula was reinserted into the lv through direct visualization via the aortotomy. upon closure of the aortotomy and filling of the left ventricle, it was discovered that the inlet of the impella had perforated the lv. a stitch repair was performed and the operation continued without issue. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The root cause of the ventricle perforation was most likely due to improper positioning after reinserting pump during procedure. the investigation into ventricle perforation has been completed. the impella device was not returned for evaluation. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23117h6 code 2135 was reported incorrectly n manufacturer device report 1220648-2024-23117. new code was added to health effect - clinical code.
14618 | 1220648-2024-23118 | 20683096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683096&pc=OZD | 10/20/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 10/21/2022 | 11/14/2024 |  |  | 12/31/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-JP |  |  | OZD |  |  | IMPELLA 2.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022023973 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the patient developed lower ischemia. this was resolved by inserting a sheath into the left leg as a preventative measure to send blood flow to the right leg. additionally, a thrombus was found near the outflow cage at time of normal impella removal, no action was necessary. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the limb ischemia and thrombus have been completed. the cause of the limb ischemia was the use of a 16fr medkit introducer that is not indicated for use with impella. per the instructions for use no introducers other than abiomed introducers are approved for use with impella due to improper technique by the end user. the cause of the thrombus was most likely biomaterial based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
14619 | 1220648-2024-23119 | 20683129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683129&pc=OZD | 12/28/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 12/28/2022 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232166 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the pump stopped and was replaced with another device. the procedure was successful with no report of harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the pump stop issue has been completed. the real time (rt) log from the console shows controller error and motor current high pump stop alarms. the cause of the issue was not determined since the product was not returned, and clinical details provided were insufficient to determine the root cause. this report is being filed as part of a retrospective review of historical records.
14620 | 1220648-2024-23120 | 20683130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683130&pc=OZD | 10/21/2024 | 12/09/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522208 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient decompensated in the last 12 hours. the patient was septic, hemorrhagic, and endured cardiogenic shock. a gastrointestinal bleed was noted. the decision was made to withdrawal care to comfort measures and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the gastrointestinal bleed and the sepsis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the gastrointestinal bleed and the sepsis could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14621 | 1220648-2024-23121 | 20683176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683176&pc=OZD | 03/20/2023 | 11/14/2024 |  | 11/14/2024 |  |  | 03/20/2023 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196816 |  | 005042 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during planned explant, the physician had difficulty removing the impella 2.5 through the valve of the impella sheath. the pump go stuck in the introducer and had to be removed by force. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the reported removal issue has been completed. the product was returned, and the issue was confirmed. the pump and introducer were returned, and a kink was observed in the sheath of the 13fr introducer. per the results of the clinical review and investigation, the root cause was determined to be a introducer sheath kink. the sheath kink was likely caused by the patient's vessel conditions. this report is being filed as part of a retrospective review of historical records.
14622 | 1220648-2024-23122 | 20683214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683214&pc=OZD | 10/21/2024 | 12/13/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477166 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and upon explanting, the patient had ventricular tachycardia. the patient was defibrillated. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause of the vt was not determined. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-23122 was submitted.
14623 | 1220648-2024-23123 | 20683227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683227&pc=OZD | 04/05/2023 | 11/14/2024 |  | 11/14/2024 |  |  | 04/05/2023 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136437 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. the pump was inserted but the pump did not start. the team made the choice to explant the impella and the procedure was aborted. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14624 | 1220648-2024-23124 | 20683292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683292&pc=OZD | 10/18/2024 | 12/12/2024 |  | 11/14/2024 |  |  | 10/18/2024 | 11/14/2024 |  |  | 11/30/2025 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434136 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had a right chest pocket washout. the rep reported that the patient was taken off anticoagulation due to continued oozing from insertion sites. the rep had no additional information as to why the washout was performed. it is unknown if heparin was in the purge or not. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23124. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14625 | 1220648-2024-23125 | 20683344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683344&pc=OZD | 07/25/2023 | 11/14/2024 |  | 11/14/2024 |  |  | 07/25/2023 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196815 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Component Missing | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being implanted with an impella 2.5 device for mechanical circulatory support. during prep, the 2.5 was opened, and the white cable was missing from the box. they had no replacements on hand. therefore, the staff made the decision to upgrade the sheath and use an impella cp. | [5 empty] | The investigation for the reported packaging issue has been completed. the impella device was not received from the customer. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14626 | 1220648-2024-23126 | 20683437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683437&pc=OZD | 12/28/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 12/28/2021 | 11/14/2024 |  |  | 11/30/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022036737 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. it was reported that the physician was unable to implant the pump, it was damaged due to the patient's anatomy. the device was removed and replaced with another 2.5. no further information is available. | [5 empty] | The investigation into the delivery issue has been completed. the impella pump was returned for investigation. the clinical details provided were sufficient to determine the cause. the root cause of the delivery issues were due to a complex patient anatomy. this report is being filed as part of a retrospective review of historical records.
14627 | 1220648-2024-23127 | 20683481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683481&pc=OZD | 01/07/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 01/07/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022120746 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during preparation, the staff noted a leak in the pressure reservoir. the resolution for this issue is unknown, no further information is available. | [5 empty] | The investigation into the pump leak/low purge pressure issue has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, and a hole was observed in the reservoir blue membrane. the sidearm was purged with fluid using a syringe and a leak was reproduced from the reservoir blue membrane hole. the cause of the reservoir leak was most likely the user inserting a needle into the reservoir piercing a hole in the membrane due to improper technique by the end user this report is being filed as part of a retrospective review of historical records.
14628 | 1220648-2024-23128 | 20683494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683494&pc=OZD | 09/06/2024 | 12/18/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1105128 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. while on support, placement signal (ps) pressure readings elevated in 200s both aorta and left ventricle. automated impella controller (aic) is being returned for investigation of placement signal issue. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into automated impella controller (aic) - loss of opm data has been completed. after data and device analysis the cause of the elevated placement signal was not determined since the failure mode was not reproduced. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23128.
14629 | 1220648-2024-23129 | 20683558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683558&pc=OZD | 01/17/2023 | 11/14/2024 |  | 11/14/2024 |  |  | 01/17/2023 | 11/14/2024 |  |  | 12/31/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022044224 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Inaccurate Information | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Customer used pump after set expiry date but before components expiry dates. labelling issue but no harm done to the patient. reportable due to packaging issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the packaging issue has been completed. the cause of the packaging issue was not determined since neither the product nor images of the package were returned. this report is being filed as part of a retrospective review of historical records.
14630 | 1220648-2024-23130 | 20683573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683573&pc=OZD | 10/20/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/20/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552949 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support, the 60-year-old male patient¿s legs were cold and had no pulses. the doctor confirmed distal flow past impella using contrast through repositioning unit side port. air in purge alarm occurred after purge bag change; line de-aired without issue. additionally, the patient had hemolyzed blood. impella position was confirmed to be under papillary muscle. showing impella flow reduced and lower than expected flows in auto and p-9. the doctor attempted several times to reposition, at both p-9 and p-2, with insertion sheath still in place. while at p-2, impella popped back across atrioventricular (av). the doctor was unable to re-advance. the doctor replaced impella with a second cp through the same site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis, low pump flow, and positioning issues has been completed. data was not returned for review. pump was returned for analysis. visual inspection found no issues on the pump. based on clinical description and product analysis, the root cause of low flow was most likely due to pump was not in the proper position under the papillary muscle leading to a cannula kink. the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of positioning issue was most likely use related since the dislodgement occurred during repositioning. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email.
14631 | 1220648-2024-23131 | 20683585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683585&pc=OZD | 10/20/2024 | 12/09/2024 |  | 11/14/2024 |  |  | 10/20/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453663 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had a gastro-intestinal (gi) scoped and was positive for upper and lower gi bleed. multiple units of prbcs were given. | It was further reported that the pump was successfully weaned and explanted after several days of support. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported gastrointestinal bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the gastrointestinal bleed could not be determined. added-b5 patient outcome. corrections to the initial mfr report: g1 mdr reporting contact email.
14632 | 1220648-2024-23132 | 20683595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683595&pc=OZD | 10/21/2024 | 01/02/2025 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  | 09/30/2025 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438576 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp placement, the physician had a problem advancing the long and short 14fr peel away sheath through the patient¿s femoral vessel. the 14fr dilators on itself went down the vessel smoothly without the peel-away sheath. the short 13cm dilator and peel-away sheath was used. there was no harm to the patient. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the introducers were returned for investigation and returned introducers had deformation/buckling found along the sheath. the root cause of delivery issue is determined to be use issue because introducers buckled due to excessive force and it was reported that the vessels were smooth. hence, the root cause is determined to be use issue. corrections to the initial mfr report: g1 mdr reporting contact email
14633 | 1220648-2024-23133 | 20683603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683603&pc=OZD | 10/21/2024 | 12/16/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544068 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Tamponade | [11 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 49-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient was emergently taken to the operation room for pericardial window for tamponade. the surgeon had to open chest. upon opening pericardium, it was discovered tamponade cause by blood and not infectious liquid. impella removed to allow physician to aggressively manipulate the heart while on cardiopulmonary bypass (cpb) to locate bleeding. the surgeon stated that there was no evidence of left ventricle (lv) perforation, lv apex, and walls were intact without bleeding. surgeon discovered ¿a split epicardium along the circumflex artery¿. the surgeon stated the bleeding was not a result of the impella. after repairing the bleeding with a suture, the patient was weaned from cpb. the patient was place on peripheral ecmo utilizing the impella left axillary graft for the arterial return cannula. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14634 | 1220648-2024-23134 | 20683609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683609&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559872 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following a successful implant, the impella cp experienced lower than expected flows. upon troubleshooting the device, it was discovered that a kink had formed on the cannula below the outlet. the pump was replaced to resolve the issue. there was no patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. complaint device not returned for investigation. data logs were reviewed, suction alarms were found with low flows since start of the case. the cause of the low pump flow most likely was the cannula kink due to improper technique. corrections to the initial mfr report: g1 mdr reporting contact email.
14635 | 1220648-2024-23135 | 20683627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683627&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555928 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump support ongoing for a 63-year-old male patient admitted in with non-st-segment elevation myocardial infarction (nstemi) and planned surgery. the pump was placed but due to poor positioning against the mitral prompted removal. the pump was removed and replaced by intra-aortic balloon pump (iabp). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. the root cause of positioning issue was most likely patient condition since the patient had a horizontally positioned heart and an aneurysmal aorta.
14636 | 1220648-2024-23136 | 20683630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683630&pc=OZD | 10/21/2024 | 12/16/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536179 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient lost pulses distal to impella. limb ischemia was managed by distal perfusion catheter. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14637 | 1220648-2024-23137 | 20683633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683633&pc=OZD | 10/21/2024 | 12/09/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications during impella 5.5 support. it was documented that following implant, continuous oozing was seen from the insertion site. a jackson-pratt drain was placed but the oozing continued and packed red blood cells were eventually transfused to counteract the bleed. the patient's heart rate climbed and they went into supraventricular tachycardia. cardioversion was completed twice to stabilize the patient. three days later, the bleeding continued and two additional units of blood were given to the patient. seven days into support, the pump began to experience flow saturation. the staff was advised that the pump was struggling and could potentially fail. despite the noted issue, the staff opted to continue use of the device. upon follow up, it was verified that the pump later experienced an unexpected stoppage. the pump was removed and an intra-aortic balloon pump was placed for ongoing support. the patient outcome is unknown. | [5 empty] | The investigation for the pump stop, saturated flow, ventricular tachycardia (vt), and access site bleed has been completed. product was not returned. data logs were reviewed which showed numerous motor current (mc) spikes followed by saturated flows. on the last day of support there is a mc spike with hmc pump stop alarm 13. there is also an increase in purge pressure (pp) and drop in purge flow after the mc spike which ultimately is resolved. the pump is unable to be restarted. there are also a few suction alarms throughout the case. the root cause of the saturated flow is most likely biomaterial. the root cause of the pump stop is most likely biomaterial. the root cause of the access site bleeding major and vt was not determined since no product was returned and lack of clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿.
14638 | 1220648-2024-23138 | 20683655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683655&pc=OZD | 10/21/2024 | 11/18/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167998 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) failed to recognize the fiberoptic sensor when the white connector cable was plugged in. troubleshooting was attempted, but the failure remained. the aic was subsequently swapped and the issue resolved. there was no patient harm. | [5 empty] | The following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23138. | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. console logs from the reported day of event are consistent with complaint and show that 3 attempts have been made to start the case. during 1st attempt, error message ¿optical sensor system error¿ presented and operators have been prompted to disconnect and reconnect the pump, which did not resolve the issue. on the 2nd attempt, the issue remained unresolved and same error message presented. after disconnecting the pump and reconnecting again, case start wizard prompted to switch the controller two minutes later, pump was disconnected and transferred to the backup console (ic5503). corresponding rtlog data shows that after 3rd connection of the pump on the original (complaint) console, the issue was resolved and optical placement signal was valid, however, staff swapped the controller to a backup due to earlier prompt from case start wizard ¿optical sensor system error - switch controller¿. during 1st and 2nd attempt, signals from the optical bench recorded in rtlogs are consistent with an air gap, producing low snr. ltlog data from ic5503 indicated that placement signal issues (optical sensor system error or placement signal not reliable) did not repeat on the backup console. the console was received in danvers and inspected, and minor contamination (debris) were found on the optical pump connector fiber. after cleaning, the 5s parameters of the optical system were well within specification. reported issues with optical system were most likely caused by an air gap between optical fiber in optical pump connector and optical fiber in console¿s pump receptacle. this might have been caused either by debris in the receptacle (dislodged after the event) or could have been use-related (e.g. pump plug not pushed all the way in). the cause of the optical signal failure was an air gap from between the aic and the patient cable plug.
14639 | 1220648-2024-23139 | 20683659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683659&pc=OZD | 10/22/2024 | 12/06/2024 |  | 11/14/2024 |  |  | 10/22/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1629362 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller (aic) ongoing for a patient admitted in acute myocardial infarction and cardiogenic shock. on day 2 of the support the team observed purge disc damage on the aic and removed the aic and replaced it with another console. no patient harm or injury was noted from the interruption in the support. | [5 empty] | The investigation for the console (aic) purge disc/flag issues has been completed. data logs were not reviewed as no data review is needed for this complaint because it was only reported that the retainer was broken. the console was returned for evaluation. the purge disc retainer was found to be broken. the root cause of the issue was broken purge disc retainer. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
14640 | 1220648-2024-23140 | 20683685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683685&pc=OZD | 08/16/2021 | 05/30/2025 |  | 11/14/2024 |  |  | 08/16/2021 | 11/14/2024 |  |  | 04/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022077506 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 92 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, they had a successful procedure with the impella 2.5 performance, however closure with proglide failed, so double angioseal was sufficient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding - major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14641 | 1220648-2024-23141 | 20683696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683696&pc=OZD | 10/20/2024 | 12/24/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484898 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that the patient was placed on impella cp for cardiogenic shock due to cardiac sarcoidosis. six days later, bleeding occurred from the lumbar artery and interventional radiology (ivr) was performed. a follow-up computed tomography (ct) scan in the morning showed dissection. impella was removed to treat the descending aorta. it is judged that the dissection of the descending aorta was caused by the ivr procedure. the cause of the bleeding was not related to impella but because heparin was administered, it may not be possible to say that there is absolutely no connection between the bleeding and impella. the physician judged that the bleeding was due to inflammation caused by cardiac. it was further reported that on the same day as the dissection of the descending aorta, lower limb ischemia was also observed, so stent placement was performed. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the vascular damage and limb ischemia has been completed. no product was returned for analysis. the clinical does not mention any instance of excessive force or relevant patient condition issues. the root cause of the vascular damage - peripheral vessel and limb ischemia was not determined as limited clinical information was provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14642 | 1220648-2024-23142 | 20683809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683809&pc=OZD | 08/18/2021 | 05/30/2025 |  | 11/14/2024 |  |  | 08/18/2021 | 11/14/2024 |  |  | 03/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022066303 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 95 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, after weaning the impella 2.5 the pro glide failed and closure was with 8 f angi seal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding - major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14643 | 1220648-2024-23143 | 20683998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20683998&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that all proper steps were taken at impella cp insertion point. while advancing the catheter, the physician felt resistance in the peel away sheath. the physician believed there was a kink in the catheter and made the decision to remove and assess. the catheter showed signs of a kink. a decision was made to place a new catheter. the impella was exchanged. there was no harm to the patient. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the product damage has been completed. product was not returned for review. the root cause of product damage that occurred during delivery was most likely due to patient condition since the patient had tortuous vessels. added h6 impact code 4629 corrections to the initial mfr report and g1 mdr reporting contact email.
14644 | 1220648-2024-23144 | 20684034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684034&pc=OZD | 01/28/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 01/28/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022111503 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during delivery, the sheath was inserted and the impella was placed through the sheath. the device was stuck and could not be advanced, a fluoroscopy was performed, and no obstruction or kink was identified. the physician decided to remove the device and replace it with another impella 2.5. it was reported that the procedure was successful. | [5 empty] | The investigation into the delivery issue has been completed. the impella pump was returned for investigation. both the impella 2.5 and the introducer were returned for analysis. the pump was placed through the introducer and got stuck with only half of the inflow cage exiting the introducer. the inner diameter (id) of the introducer was measured with gauge pins to be approximately 0.171in. the outer diameter (od) of the pump was checked using gauge rings and was found to be between .1732in and 0.1736in. the od of the pump was larger than the id of the introducer. the root cause of the pump being unable to fit through the introducer was determined to have a large pump outer diameter as a result of the manufacturing process. this report is being filed as part of a retrospective review of historical records.
14645 | 1220648-2024-23145 | 20684044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684044&pc=OZD | 10/21/2024 | 12/17/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541809 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had limb ischemia. pulses were lost and the impella was weaned and removed. post removal, there was a dissection noted to the right iliac. the physician felt that the dissection occurred during the implant and caused the ischemia. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and vascular damage could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14646 | 1220648-2024-23146 | 20684126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684126&pc=OZD | 10/21/2024 | 01/03/2025 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544539 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with animpella cp. however, it was aborted due to patient anatomy. no patient harm has been reported. | [5 empty] | The pump has been returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related. due to the patient's anatomy, implantation was not possible as per the clinical description provided. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-23146. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-23146 was submitted in accordance with updated procedures.
14647 | 1220648-2024-23147 | 20684185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684185&pc=OZD | 02/23/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 02/23/2022 | 11/14/2024 |  |  | 05/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022085296 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that after insertion of the repositioning sheath, the physician was not able to close the touhy-borst valve. the pump was not secured; however the patient was not harmed during the procedure. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. the returned tuohy-borst was visually inspected and no abnormalities were observed. the tuohy-borst was able to sufficiently tighten down. the tuohy-borst was torn down and no thread damage was observed, and the mating components were able to fully thread together. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
14648 | 1220648-2024-23148 | 20684199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684199&pc=OZD | 07/01/2021 | 09/23/2025 |  | 11/14/2024 |  |  | 07/01/2021 | 11/14/2024 |  |  | 09/30/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1499302 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was going to be on impella 2.5 for mechanical circulatory support. during delivery of the impella, the implant was aborted due to the patient's vascular anatomy. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23148 in accordance with updated procedures.
14649 | 1220648-2024-23150 | 20684236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684236&pc=OZD | 03/08/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 03/08/2022 | 11/14/2024 |  |  | 10/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173247 |  | 004413 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that during delivery, a small resistance was felt at the level of the aortic arch. additionally, while explanting the 2.5 pump had a resistance pulling through the 13 fr abiomed sheath. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the delivery/removal issue has been completed. the cause of the removal issue was not determined since product was not returned data analysis logs were not available and the clinical details provided were insufficient to determine a root cause. the cause of the delivery issue was patient condition as determined from the clinical details indicating the resistance at the aortic arch while removing the device. this report is being filed as part of a retrospective review of historical records.
14650 | 1220648-2024-23149 | 20684238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684238&pc=OZD | 10/18/2024 | 01/02/2025 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025469879 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had an infection. the pump was explanted and was replaced on the other side. the patient survived the explant. | Additional information was received. it was reported that the patient's condition did not improve. the reported infection was unable to be controlled and the cardiac function did not recover and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b.2 death and date of death were added. b.5 new information was received. h.1 type of reportable event was updated to reflect new information. h.6 health effect - impact code updated to reflect new information. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-23149 was submitted in accordance with updated procedures.
14651 | 1220648-2024-23151 | 20684739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684739&pc=OZD | 03/22/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 03/22/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128916 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the device exhibited suction issues indicating low flow. the patient was hemodynamically stable and did not require any vasopressors. the physician decided to explant the device. it was reported that the patient survived explant and the procedure. | [5 empty] | The investigation into the product low pump flow has been completed. the impella pump was returned for investigation. no abnormalities were found on the returned product. the low flow was not reproduced in the lab and the pump ran at 3.0 liters per minute. thus, the root cause of the low flow was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
14652 | 1220648-2024-23152 | 20684881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684881&pc=OZD | 08/22/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 08/22/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022152462 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation of an impella 2.5, when connecting the cable of the 2.5 catheter to the console, the automated impella controller would not advance to the next step. the complainant replaced the connecter cable and the console, however this did not resolve the reported issue. a new impella 2.5 catheter was prepped and used with no issues. there was no patient harm reported. | [5 empty] | A2 through a5 patient information was unavailable. d6a, d6b: implant and explant information was unavailable. the impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the product damage cannula kink has been completed. the root cause of the damage was unable to be determined as no product or logs were returned for evaluation and clinical details provided were insufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
14653 | 1220648-2024-23153 | 20684908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20684908&pc=OZD | 10/19/2024 | 12/09/2024 |  | 11/14/2024 |  |  | 10/19/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552300 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 82-year-old male patient suffered from an unknown bleed during impella cp support. the patient was transfused two units of packed red blood cells and heparin was withdrawn. planned wean and explant of the device was successful later that same day. it was noted that the perclose failed following explant. manual pressure was applied until hemostasis was achieved. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported unknown blood loss has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the unknown blood loss could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email.
14654 | 1220648-2024-23154 | 20685021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685021&pc=OZD | 09/19/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 09/19/2022 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023220792 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported at the removal of the pump, the pump became stuck within the 13fr sheath. the medical team decided to pull the sheath and the catheter as one unit with no manual pressure until completely removed. the impella 2.5 and sheath were successfully removed with no patient injury reported. | [5 empty] | The investigation into the delivery/removal issue has been completed. the impella pump was returned for investigation. the issue was reproduced. the cause of the pump stuck in the introducer issue was an insufficient sheath inner diameter. the root cause of the sheath inner diameter issue was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14655 | 1220648-2024-23155 | 20685149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685149&pc=OZD | 10/21/2024 | 12/24/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482371 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the female patient was on impella cp support. while on support, the patient developed ischemia in the impella leg. the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: d4-udi number updated. g1 mdr reporting contact email.
14656 | 1220648-2024-23156 | 20685262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685262&pc=OZD | 10/18/2024 | 12/05/2024 |  | 11/14/2024 |  |  | 10/18/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1361559 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) did not recognize the purge cassette even after several attempts. another aic was used and was successful. support was delivered to the patient uninterrupted. | [5 empty] | The investigation into automated impella controller (aic) - purge disc/flag issues has been completed. after reviewing the imc logs reported issue that purge disc couldn't be detected was confirmed. during the device analysis the purge disc flag was found to be broken (missing at blue retainer). g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23156. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two medical device reports associated with the event. this report represents the automated impella controller and manufacturer device report number 1220648-2024-23082 represents the pump.
14657 | 1220648-2024-23157 | 20685307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685307&pc=OZD | 03/22/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 03/22/2022 | 11/14/2024 |  |  | 10/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023164326 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that the impella 2.5 got stuck in the peel away sheath during removal. the physician pulled out everything as a single unit, no bleeding or harm to patient. the patient survived explant and the procedure. | [5 empty] | The investigation into the product removal issue has been completed. the impella pump was returned for investigation. the evaluation into the removal issue has been completed. the visual inspection of the returned pump noted the pump was stuck in the introducer. significant amounts of clotting was observed in the introducer. the impella was able to be removed after soaking the products. upon removal, significant amounts of clotting and blood exited the introducer. the impella was measured with the ring gauges and its diameter was within specification as all ring gauges were able to pass over the motor housing. the inner diameter of the introducer was also confirmed to be within specification using pin gauges. the impella was able to pass through the introducer upon reinsertion, but the clotting in the introducer did not allow for further advancement. therefore, it was determined that the cause of the removal issue was patient condition related. the significant clotting within the introducer sheath prevented the impella from being removed through the introducer. this report is being filed as part of a retrospective review of historical records.
14658 | 1220648-2024-23159 | 20685358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685358&pc=OZD | 10/17/2024 | 12/12/2024 |  | 11/14/2024 |  |  | 10/17/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534193 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support was switched back to heparan sulfate after seeing a left ventricle thrombus on echocardiogram. additionally, blood and condensation was found inside the sterile sleeve overnight. volume of blood was very minor and has stopped bleeding since originally found. at end of unit support the patient was bridged to a left ventricular assist device. the patient survived the explant. | [5 empty] | The investigation into product damage (sterile sleeve damage) and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23159. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
14659 | 1220648-2024-23160 | 20685364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685364&pc=OZD | 03/30/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 03/30/2022 | 11/14/2024 |  |  | 10/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC |  |  | US |  | 2022148165 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the pump stopped. the staff tried to restart the pump; however, these attempts were not successful. the device was explanted and replaced with another impella 2.5. it was reported that the new device worked without issue during the entire procedure. no further information is available. | [5 empty] | The evaluation into the pump did not start has been completed. the returned product was visually inspected, and a fractured red lumen was observed. the location of the fracture was at the pump outflow. the cause of the issue was a red lumen removal issue due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
14660 | 1220648-2024-23158 | 20685391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685391&pc=OZD | 03/28/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 03/28/2022 | 11/14/2024 |  |  | 10/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170290 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 2.5 for mechanical circulatory support. however, during delivery the physician was unable to advance the impella past the iliac bifurcation due to calcium and tortuosity. the device was removed and instead utilized an intra-aortic balloon pump. no further information is available. | [5 empty] | The investigation into the product delivery issue has been completed. the impella pump was returned for investigation. the clinical details provided were sufficient to determine the cause of the delivery issues as calcium and tortuosity as the physician was unable to advance the impella 2.5 past the iliac. this report is being filed as part of a retrospective review of historical records.
14661 | 1220648-2024-23161 | 20685426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685426&pc=OZD | 04/27/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 04/27/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022148728 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. during delivery, the impella was placed to the level of the aortic valve and would not cross. the wire came out of the left ventricle and the impella was removed and replaced with a new impella 2.5. it was noted that the patient's aortic valve was calcified no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the delivery issue has been completed. the cause of the issue was a calcified valve. the pump would not pass the valve and likely was damaged upon attempts to pass the valve. the pump would not cross the iliac bifurcation upon re-insertion likely due to the pump damage. aortic valve stenosis/calcification is a contraindication for use, this was misuse by the end user. this report is being filed as part of a retrospective review of historical records.
14662 | 1220648-2024-23163 | 20685521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685521&pc=OZD | 06/10/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 06/10/2022 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188636 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. it was reported that during delivery the impella was unable to advance past the iliac. it was reported that the device was removed. no additional information is available. | [5 empty] | The investigation into product delivery issue has been completed. the impella pump was returned for investigation. a visual inspection of the returned product indicated there was a catheter kink. based on the clinical details and visual inspection the cause of the delivery issue was the patient condition as the physician was unable to advance the impella 2.5 past the iliac. this report is being filed as part of a retrospective review of historical records.
14663 | 1220648-2024-23162 | 20685539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685539&pc=OZD | 10/21/2024 | 12/27/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481592 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that shortly after the impella cp implantation and during the intervention, high motor current values and low flow rates were noticeable. pump stop occurred after about half an hour after impella started. the cp was explanted, and a thrombus was clearly visible in the inlet and cannula. activated clotting time values were all the time in the preferred range. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and thrombus has been completed. data logs were not returned for investigation. returned product was investigated and the pump passed the electrical test, confirming it likely was not an electrical issue. additionally, there were no other abnormalities with the pump upon visual inspection, including any biomaterial in the purge gap or cannula. the pump was then connected to a console and run for 30 minutes without issue. though clinical details report high motor currents and biomaterial upon explant, the root cause of the pump stop could not be determined without data logs. as the necessary information was not provided, the root cause of the thrombus was not determined." the following have been reported incorrectly or missing from manufacturer device report. b1 should have been both adverse event and product problem g1 reporting contact email revised in accordance with updated procedures.
14664 | 1220648-2024-23164 | 20685610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685610&pc=OZD | 06/20/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 06/20/2022 | 11/14/2024 |  |  | 06/30/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1486020 |  | 005042-RO |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the physician was unable to connect the gray-to-gray adapter. staff noted that the prongs were bent on the connection side of the pump. the device was removed and replaced with a new impella 2.5. it was reported that the procedure was successful. | [5 empty] | The investigation into product damage issue has been completed. the impella pump was returned for investigation. a visual inspection of the returned product confirmed that the pins inside the impella 2.5 grey connector were broken. the root cause of the broken pins was determined to be use related as the pins are subject to breaking if the connection is not made properly. the cause of the issue was damaged connector pins due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
14665 | 1220648-2024-23165 | 20685699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685699&pc=OZD | 11/12/2021 | 04/22/2025 |  | 11/14/2024 |  |  | 11/12/2021 | 11/14/2024 |  |  | 04/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022073971 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Sepsis | [11 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was an onset of renal failure and sepsis. the patient continued to deteriorate with multiorgan failure due to sepsis and expired while on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14666 | 1220648-2024-23166 | 20685772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685772&pc=OZD | 10/17/2024 | 12/26/2024 |  | 11/14/2024 |  |  | 10/21/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546808 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 62 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Ischemia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured limb ischemia with a "unknown (undetermined)" relationship to the impella 5.5 device used: ¿10/17 mi from osh, iabp placed, vt code upon induction of anesthesia, insertion impella left ventricular support device (impella 5.5), right axillary conduit creation with 10 mm gelweave graft direct anastomosis, emergent peripheral venoarterial ecmo initiation via 21 french left common femoral artery cannula and 25 french multi stage left common femoral vein cannula, removal of intra-aortic balloon pump from the right common femoral artery, limb ischemia requiring limb reperfusion cannula (901) antegrade perfusion cannula placement: 7fr antegrade perfusion cannula placed in l sfa with us and fluoro guidance.¿ the patient recovered with no sequelae. it was further reported that the patient expired while on impella 5.5 support. | A notification was received on 10/28/2024 via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella device. it was reported that the patient/event had not recovered/not resolved. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. | The investigation of limb ischemia, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23166 in accordance with updated procedures. g2 report source missing from manufacturer device report 1220648-2024-23166. | Added-b5 hemolysis failure mode added, b6-relevant lab data, h6 clinical code 1886.
14667 | 1220648-2024-23167 | 20685843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685843&pc=OZD | 01/05/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 01/05/2022 | 11/14/2024 |  |  | 02/28/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064679 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, upon explant, attempted closure with preclose without success, wire access maintained, after several attempts one perclose did take in known highly calcified vessel. a 6fr angioseal was also used, and pressure held for 20 min to make sure no oozing at site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14668 | 1220648-2024-23168 | 20685907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20685907&pc=OZD | 09/12/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 09/12/2021 | 11/14/2024 |  |  | 10/31/2021 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1463394 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, bleeding was observed from all access sites. the medical staff was concerned that the patient would experience disseminated intravascular coagulation. therefore, several blood products were provided, and a fem stop was placed in an effort to manage the complication. the device was explanted, and the procedural outcome was the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14669 | 1220648-2024-23169 | 20686031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686031&pc=OZD | 10/09/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 10/09/2021 | 11/14/2024 |  |  | 02/28/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022070506 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support bleeding was noted from all sites, purge blocked alarm occurred, and the patient experienced a stroke. no heparin was in the purge or administered systemically due to stroke and the reported bleeding. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14670 | 1220648-2024-23170 | 20686097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686097&pc=OZD | 08/24/2022 | 05/27/2025 |  | 11/14/2024 |  |  | 08/24/2022 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219860 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during implantation, the procedure was aborted due to iliac anatomy. no other information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14671 | 1220648-2024-23171 | 20686119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686119&pc=OZD | 09/26/2022 | 05/27/2025 |  | 11/14/2024 |  |  | 09/26/2022 | 11/14/2024 |  |  | 11/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217496 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced moderate hemolysis evident in tea colored urine. hematocrit (hct) had been dropping requiring transfusions. in addition, the patient was intubated and sedated; source of gi bleed found was ascending colon, embolization not done. care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14672 | 1220648-2024-23172 | 20686175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686175&pc=OZD | 11/29/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 11/29/2021 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130993 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, the impella was removed due to imperfect position. it was reported that after 25.50 hours of support that care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14673 | 1220648-2024-23174 | 20686674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686674&pc=OZD | 06/01/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 06/01/2022 | 11/14/2024 |  |  | 10/31/2023 |  | 11/14/2024 |  | NAME: TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023178256 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. the impella was weaned slowly, however at the completion of case, a bicep hematoma was noted. protamine was given to resolve the issue and the device explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14674 | 1220648-2024-23173 | 20686739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20686739&pc=OZD | 11/30/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 11/30/2021 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022112600 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was some bleeding around the access site and oozing overnight with no hematoma noted. the case is absent of how the access site bleeding and oozing was controlled. two days later, ground emesis was noted. 300cc drainage from chest tube. the patient had received a total of 5 units of packed red blood cells. ultimately the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14675 | 1220648-2024-23175 | 20687195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20687195&pc=OZD | 07/11/2022 | 05/27/2025 |  | 11/14/2024 |  |  | 07/11/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120628 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. it was reported that the patient's urine was red, so heparin was placed on hold. labs were not hemolyzed, however the patient had significant bleeding noted at the urethra. dopamine was weaned to a low dose. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14676 | 1220648-2024-23176 | 20687251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20687251&pc=OZD | 06/17/2021 | 08/05/2025 |  | 11/14/2024 |  |  | 06/17/2021 | 11/14/2024 |  |  | 03/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060342 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | Vascular Dissection | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, it was noted that weaning failed as the patient experienced tachycardia and hypotension. the echo showed significant pericardial effusion and pericardial puncture. 150 milliliters was directly aspirated and the patient stabilized immediately. second cardiac catheterization showed no more covered dissection and active bleeding. the dissection was fixed surgically. impella was explanted afterwards in the operating room. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23176 in accordance with updated procedures.
14677 | 1220648-2024-23177 | 20687856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20687856&pc=OZD | 02/07/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 02/07/2022 | 11/14/2024 |  |  | 06/30/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139302 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. during impella placement, a perforation of the iliac artery was observed after cannulation was complete. no further information was provided as to treatment or medical intervention. additional information was provided that noted that the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14678 | 1220648-2024-23178 | 20687859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20687859&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and had decreased blood flow in the right leg. a decision was made to place a right femoral artery anti grade sheath for external femoral bypass to increase flow. doppler patient pulses were present post procedure. it was further reported that the family decided to withdraw care and the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23178.
14679 | 1220648-2024-23179 | 20688002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688002&pc=OZD | 02/12/2022 | 11/14/2024 |  | 11/14/2024 |  |  | 02/12/2022 | 11/14/2024 |  |  | 09/30/2022 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021004085 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Cardiac Arrest | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. during impella placement, flows were noted to be limited due to a kink in the cannula. it was stated that the patient began to decompensate, and the decision was made to remove the impella 2.5 and escalate care to an impella cp. it was reported that the patient coded after the impella 2.5 was explanted and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14680 | 1220648-2024-23180 | 20688006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688006&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544041 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was admitted with acute myocardial infarction/cardiogenic shock had impella 5.5 pump support ongoing. the team placed the pump as an escalation from an impella cp. the patient developed thrombocytopenia. the pump remained on until support was withdrawn after 19 days. the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23180 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14681 | 1220648-2024-23181 | 20688313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688313&pc=OZD | 10/23/2024 | 12/20/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551014 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had access site bleeding. a femstop was applied and heparin was withheld. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email.
14682 | 1220648-2024-23182 | 20688654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688654&pc=OZD | 10/23/2024 | 01/02/2025 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558945 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella cp pump for mechanical circulatory support. it was documented that following insertion, the pump was unable to advance into the aorta due to the patient's vascular anatomy. the implant was subsequently aborted due to the complication. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email revised on manufacturer device report 1220648-2024-23182 in accordance with updated procedures.
14683 | 1220648-2024-23184 | 20688942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688942&pc=OZD | 10/23/2024 | 12/18/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  | 05/31/2025 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375101 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for the support of a 17-year-old patient with question of myocarditis and possible right sided failure. the impella cp was placed and patient was intubated due to the condition. the access site oozed and the team placed a suture to remedy the issue. the pump remained on for support while diagnosis was being made. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since hemostasis was achieved by adding additional suture. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23184 was submitted.
14684 | 1220648-2024-23183 | 20688984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20688984&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547456 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported delivery issues with an impella cp. the doctor had difficulties placing in the mid left ventricle. wire access was lost and the pump was pulled out to rewire. once the cannula was out, it was inspected and a kink was observed. no patient harm has been reported. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues has been completed. the pump was returned for investigation and kink was not on the cannula. the root cause of delivery issue was determined to be use issue because the second pump was placed successfully. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23183 was submitted in accordance with updated procedures.
14685 | 1220648-2024-23185 | 20689270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20689270&pc=OZD | 10/23/2024 | 12/05/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1121973 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had sent an automated impella controller (aic) in for preventative maintenance. when logged in, the aic was found to have damage, possibly from being dropped. the aic malfunction could not have caused any harm or injury as the issue was found outside of patient use. | [5 empty] | The investigation into automated impella controller (aic) - damage before therapy has been completed. the console logs did not contain any data relevant to this complaint. the reported damages to the rear housing and the purge door were confirmed. aside from the reported damages found on the rear housing and purge door no other issues were observed after inspection of the aic and its internal components. the cause of the issue was most likely use related. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23185. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14686 | 1220648-2024-23186 | 20689533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20689533&pc=OZD | 10/23/2024 | 11/29/2025 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094476 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella training, the automated impella controller screen would not light up. a green light came on but the screen was blank and there were no lights. after troubleshooting, the console was still not working. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14687 | 1220648-2024-23187 | 20689758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20689758&pc=OZD | 10/23/2024 | 11/30/2025 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168285 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) was displaying suboptimal flows and the left ventricle waveforms were low with suction. the team replaced the aic and the waveforms and flows were back to normal limits. echocardiogram confirmed good position. no harm or injury wasn reported. the patient remained on support with the second console. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. data analysis: the case associated with the reported complaint was started on 22 oct 2024 with the pump sn: (b)(6). after 13 hours since the pump was started, the console displayed ¿suction (det. via char. curve or placement signal) ¿ alarm.¿ this confirms the reported failure mode. rt log shows motor current spike to 1242, indicating signs of biomaterial ingestion. lt log shows the mean flow (l/min) was reducing since the suction alarm. after 3 hours, the pump was moved to (b)(6). though the mean flow and the motor current fluctuated initially, they eventually became stable; most likely, the biomaterial could be removed from the impella, probably by adding blood or by changing the pump position. the second console (b)(6). does not have a suction alarm as per the lt log alarms. device analysis: this console was tested after arriving in fs. unable to reproduce the reported failure mode. upon running the optical test pump at p-9, the motor current was around 950 ma, the same as tested with known good test fixtures. the applied external pressure was measured and displayed in the placement signal window without any issues. root cause: no issue was found with the console (b)(6). hardware related to the reported failure mode. according to the disposable investigation (b)(4), the root cause of the low pump flow was most likely the ingestion of biomaterial. d2 common device name updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
14688 | 1220648-2024-23188 | 20689867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20689867&pc=OZD | 10/23/2024 | 12/23/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520153 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, the purge system was blocked. the cassette was changed and tissue plasminogen activator was started and there were no changes. a pump stop occurred. the impella 5.5 was explanted and replaced. the patient survived the explant. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation was completed. the impella 5.5 was returned for evaluation. the pump was returned with a cut at the catheter. upon visual inspection, a kink in the catheter just distal of the kink protector and red handle was noted. blood was present in the purge gap and in the purge line just below the motor housing. the pump was unable to be electrically tested and was unable to be run in attempt to reproduce the issues. insufficient data logs were returned for evaluation. the pump stop occurred later that same day during pump exchange. the root cause of the pump stop was most likely due to a purge line kink just distal of the red handle which eventually lead to saturated flow due to blood ingress in the motor and purge line which eventually lead to a pump stop. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23188 in accordance with updated procedures.
14689 | 1220648-2024-23189 | 20690991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20690991&pc=OZD | 10/21/2024 | 12/26/2024 |  | 11/14/2024 |  |  | 10/23/2024 | 11/14/2024 |  |  |  |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564643 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump being placed to support a patient admitted in for aortic repair. it was reported that the impella 5.5 tried to be inserted but it was not possible to pass the artificial blood vessel anastomosis after arch replacement. the case was aborted and there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since it was not possible to pass the artificial blood vessel anastomosis after arch replacement. corrections to the initial mfr report: g1 mdr reporting contact email
14690 | 1220648-2024-23190 | 20691959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20691959&pc=OZD | 08/13/2021 | 11/14/2024 |  | 11/14/2024 |  |  | 08/13/2021 | 11/14/2024 |  |  | 05/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022084815 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during prep of an impella 2.5 that they were unable to use the easy guide lumen to place the guidewire within the device. as a result, a new impella cp was used. there was no adverse impact to the patient. | [5 empty] | The investigation for the reported introducer damage issue has been completed. the product was returned, and the issue was confirmed. the impella 2.5 was returned and the red lumen was still in place. a sample guidewire was attempted to be backloaded and the failure was reproduced. the guidewire was unable to advance past the outflow/impeller area. the red lumen was removed and a kink was found in that area. per the results of the clinical review and investigation, the root cause was determined to be backloading issues as a result of a kinked red lumen. the cause of the kink in red lumen could potentially be the lumen being caught between the impeller and the outflow cage. this report is being filed as part of a retrospective review of historical records.
14691 | 1220648-2024-23191 | 20691982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20691982&pc=OZD | 11/05/2022 | 05/27/2025 |  | 11/14/2024 |  |  | 11/05/2022 | 11/14/2024 |  |  | 01/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022047533 |  | 005046 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 15 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during the procedure the pigtail got snagged on the mitral valve leaflet. the impella was removed and reinserted. the correct position was obtained and the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14692 | 1220648-2024-23192 | 20692020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20692020&pc=OZD | 11/21/2022 | 10/20/2025 |  | 11/14/2024 |  |  | 11/21/2022 | 11/14/2024 |  |  | 05/31/2023 |  | 11/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022081333 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, after the removal of the device, hemostasis was achieved with angioseal and an additional femostop due to proglide not fixed. | Us complainant reported that hemostasis was achieved after the removal of the impella 2.5 with angioseal and a additional femostop due to proglide not fixed. mdr, mir and pmda reports required. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
14693 | 1220648-2024-23003 | 20670830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670830&pc=OZD | 06/08/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 06/08/2021 | 11/13/2024 |  |  | 08/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1494563 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the impella pump experienced suction alarms despite the pump being well positioned. it was stated that the pump speed level was decreased, and no additional suction alarms were received. the weaning was successful, and the device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14694 | 1220648-2024-23004 | 20670845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670845&pc=OZD | 07/15/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/15/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022065683 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, a small hematoma at the end of case which was resolved. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14695 | 1220648-2024-23005 | 20670853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670853&pc=OZD | 07/20/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/20/2021 | 11/13/2024 |  |  | 11/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022039525 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during implantation, it was noted that due to femoral artery stenosis, the impella placement was not possible and implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14696 | 1220648-2024-23006 | 20670855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670855&pc=OZD | 07/21/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/21/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022069114 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, it was noted that there was a swap over to sodium bicarbonate for purge alternative due to the patient bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14697 | 1220648-2024-23007 | 20670860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670860&pc=OZD | 07/27/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/27/2021 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022086182 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support the pump dislodged due to manipulation of the catheter. it was totally removed and implanted a second time after rewiring. it was also noted that the patient experienced heparin induced thrombocytopenia during support. actions taken to resolve are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14698 | 1220648-2024-23008 | 20670871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670871&pc=OZD | 07/28/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/28/2021 | 11/13/2024 |  |  | 12/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC |  |  | US |  | 2022046762 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, there was a clot noted in right axillary artery and plans to amputate all extremities except for left arm due to lost pulses at the limb was noted. continuous renal replacement therapy was also initiated during impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14699 | 1220648-2024-23009 | 20670919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670919&pc=OZD | 07/28/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/28/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072084 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted impella could not be advanced due to tortuosity and calcium, so exchanged for long oscor sheath. the impella 2.5was rewired and inserted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14700 | 1220648-2024-23010 | 20670986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670986&pc=OZD | 08/03/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/03/2021 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022096732 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that tip of the wire fractured off. action taken was unknown. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14701 | 1220648-2024-23011 | 20670999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20670999&pc=OZD | 08/16/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/16/2021 | 11/13/2024 |  |  | 06/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1485970 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. after impella weaning and removal, the attempt to close the access site utilizing perclose and angioseal were unsuccessful. manual pressure was held in an effort to resolve the issue. after compression was complete, it was stated that the patient's distal limb was mottled and doppler pulses were not present. a decision was made to perform a contralateral distal angiogram. it was reported that the angiogram identified an occlusion of the common femoral artery. a balloon angioplasty was initiated and flow to the distal limb was restored. it was reported that a thrombus in the patient's popliteal was aspirated. good doppler pulses post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14702 | 1220648-2024-23012 | 20671024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671024&pc=OZD | 12/04/2021 | 03/18/2025 |  | 11/13/2024 |  |  | 12/04/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022131138 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | Life Threatening | Death | Other | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, it was reported that the patient went into ventricular tachycardia, was resuscitated twice, then went into atrial fibrillation. additionally, an x-ray was performed and showed the motor current was flat. it was confirmed that the catheter was not in the left ventricle. the physician decided to explant the device. the patient survived the procedure, no report of harm or injury. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14703 | 1220648-2024-23013 | 20671040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671040&pc=OZD | 12/02/2021 | 02/24/2025 |  | 11/13/2024 |  |  | 12/02/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022124981 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during delivery, the physician reported that the sheath placement in the right groin was difficult. the second attempt in the left groin was successful. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issues has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14704 | 1220648-2024-23014 | 20671046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671046&pc=OZD | 12/06/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 12/06/2021 | 11/13/2024 |  |  | 11/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022039158 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 2.5 for mechanical circulatory support. during support, the device was accidently pulled back into the aorta, this occurred while the balloon was removed. the physician tried to re-cross unsuccessfully and decided to explant the device. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14705 | 1220648-2024-23015 | 20671048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671048&pc=OZD | 12/10/2021 | 03/17/2025 |  | 11/13/2024 |  |  | 12/10/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120634 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Deformation Due to Compressive Stress | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. it was noted that the patient has peripheral vascular disease, an angiogram was performed, and it was noted that the 13fr sheath was occlusive. the patient complained of severe leg pain. the physician decided to explant the impella 2.5 and replace with an intra-aortic balloon pump. hemostasis was achieved and it was reported that the procedure was successful. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
14706 | 1220648-2024-23016 | 20671055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671055&pc=OZD | 12/16/2021 | 03/18/2025 |  | 11/13/2024 |  |  | 12/16/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022057726 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Death | Other | Required Intervention | Life Threatening | [21 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the patient went into atrial fibrillation and was administered metaprotocol. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14707 | 1220648-2024-23017 | 20671059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671059&pc=OZD | 12/21/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 12/21/2021 | 11/13/2024 |  |  | 12/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078200 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. during delivery, it was noted that the patient's aorta was calcified. the physician decided to proceed with the procedure without the impella due to patient anatomy complications. no further information is available. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14708 | 1220648-2024-23018 | 20671066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671066&pc=OZD | 12/22/2021 | 04/03/2025 |  | 11/13/2024 |  |  | 12/22/2021 | 11/13/2024 |  |  | 11/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022037558 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was on an impella 2.5 for mechanical circulatory support. during insertion, the staff noted bleeding at the peel away sheath valve and the patient's mean arterial pressure dropped to 48. the sheath was exchanged, fluid was administered, and the patient's mean arterial pressure returned to 83. additionally, the placement signal lumen was blocked, the staff tried to aspirate at the red luer, however there was still no signal. unknown how this issue was resolved. no additional information is available. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding - major has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14709 | 1220648-2024-23019 | 20671072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671072&pc=OZD | 12/22/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 12/22/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022135740 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that the perclose procedure was completed, however there was a slight bleeding. to resolve this issue, an 8fr angioseal was deployed leaving the groin dry and soft. it was reported that the procedure was successful. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14710 | 1220648-2024-23020 | 20671099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671099&pc=OZD | 12/29/2021 | 03/17/2025 |  | 11/13/2024 |  |  | 12/29/2021 | 11/13/2024 |  |  | 10/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021010342 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemothorax | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. it was reported the patient developed hemothorax while on support. the resolution for this issue is unknown. the patient's status at normal explant was survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14711 | 1220648-2024-23021 | 20671106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671106&pc=OZD | 01/04/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 01/04/2022 | 11/13/2024 |  |  | 10/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2021015832 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the patient experienced bleeding at all puncture sites and mouth. heparin was reduced to try and control the bleeding. it was reported that the bleeding was under control. the device was normally weaned and the patient outcome at explant is survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14712 | 1220648-2024-23022 | 20671147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671147&pc=OZD | 04/30/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 04/30/2022 | 11/13/2024 |  |  | 03/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022060535 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the patient bleeding at the groin and heparin was withheld. the device was normally weaned and the patient outcome at explant is survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14713 | 1220648-2024-23023 | 20671153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671153&pc=OZD | 05/02/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/02/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109059 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was on an impella 2.5 for mechanical circulatory support. it was reported that upon insertion of the catheter, the catheter "buckled/bent". the impella 2.5 was explanted and replaced with an impella cp. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14714 | 1220648-2024-23024 | 20671164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671164&pc=OZD | 05/04/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/04/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022148694 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. positioning issues were reported, during positioning the sheath got pulled back, possibly out of the vessel so decision was made to remove peel-away and advance the repositioning sheath. it was reported that the procedure was successful with the device and the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14715 | 1220648-2024-23025 | 20671176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671176&pc=OZD | 05/05/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/05/2022 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022090462 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. it was reported that the patient's vasculature was heavily calcified and there was stenosis in the right and left iliac arteries. the right iliac was occluded. the physician attempted to pass the impella 2.5 through the 7mm left iliac stent but couldn¿t advance due to anatomy. the physician attempted device delivery several times but was successful. the impella was removed and replaced with an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14716 | 1220648-2024-23026 | 20671182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671182&pc=OZD | 05/09/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/09/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022142767 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during impella support, access to the left femoral artery noted disease in the iliac. the medical staff continued with impella implantation. it was reported that the impella 2.5 was unable to be advanced due to the patient's vasculature. ischemia was observed in the left femoral artery. this issue resulted in a surgical cutdown, and removal of the impella device. the procedural outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14717 | 1220648-2024-23028 | 20671214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671214&pc=OZD | 08/15/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 08/15/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022100945 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 3 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with heart failure was implanted with an impella 2.5 device for mechanical circulatory support. while on support continuous veno-venous hemofiltration, beginning of kidney failure was noted. pump was explanted and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14718 | 1220648-2024-23027 | 20671219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671219&pc=OZD | 05/18/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/18/2022 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078174 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the pump stopped. it was reported that the procedure was successful and the patient stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14719 | 1220648-2024-23029 | 20671226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671226&pc=OZD | 06/23/2021 | 03/17/2025 |  | 11/13/2024 |  |  | 06/23/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078272 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the patient experienced thrombus. the resolution for this issue is unknown. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14720 | 1220648-2024-23030 | 20671233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671233&pc=OZD | 06/27/2021 | 03/17/2025 |  | 11/13/2024 |  |  | 06/27/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072099 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the patient experienced hemolysis. the resolution for this issue is unknown. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14721 | 1220648-2024-23031 | 20671246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671246&pc=OZD | 06/30/2021 | 03/14/2025 |  | 11/13/2024 |  |  | 06/30/2021 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be on an impella 2.5 for mechanical circulatory support. during delivery, impella insertion was aborted due to the patient anatomy. the device was unable to pass through the patient's anatomy. no further information was provided. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | D4: the serial number, lot number, and udi is unknown. the impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14722 | 1220648-2024-23032 | 20671274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671274&pc=OZD | 05/22/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 05/22/2022 | 11/13/2024 |  |  | 10/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170316 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the staff noted gross amounts of blood in the urine. an echocardiogram was performed and the staff discovered that the pump was caught in the mitral apparatus. the physician tried to reposition the device without success. after multiple attempts, the physician decided to explant the device and utilize an intra-aortic balloon pump. no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14723 | 1220648-2024-23033 | 20671306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671306&pc=OZD | 06/17/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 06/17/2022 | 11/13/2024 |  |  | 12/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022046756 |  | 005042-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the impella pump stopped. the pump was able to be restarted, and the procedure was successful with no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14724 | 1220648-2024-23034 | 20671313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20671313&pc=OZD | 06/20/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 06/20/2022 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022087124 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. it was reported that during delivery the table j wire was meeting some resistance when inserted into the right femoral artery. the physician called for a glide advantage wire, the wire was inserted and would not advance into the descending aorta. the impella cp implant was aborted. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14725 | 1220648-2024-22690 | 20672496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672496&pc=OZD | 06/17/2021 | 11/08/2024 |  | 11/13/2024 |  |  | 06/24/2021 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1131454 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) was being prepped and the system experienced a purge system failure alarm and air in system alarm on the aic. troubleshooting measures did not resolve the issue. this aic was replaced with another aic and successfully support was initiated for the procedure. there was no patient harm or injury reported. | [5 empty] | The investigation into the pump not detected issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs indicated the purge cassette was inserted, and the console began purge fluid detection, almost immediately after this, the logs indicate a piston block and purge system failure (piston blocked). shortly after the pump was connected to the console, but the console did not detect purge fluid present, and several log item lines indicated that no purge fluid was present. the returned console purge assembly and stepper motor were inspected, very little dextrose build up was found and was cleaned. the console was left running on a pump for over 24hrs without issue. the complaint cause in the field of being unable to purge is due to staff beginning to prep without purge fluid in the line due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
14726 | 1220648-2024-23035 | 20672559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672559&pc=OZD | 09/13/2022 | 03/10/2025 |  | 11/13/2024 |  |  | 09/13/2022 | 11/13/2024 |  |  | 09/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021017968 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during support, the patient developed limb ischemia. the physician decided to remove the impella 2.5, it was reported that the patient was stable. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia ¿ reversible issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14727 | 1220648-2024-23036 | 20672571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672571&pc=OZD | 09/29/2022 | 03/07/2025 |  | 11/13/2024 |  |  | 09/29/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128994 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. during insertion, the sheath kinked and was replaced without any reported issues. it was reported that the patient was stable. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14728 | 1220648-2024-23037 | 20672587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672587&pc=OZD | 10/07/2022 | 03/10/2025 |  | 11/13/2024 |  |  | 10/07/2022 | 11/13/2024 |  |  | 11/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023240241 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. during delivery of the device, the pump was unable to pass through the patient's vasculature. the device was removed and delivery aborted. no additional information is available. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14729 | 1220648-2024-23038 | 20672607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672607&pc=OZD | 10/20/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 10/20/2022 | 11/13/2024 |  |  | 11/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023202980 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an impella 2.5 for mechanical circulatory support. post removal of peel away sheath significant bleeding was noted at the groin site. pressure was held and perclose sutures were tightened, however the bleeding was not resolved, and a hematoma was identified. the physician decided to explant the device. it was noted that the patient was stable. the doctor noted that the impella device saved the patient's life, and the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14730 | 1220648-2024-23039 | 20672642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672642&pc=OZD | 11/06/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 11/06/2022 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022101324 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an impella 2.5 for mechanical circulatory support. during support, the physician decided to explant the device due to concerns of limb ischemia. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14731 | 1220648-2024-23040 | 20672969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20672969&pc=OZD | 04/23/2021 | 02/24/2025 |  | 11/13/2024 |  |  | 04/23/2021 | 11/13/2024 |  |  | 10/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1505190 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Septic Shock | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced sepsis. it was reported that the patient was very critical overall situation, and was in both cardiogenic and septic shock with a bag prognosis. the medical team discussed a possible escalation of therapy; however, the patient was too unstable to transfer to another hospital for additional therapies. the cause of the sepsis is unknown, however as the sepsis was discovered after impella was implanted and therefore the impella 2.5 cannot be ruled out as a contributing factor. it was reported there were no issues with the impella 2.5. several days later, it was reported the impella was successfully weaned and removed as the patient's left ventricle performance improved. the patient was reported to still be septic. no further information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported infection/sepsis has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14732 | 1220648-2024-23046 | 20673011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673011&pc=OZD | 05/14/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 05/14/2021 | 11/13/2024 |  |  | 03/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022062093 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during set up, the physician couldn¿t place the impella because of kinking. the impella was explanted with closure with proglide. percutaneous coronary intervention (pci) without impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14733 | 1220648-2024-23043 | 20673038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673038&pc=OZD | 05/05/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 05/05/2021 | 11/13/2024 |  |  | 10/31/2021 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1466452 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, upon access the 2.5 sheath was too short to manage the patient's tortuosity and could not cross the iliac. the short sheath was replaced with a long cp sheath and the device crossed with ease. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14734 | 1220648-2024-23044 | 20673064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673064&pc=OZD | 10/21/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/21/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520007 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with a history of prior coronary artery bypass graft and was suffering from cardiac symptoms at home. the team gave compressions and was attempting to deliver the impella 5.5 to the patient. the impella 5.5 failed to delivered and the patient expired. the team had been attempting to save the patient for 50 minutes when the time of death was called. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella inserted but couldn¿t advance past brachiocephalic as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14735 | 1220648-2024-23047 | 20673073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673073&pc=OZD | 12/30/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 12/30/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022135748 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an impella 2.5 for mechanical circulatory support. during support, it was reported that pulses were absent following insertion of the reposition sheath. the resolution for this issue is unknown. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14736 | 1220648-2024-23045 | 20673095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673095&pc=OZD | 09/20/2024 | 11/13/2024 |  | 11/13/2024 |  |  | 10/14/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085222 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a cp placed for support after the 5.5 failed to deliver. the cp was supporting and had loss of optical signal. no harm or injury noted and console was replaced. this was seen on day 2 of the support. the team replaced the automated impella controller (aic) and transferred the pump to a new aic. later on day 3 of the support the patient experienced ventricular tachycardia (vt) that prompted shock x3. the pump additionally had position adjusted within the left ventricle. later on day 6 the patient had more vt runs x6. the patient this time was not shocked. the patient later expired on day 7 after care was withdrawn. | [5 empty] | The impella device was returned and evaluated. the root cause of this issue was an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. this is one of two medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: november 12, 2024.
14737 | 1220648-2024-23048 | 20673105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673105&pc=OZD | 06/24/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 06/24/2022 | 11/13/2024 |  |  | 11/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210749 |  | 005042-RO |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an attempt to gain access to the right femoral artery was unsuccessful. the patient's anatomy would not allow for the final placement of the impella to be achieved. the device was explanted, and the procedural outcome was unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14738 | 1220648-2024-23049 | 20673260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673260&pc=OZD | 08/06/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/06/2021 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022090429 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was significant oozing noted from extracorporeal membrane oxygenation cannula sites, and gastrointestinal bleeding was also noted. the purge solution was switched to bicarbonate with systemic heparin still infusing. blood was also administered to the patient due to drop in hemoglobin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14739 | 1220648-2024-23050 | 20673262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673262&pc=OZD | 10/21/2024 | 12/20/2024 |  | 11/13/2024 |  |  | 10/21/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558897 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was plugged in and did not zero correctly. after priming, the pressure read 20/19. the cable was unplugged and re-plugged and still did not zero. it was recommended to replace the impella. the impella was removed and the surgical team opted to place an intra-aortic balloon pump for patient support. there was no patient harm. | [5 empty] | The investigation for the priming issues has been completed. clinical details noted that after pump was primed and connected to console, the pump would not zero and placement signal was reading at 20/19. the pump was able to complete priming and the clinical team reported an issue with the placement signal. the impella 5.5 was returned for evaluation. no visual abnormalities noted. optical parameters of returned pump checked and all were found to be within range. pump was connected to console and no optical alarms were triggered. placement signal was reading at -8/-8 while pump was connected to console, pump was not in a bucket of water and was dry on the lab table. data logs from field were reviewed and rt log when pump was first connected to console shows no issue with pump optical parameters or hard failures of the optical sensor. pump was unplugged and reconnected to console and again all optical parameters were within range. as it was noted in the clinical details, the pump had just completed priming but had not been inserted in the patient when placement signal was reading 20/19. no problem was found with the pump in the field as pump was outside of the body when reading 20/19 placement signal and no issues were found on the returned pump. corrections to the initial mfr report: g1 mdr reporting contact email | The impella pump and dta stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14740 | 1220648-2024-23051 | 20673381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673381&pc=OZD | 10/16/2024 | 12/06/2024 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 32-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had pink urine while on support. the pump was removed due to suspected hemolysis. labs did not confirm hemolysis, plasma free hemoglobin not utilized for decision, urine pink was reason hemolysis concern. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. no product returned for analysis. the data logs show no motor current spikes or positioning issues. the placement signal was trending downwards after which there was an upward trend for the rest of the procedure. onset of hemolysis is unclear. the root cause of hemolysis was not determined as insufficient clinical information was provided.
14741 | 1220648-2024-23052 | 20673488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673488&pc=OZD | 10/20/2024 | 12/24/2024 |  | 11/13/2024 |  |  | 10/20/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481538 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported a patient was on impella cp support. while on support, a pump stop occurred. the pump was explanted. no patient harm has been reported. | [5 empty] | D.6b the date of explant is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23052 in accordance with updated procedures.
14742 | 1220648-2024-23053 | 20673490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673490&pc=OZD | 10/16/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 10/16/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022081734 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, bleeding from multiple points was noted. advised to switch to plain d5 until the bleeding was under control before adding heparin back on the purge. air in purge tubing, and impella slightly out of position was also noted. the nurse was provided with de-airing procedure to flush the tubing. no resolution provided for positioning issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14743 | 1220648-2024-23057 | 20673564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673564&pc=OZD | 10/18/2024 | 01/02/2025 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025459147 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Pericardial Effusion | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and there was high purge pressure. lysis therapy was initiated. additionally, the patient had a pericardial effusion while on support. it is unknown if the pericardial effusion is impella related. furthermore, the patient expired after 184.82 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿. | The investigation the high purge pressure and vascular damage has been completed. product was returned. there was no biomaterial noted around the impeller or on the purge gap. there was a kink on the catheter right before the motor housing. the pump was purged and purge pressure slowly increased over 1000 mmhg. the catheter was cut in between the kink and the motor housing and the purge pressure instantly resolved. field log review showed that the high purge pressure resolved after ~8 hours and lysis therapy was used. the root cause of the high purge pressure was most likely biomaterial. without additional clinical details, the root cause of the vascular damage was not determined since insufficient clinical details were provided. added h6 impact code 4644 corrections to the initial mfr report: g1 mdr reporting contact email.
14744 | 1220648-2024-23056 | 20673593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673593&pc=OZD | 10/18/2024 | 11/15/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537992 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to multi-vessel disease. while on support, echocardiography showed clot in left ventricle and descending aorta attached to impella catheter, per anesthesia. the cp was then replaced/escalated to an impella 5.5. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. the root cause of the low pump flow issue was not determined since product was not returned and data logs were not returned as well. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for low pump flow and thrombus has been completed. pump was not returned for investigation. data logs were not returned as well. since insufficient clinical information was provided and product not returned, the root cause of the access site bleeding issue was not determined. the root cause of the low pump flow issue was not determined since product was not returned and data logs were not returned as well.
14745 | 1220648-2024-23054 | 20673601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673601&pc=OZD | 10/17/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535624 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, purge flows continued to decrease. high motor current and pump stop occurred. once chest was opened for explant, doctor was able to walk out the impella, graft was full of clot. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The product was returned for investigation. no physical abnormalities were reported. biomaterial was observed on the purge gap. high purge pressure and pump stop were reproduced during the case. the pump was cut near the motor to observe the blockage in the purge line. purge was able to come out from the cut catheter. no further investigation is required for this case. data logs show an 8-minute gradual rise in purge pressure, followed by the saturated pump flow and pump stop, with several motor current high-pump stop alarms. as per the given clinical data, the pump stop occurred shortly after noted low purge flow and high purge pressure. the pump was successfully explanted, and during removal, some clot was reported attached to the graft. the cause of the pump stop issue was biomaterial, based on the data log review and investigated data logs. the saturated pump flow failure was removed and replaced with pump stop since the cause of the issue was addressed under the pump stop failure mode with sev 5. the same biomaterial could cause the high purge pressure. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23054. b1 adverse event was inadvertently selected, product problem only b2 should have been left blank e4 should have been left blank g1 reporting contact email h1 type of reportable event corrected h6 code 4440, 4627, and 2964 were reported incorrectly. new code was added to health effect - clinical code, health effect - impact code, and medical device problem code. h10 instructions for use missing.
14746 | 1220648-2024-23064 | 20673605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673605&pc=PYX | 10/20/2024 | 12/11/2024 |  | 11/13/2024 |  |  | 10/20/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025528976 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. post removal of impella peel away sheath it was noted that flows of impella rp flex significantly dropped to 2l/min, it was also seen that placement signal had spiked and motor current had increased, purge pressures elevated, and purge flows decreased. tech checked purge line tubing for kinks and did not observe any. the physician repositioned impella rp flex looking for kinks on catheter and did not observe any. the physician had multiple attempts pulling the catheter back and advancing catheter to new positions but did not observe any changes with impella flow. fluids were given in the cath lab and it did not improve flows. impella rp flex was removed and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. the data logs confirmed the clinical details provided. there was a spike in motor current with no change in p-level and a spike in placement signal with a decrease in pump flow. purge pressure began to rise shortly after. the root cause of the low pump flow is most likely due to ingested biomaterial. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23064. e4 should have been left blank. g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
14747 | 1220648-2024-23055 | 20673609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673609&pc=OZD | 10/17/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550604 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient had short runs of ventricular tachycardia (vt), amio gtt was decreased and epinephrine was increased from 0.02 to 0.04. echocardiography showed impella was slightly deep; withdrawn 1cm. better positioning. ectopy minimal. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14748 | 1220648-2024-23060 | 20673621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673621&pc=OZD | 10/19/2024 | 12/16/2024 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501582 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that upon trying to stand the patient from bed to chair, the patient fell back on the bed after partially standing, resulting in kinking of the impella. the kink was between the insertion site and the first securement of the 3-point fixation. restarting the impella was attempted but unsuccessful. automated impella controller transfer attempted but pump did not restart. the patient was taken to the operation room for 5.5 explantation. the patient remained stable and survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into pump stop has been completed. the data logs showed that there was an electrical pump stop. the pump was looked at and there was a catheter kink and open lines. the red handle was opened and the motor lines were necked on the catheter side. the root cause of the pump stop is due to open electrical lines that most likely occurred when the patient fell and pulled the catheter. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23060.
14749 | 1220648-2024-23063 | 20673629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673629&pc=OZD | 10/19/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519330 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed a left ventricle thrombus and impella was explanted. the patient is stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email.
14750 | 1220648-2024-23065 | 20673658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673658&pc=OZD | 10/20/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/20/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486098 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had ventricular fibrillation. the patient was shocked a couple times. the issue reoccurred and pump was explanted shortly after. it is unknown if the explant was related to the patient's arrhythmia issues. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf could not be determined. corrections to the initial mfr report: d4 model number. g1 mdr reporting contact email.
14751 | 1220648-2024-23059 | 20673664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673664&pc=OZD | 10/19/2024 | 12/17/2024 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482948 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump implanted via the femoral artery to support a patient admitted in for a st-segment elevation myocardial infarction. during initial implant the physician was unable to push sterile sleeve and repositioning sheath all the way back, sterile sleeve was stuck on catheter. he had to remove the peel away sheath to advance the catheter all the way into the left ventricle. it was also reported that the leg became ischemic and the pump was removed and escalated to the impella 5.5 via the axillary after 14hours. the team additionally had to perform a fasciotomy. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the limb ischemia, product damage, and vascular damage has been completed. data logs are not relevant to the complication. returned product was investigated and there were no signs of ripping or excessive force on the sterile sleeve, though it had been detached from the hub on the repositioning sheath side. it was confirmed on other cp pumps that the sterile sleeve could be pulled off of the hub without tearing. though it was reported that the sleeve and/or sheath were stuck and couldn't be moved, both were able to slide up and down the catheter easily on returned product. based on clinical details and returned product, the product damage was determined to be a use issue. according to clinical details, vascular damage was at the access site, its investigation will also cover the investigation into the reported access site bleed noted when the patient got back to the icu. data logs are not relevant to the complication. returned product was investigated and there were no visual abnormalities. clinical details provide very limited detail as to what could have caused the damage or when it occurred, as it was only noted once the pump was explanted. since limited details were provided and there were no abnormalities with returned product, the root cause of the vascular damage could not be determined. the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14752 | 1220648-2024-23058 | 20673670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673670&pc=OZD | 10/18/2024 | 12/19/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515310 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, the patient developed slow ventricular tachycardia (vt) when attempting to place pa catheter. patient sustained slow vt resulting in need for cardioversion and vt resolved when swan was removed and a single 200 j cardioversion. upon assessment of right leg extremity, catheterization lab staff was unable to obtain doppler pedal pulses and reported foot and leg to be cold. an external femoral bypass was discussed but ultimately not performed. the patient survived the event. | The extent of machines and pressors to keep her alive was discussed with the family by the physician and the decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation vf/vt/af/at and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-23058. b2 death added, date of death is unknown. b5 updated to include the patient outcome of .death e4 should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. h1 type of reportable event. h6 health effect - impact code 1802 death.
14753 | 1220648-2024-23062 | 20673675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673675&pc=OZD | 10/19/2024 | 12/17/2024 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532868 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemolysis | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to cardiogenic shock. while on support, there were concerns about pump saturation which seemed to have resolved on its own. it was also reported that the patient had developed hemolysis. there was no mention of intervention for hemolysis but the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and saturated flow has been completed. no product was returned for evaluation. data logs were reviewed and real time (rt) logs show motor current and pump flow step up with no change in p-level, pump flow also becomes saturated at that time. purge pressure began to rise right at time of motor current step up and rise to high purge pressure alarms after an 8-minute rise with a drop in purge flows. pump was turned down to p-0 then back to p-5 and purge pressure and flow began to return to normal values after. additionally, once purge pressure returns to normal range, motor current steps down and pump flows are no longer saturated. no persistent suction alarms or positioning issue alarms were noted in the lt logs. clinical details noted that pump flow was saturated in the field and high purge pressure alarms were being triggered. additionally, it was noted that later on in support after high purge pressure resolved, patient had tinged urine but labs were normal and no intervention was performed. the root cause of the saturated flow and high purge pressure was most likely biomaterial ingestion. the root cause of the hemolysis cannot be definitively determined as limited clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email
14754 | 1220648-2024-23061 | 20673692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673692&pc=OZD | 10/19/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 10/19/2022 | 11/13/2024 |  |  | 11/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC |  |  | US |  | 2023217467 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon plugging impella 2.5 into console a yellow alert message came on screen saying catheter defective. the impella 2.5 was replaced. there was no patient harm. | [5 empty] | The investigation into the reported issue has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the cause of the issue was not determined since the product was not returned and the clinical details provided were insufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
14755 | 1220648-2024-23066 | 20673708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673708&pc=OZD | 10/20/2024 | 12/16/2024 |  | 11/13/2024 |  |  | 10/20/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515184 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient suffered from an ischemic infarction during impella 5.5 support. the patient's neuro status was monitored and support with the implanted pump continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-23066.
14756 | 1220648-2024-23067 | 20673729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673729&pc=OZD | 10/15/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 10/15/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022112623 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, a low purge pressure alarm was noted. all connections were tight and dry. attempted to reinsert the purge cassette and increased purge fluid to d10, with no success. flows were maintained 2.3-2.5 lpm for the entirety of the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14757 | 1220648-2024-23068 | 20673751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673751&pc=OZD | 10/04/2024 | 12/18/2024 |  | 11/13/2024 |  |  | 10/04/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump stopped unexpectedly during patient support. the nursing team was able to restart the pump at a lower level until a replacement pump could be implanted later that day. there was no patient harm resulting from the noted failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: entering cardiac output: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: troubleshooting the purge system : ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message: could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings and cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the temporary pump stop has been completed. the pump was not returned. data logs showed high motor current (mc) pump stop alarm on the last day of support. there was a mc spike with an increase in purge pressure and drop in purge flow but no purge alarms were triggered. the pump was able to be restarted at p-2 and the purge also was returned to normal. the pump continued to run with no other issues for two hours before explant. the root cause of the temporary pump stop was most likely biomaterial. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23068 was submitted.
14758 | 1220648-2024-23070 | 20673811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673811&pc=OZD | 10/17/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526663 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received voluntary medwatch report # mw5161762 providing additional details. the medwatch states that while attempting to remove the device, the patient lost a significant amount of blood. the patient received a unit of packed red blood cells along with manual pressure. during the emergency surgery to repair the femoral artery, the patient was placed in intra aortic balloon pump in the right femoral artery. the patient was eventually weaned and survived the explant. | The complainant reported the 65-year-old male patient was on impella cp support and upon placement the impella was stuck in the mitral apparatus. when the doctor pulled the pump back, the pigtail with the inlet cage became detached and it floated down to the left femoral artery. several attempts to snare it were made but ultimately the decision was made to take the patient to surgery for a femoral cutdown, foreign body removal and urgent upgrade to an impella 5.5 via the right axillary artery. |  |  |  |  | The investigation for the positioning issue and the product damage has been completed. as per clinical, the pigtail was stuck in mitral apparatus and was detached but without the pump, the complaint cannot be investigated. therefore, the root cause of the product damage issue was not determined since product was not returned for investigation. clinical details also state that at the time of the initial placement, the pump was stuck in the mitral apparatus. the patient had mitral regurgitation (mr) which was repaired. the cause of the positioning issue was use related since the pump was stuck in the mitral apparatus upon initial placement. voluntary medwatch #mw5161762 was received and corresponds to this report, initial mfr report #1220648-2024-23070. b5-added details of the event that were not provided in the initial mfr report h6-clinical code 1888 h6 impact code 4643 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14759 | 1220648-2024-23069 | 20673838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673838&pc=OZD | 11/10/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 11/10/2021 | 11/13/2024 |  |  | 12/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022050027 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the red alarm indicating the pump dislocated in the aorta occurred due to being too deep. impella dislocated and repositioning failed and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14760 | 1220648-2024-23071 | 20673904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673904&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/13/2024 |  |  | 10/15/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546173 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 58 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced oozing from their impella access site following impella 5.5 implant. packed red blood cells, fresh frozen plasma, platelets, and cryoprecipitate were given and a pocket washout was scheduled for the impella pocket. support was continued with the implanted pump following the noted intervention. roughly two weeks later, care was withdrawn and the patient passed away. it is unknown if the condition caused or contributed to the patient's passing two weeks later. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-23071 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14761 | 1220648-2024-23072 | 20673930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673930&pc=OZD | 07/06/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 07/06/2021 | 11/13/2024 |  |  | 08/31/2021 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1453454 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 11 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient's leg became ischemic when the impella sheath was inserted. an external percutaneous femoral bypass was performed to manage the complication. as cardiac function was beginning to recover, the impella side, where ischemia in the lower limbs was continuing, was first removed, and the patient was then taken off extracorporeal membrane oxygenation (ecmo). the patient was said to be stable after that. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14762 | 1220648-2024-23073 | 20673952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673952&pc=OZD | 11/16/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 11/16/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022108972 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. during an intervention supported by the impella 2.5, the pump was removed due to suspected limb ischemia. the device was explanted, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14763 | 1220648-2024-23074 | 20673994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20673994&pc=OZD | 08/13/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/13/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022087191 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced sepsis. drug therapy was provided in an effort to manage the complication. the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14764 | 1220648-2024-23076 | 20674066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674066&pc=OZD | 10/19/2024 | 03/24/2025 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516915 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a broken purge filter. the yellow luer lock was leaking. the impella 5.5 was explanted. no patient survived the explant. | [5 empty] | The pump has been returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the product was returned and there was a crack down the side of the luer. it was purged with fluid and fluid expelled from the crack. it did not appear that there was excessive force applied. the root cause of the purge leak - pump was due to a damaged sidearm yellow luer most likely occurring during use due to the high tensions in combination with disinfectants and certain drug ingredients. g1 reporting contact email revised on manufacturer device report 1220648-2024-23076 in accordance with updated procedures.
14765 | 1220648-2024-23077 | 20674089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674089&pc=OZD | 11/18/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 11/18/2022 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022068023 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. during closure of the impella, it was noted that a thrombus in the left leg was obstructing blood flow. a balloon was utilized to restore the flow. the device was explanted, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14766 | 1220648-2024-23078 | 20674174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674174&pc=OZD | 05/18/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 05/18/2021 | 11/13/2024 |  |  | 11/30/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022036742 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was going to be on an impella 2.5 for mechanical circulatory support. it was reported that placement was not possible due to calcification and tortuous anatomy femoral vessels. implantation was aborted, no further information is available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14767 | 1220648-2024-23079 | 20674243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674243&pc=OZD | 10/18/2024 | 12/27/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513021 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during an attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion, it was noted that the implant was not possible due to calcification in the vessel. the implant was aborted and an computed tomography angiography was ordered for all vessels. there was no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, not possible to implant impella due to big calcification as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14768 | 1220648-2024-23080 | 20674406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674406&pc=OZD | 08/13/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/13/2021 | 11/13/2024 |  |  | 02/28/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022078298 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during impella placement, a dissection of the left anterior descending artery was observed. a pericardial effusion was performed and one liter of blood pulled from pericardial effusion¿s, auto transfusion initiated, patient refused blood due to religious reasons. the pump was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14769 | 1220648-2024-23082 | 20674483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674483&pc=OZD | 10/18/2024 | 12/24/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529708 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, suction alarms were present. position in aorta alarm present and resolved by staff. the patient developed sepsis. the physician reported leg ischemia on the right leg where the impella is located. the physician discussed the crossover bypass. eventually, care was withdrawn and the patient expired. patient outcome (demise) cannot be dissociated with the use of pump. the patient had pump related critical limb ischemia and the pump may have been malpositioned and may have caused or contributed to the patient's demise. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: table 3.2 impella catheter components . ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. this is one of two medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: november 13, 2024. | The investigation limb ischemia, low pump flow, and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-23082 in accordance with updated procedures.
14770 | 1220648-2024-23081 | 20674519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674519&pc=OZD | 08/16/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 08/16/2021 | 11/13/2024 |  |  | 10/31/2022 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022027985 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. during an effort to place the impella pump, the device could not be delivered due to issues with the patient's vasculature and the implantation process was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14771 | 1220648-2024-23083 | 20674547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674547&pc=OZD | 10/28/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 10/28/2021 | 11/13/2024 |  |  | 12/31/2021 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1477943 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient scheduled for a balloon aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. while on support, six perclose sutures broke and failed to close the impella access site. the patient experienced bleeding at the impella access site and the patient was taken for surgical cut down with artery repair to close the impella access site. the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14772 | 1220648-2024-23084 | 20674845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674845&pc=OZD | 10/18/2024 | 12/06/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849119 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a broken purge clip. there was no patient involvement. the aic was removed from service. | [5 empty] | The investigation for the console (aic) purge flag/disc issue has been completed. console logs did not contain any data relevant to this complaint. the console was returned for evaluation. the consoles purge sensor was inspected but both the purge disc retainer and flag were intact. after inspection of the rest of the console, it was discovered that pump connector cover was missing/broken and the pump connector was lodged into the console the reported broken purge ¿clip¿ on (b)(6) could not be confirmed. both the retainer and the flag of the aic was still intact. however, it was observed that the pump connector cover was missing and the pump connector was lodged inside the console. it is likely that these damages were what the clinical staff meant to report. the damages to the pump connector cover and pump connector were likely use related. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
14773 | 1220648-2024-23085 | 20674947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20674947&pc=OZD | 10/18/2024 | 12/16/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542533 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp placement was unsuccessful due to the patient having an abdominal aneurysm. the impella wire kept buckling inside the aneurysm. several attempts were made but ultimately impella placement was aborted. the patient was stable and survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through due to patient anatomy. corrections to the initial mfr report: g1 mdr reporting contact email.
14774 | 1220648-2024-23086 | 20675431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20675431&pc=OZD | 10/18/2024 | 12/16/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519315 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 82-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had lower limb ischemia. an antegrade stick to the left superficial femoral artery for an external bypass. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral. vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal. failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14775 | 1220648-2024-23087 | 20675783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20675783&pc=OZD | 10/17/2024 | 01/02/2025 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510286 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported multiple complications were encountered during impella 5.5 support. one day after implant, it was noted that the pump began to experience high purge pressure. best practice was discussed with the site and a lyse protocol was forwarded to the staff to troubleshoot the issue. roughly a week later, the patient was diagnosed with a cerebral hemorrhage. it was documented that heparin had been utilized for support leading up to the cerebral hemorrhage. therapy was discontinued and the patient passed away the following day. | [5 empty] | The investigation for the high purge pressure and stroke/cerebrovascular accident has been completed. no product was returned. log review showed an increase in purge pressure over about 13 hours. the purge pressure leveled out for about two hours before dropping back to normal. logs showed that heparin was added to the purge fluid shortly after the purge pressure decrease began. there were no abnormalities with the purge after that. the root cause of the high purge pressure was most likely biomaterial buildup. the root cause of the stroke could not be determined without additional clinical details. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-23087 was submitted in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | B.2 date of death was incorrectly reported on initial manufacturer device report number 1220648-2024-23087. the date was corrected on follow up report 01 but corrected data information was inadvertently omitted from h.11 on the follow up report 01.
14776 | 1220648-2024-23088 | 20676006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676006&pc=OZD | 12/01/2017 | 12/30/2024 |  | 11/13/2024 |  |  | 10/17/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The abiomed team in japan conducted literature reviews, and one was from a japanese city hospital reviewing their outcomes for 2017-2023 impella cases. the team found that prior to the use of the impella cp pump, the deaths were from circulatory failure when in-house/in-hospital. after the impella cp was utilized, the deaths were attributed to access site and chronic infections. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. literature citation: short-term and long-term outcomes of impella (percutaneous left ventricular assist device) for cardiogenic shock in a (b)(6) hospital (b)(6), et al. journal of the japanese society of internal medicine 2024; vol.113. no.146-. | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the infection could not be determined. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-23088 was submitted in accordance with updated procedures.
14777 | 1220648-2024-23089 | 20676056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676056&pc=OZD | 10/18/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | Ischemia Stroke | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 58x-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient suffered an ischemic brain injury that converted to hemorrhagic. the patient was actively herniating. comfort care was initiated and the patient expired. the relationship between the cause of death and the impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system . section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. g1 mfr site name and address.
14778 | 1220648-2024-23090 | 20676155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676155&pc=OZD | 10/18/2024 | 03/03/2025 |  | 11/13/2024 |  |  | 10/18/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Disability | Death | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 72-year-old female noted as high risk intervention was implanted with an impella 5.5 device for mechanical circulatory support. the patient suffered multiple episodes of ventricular tachycardia overnight during impella 5.5 support. the patient went into torsades de pointes for approximately four minutes and required defibrillation to stabilize. the impella support level was reduced in immediate response to the episode and the device was repositioned via echo guidance. support continued with the implanted pump following the event. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. device history record review confirmed the pump passed all post sterile inspection checks. no device was received for evaluation. to determine the true root cause of the issue sufficient clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the reported issue was not determined. | B2-selected death and added date of death. b5-mso review. h1-changed to death.
14779 | 1220648-2024-23092 | 20676642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676642&pc=OZD | 09/20/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 09/21/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022108339 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. upon implantation the pump failed to start. the team replaced the pump and support proceeded successfully. there was no harm or injury. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field logs were reviewed, and an impella stopped ¿ motor current high alarm was observed on (b)(6) 2021 which was immediately followed by an impella stopped ¿ reverse flow alarm. multiple pump start attempts followed these alarms which were associated with motor current spiking to approximately 1500 ma. all start attempts were unsuccessful. the returned pump was visually inspected, and it was noted that the impeller rotated with resistance upon manipulation by hand. further inspection revealed that the outflow cage was bent, and multiple struts made contact with the blades of the impeller. the root cause of the pump being unable to start was a deformed outflow cage most likely due to excessive force. this report is being filed as part of a retrospective review of historical records.
14780 | 1220648-2024-23093 | 20676661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676661&pc=OZD | 10/15/2024 | 01/14/2025 |  | 11/13/2024 |  |  | 10/15/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025557318 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella cp support ongoing for a patient admitted in with acute myocardial infarction/cardiogenic shock. it was reported that there was an access site bleed that prompted the team to re-dress the site and infuse blood products back to the patient. the pump remained on for support. it was also reported that the pump stopped working eight days later. after automated impella controller (aic) to aic transfer, the pump still stopped and could not be restarted. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the pump stop and access site bleed has been completed. the product was returned for investigation. data log was not provided for this case run. no physical damage was observed on the returned product. all the motor cable lines were within specifications. further investigation revealed that the large impeller flap, which is a sign of bearing wear, was present. the pump stopped on the 10th day of implant. the root cause of the pump stop was bearing wear after 10 days of use, which is beyond the 8-day design life per design trace matrix 0046-9910. as per the clinical details, heavy bleeding was reported on the second day of pump use. the doctor changed the dressing and resolved the issue by applying manual pressure. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14781 | 1220648-2024-23091 | 20676773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20676773&pc=OZD | 10/19/2024 | 12/13/2024 |  | 11/13/2024 |  |  | 10/19/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466790 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 81-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the patient suffered a large right sided stroke during impella cp support. the patient was able to move their extremities on the right side and was reported to be stable. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23091.
14782 | 1220648-2024-23096 | 20677102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677102&pc=OZD | 10/20/2024 | 12/17/2024 |  | 11/13/2024 |  |  | 10/21/2024 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548334 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp placement the .018 guidewire was difficult to remove and frayed after removal. the physician felt cp may still have wire near the motor, so decision was made to replace impella with a new cp. no patient harm has been repoted. | [5 empty] | The impella device has been received and an evaluation will be taking place. upon completion of the investigation, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. returned compliant pump visually inspected. no catheter/cannula kink observed. no broken blade and no damage observed. no guidewire piece found in side the cannula or in the pump. the cause of the delivery issue cannot be determined since the compliant 0.018' gw not returned for analysis and lack of details. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-23096
14783 | 1220648-2024-23094 | 20677130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677130&pc=OZD | 09/22/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 09/23/2021 | 11/13/2024 |  |  | 05/31/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080592 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 2.5 device for mechanical circulatory support. upon implantation, the pump did not start and another impella 2.5 was used. there was no adverse event or harm to the patient. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no abnormalities were found and that it successfully started in the lab. the cause of the pump unable to start was not determined as it could not be reproduced. this report is being filed as part of a retrospective review of historical records.
14784 | 1220648-2024-23095 | 20677164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677164&pc=OZD | 09/22/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 09/24/2021 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella 2.5 device for mechanical circulatory support. during prep, the physician and scrub noticed that the 2.5 pump had a damaged cage, a cp pump was used, and there was no harm to the patient. | [5 empty] | D4 serial number, catalog number, lot number, expiration date, and unique identifier (udi) #: unknown h4 device manufacture date: unknown. the investigation for the reported product damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine that the cause of the pump damage was a damaged inflow cage. this report is being filed as part of a retrospective review of historical records.
14785 | 1220648-2024-23099 | 20677227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677227&pc=OZD | 05/06/2023 | 11/13/2024 |  | 11/13/2024 |  |  | 05/06/2023 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022152469 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. during support, the patient developed large hematoma in right groin at access site, therefore, the physician decided to remove impella, sheath and close arteriotomy with previously deployed perclose suture. patient condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the access site bleed/hematoma has been completed. the cause of the issue was not determined since the product was not returned, and clinical details provided were insufficient to determine the root cause. this report is being filed as part of a retrospective review of historical records.
14786 | 1220648-2024-23098 | 20677278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677278&pc=OZD | 11/22/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 11/22/2021 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues that were resolved by changing to another console. no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was not confirmed. failure mode is likely due to misuse of the pump and automated impella controller (aic). after testing results were observed to be very similar to complaint data, it can likely be concluded that staff did not fully snap the pt cable into the blue optical connector in the aic. the placement signal issue was use related (ex. reading the waveforms or metrics). this report is being filed as part of a retrospective review of historical records.
14787 | 1220648-2024-23097 | 20677315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677315&pc=OZD | 09/23/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 09/23/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109034 |  | 005042 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. upon implantation, the pump was unable to be implanted due to large calcifications in the right femoral artery (rfa). the decision was made to abort impella and insert a new intra-aortic balloon pump (iabp). | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned, and the issue was not confirmed. the returned pump was visually inspected, and no abnormalities were noted. it was reported that the patient¿s left and right femoral arteries were calcified. per the results of the clinical review and investigation, the root cause was determined to be related to patient condition. this report is being filed as part of a retrospective review of historical records.
14788 | 1220648-2024-23100 | 20677406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677406&pc=OZD | 06/30/2023 | 11/13/2024 |  | 11/13/2024 |  |  | 06/30/2023 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5-EU |  |  | OZD |  |  | IMPELLA 2.5-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022126121 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella 2.5 device for mechanical circulatory support. during implant procedure, physician was having trouble inserting the impella due to the patient's anatomy. as a result, the procedure was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the evaluation into the delivery issue has been completed. as per clinical review, the patient had small vessels in the right femoral artery and 13 french introducer sheath was unable to be inserted. the cause of difficulty to insert introducer issue was patient condition related due to patient having small vessels. this report is being filed as part of a retrospective review of historical records.
14789 | 1220648-2024-23102 | 20677489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677489&pc=OZD | 10/16/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 10/21/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104314 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the flow was too low for the chosen p level. the system couldn¿t provide more than 1.3l/min at p6 despite several attempts of repositioning and optimization in a very experienced center. the device was explanted and replaced by another impella 2.5 catheter that performed a normal support. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the product was returned, and the issue was confirmed. the data logs show a significant drop in flow and mc pulsatility 5 hours into support. flow remains low until the pump is explanted. the product was returned and biomaterial was found wrapped around the impeller and in the outflow. the pump was run in a bucket of water and the low flow and suction were reproduced. the pump was then soaked in tetraglyme and run again. the pump was able to operate within specification once the biomaterial was removed. per the results of the clinical review and investigation, the root cause was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
14790 | 1220648-2024-23103 | 20677504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677504&pc=OZD | 11/25/2021 | 11/13/2024 |  | 11/13/2024 |  |  | 11/25/2021 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022114807 |  | 005040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | Yes | Defective Component | Deformation Due to Compressive Stress | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was resistance while advancing the sheath beneath the skin-level. after removing the dilator from the sheath, a strong bleeding form the sheath was visible. 13 f sheath was cracked open at the proximal portion (close to the hemostatic valve). the dilator was re-inserted to stop the bleeding, and the sheath was replaced by a new sheath. no further issues with the new sheath or procedure. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the product was returned, and the issue was confirmed. the returned introducer was visually inspected and a scoreline split was observed near the hub. per the results of the clinical review and investigation, the root cause was determined to be most likely an introducer sheath split along the scorelines this report is being filed as part of a retrospective review of historical records.
14791 | 1220648-2024-23101 | 20677511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677511&pc=OZD | 02/09/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 02/09/2022 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239165 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by rebooting the console. the patient ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - purge issue - other has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the aic - purge issue - other could not be determined. this report is being filed as part of a retrospective review of historical records.
14792 | 1220648-2024-23105 | 20677888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677888&pc=OZD | 02/23/2022 | 04/22/2025 |  | 11/13/2024 |  |  | 02/23/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022110415 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. during implantation the procedure was aborted due to the patient's anatomy. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
14793 | 1220648-2024-23106 | 20677961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677961&pc=OZD | 04/05/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 04/05/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139228 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. while on support hemolysis was noted and volume was given. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14794 | 1220648-2024-23104 | 20677993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20677993&pc=OZD | 11/09/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 11/09/2022 | 11/13/2024 |  |  | 11/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237153 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced limb ischemia which was noted as a loss of pulses in the patient's left leg. the patient was successfully weaned and the impella device was explanted the same day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14795 | 1220648-2024-23107 | 20678104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20678104&pc=OZD | 04/05/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 04/05/2022 | 11/13/2024 |  |  | 06/30/2023 |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125822 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that after removal of the impella 2.5 they were unable to deploy perclose. manual pressure was held for hemostasis and protamine given for reversal. the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
14796 | 1220648-2024-23108 | 20679245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20679245&pc=OZD | 04/21/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 04/21/2022 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145436 |  | 0042-0000 |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support and during mid procedure, the automated impella controller experienced an error message and loss of placement signal tracing. the console was replaced which resolved the issue. the patient was later determined not to be a candidate for advanced therapies, care was withdrawn, and the patient expired. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore an evaluation of the device was not possible. event history logs were obtained and reviewed. the reported controller error alarm and the reported optical placement signal was no longer available was confirmed. the root cause of the automated impella controller issue was a software issue. this report is being filed as part of a retrospective review of historical records.
14797 | 1220648-2024-23109 | 20679378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20679378&pc=OZD | 02/23/2023 | 10/21/2025 |  | 11/13/2024 |  |  | 02/23/2023 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371242 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for a high risk percutaneous coronary intervention procedure, the automated impella controller alarmed for "battery fault". abiomed supported provided guidance over the phone for the procedure of activating the battery switch underneath the console. after pressing the switch, the error message disappeared. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14798 | 1220648-2024-23110 | 20679405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20679405&pc=OZD | 09/20/2022 | 11/13/2024 |  | 11/13/2024 |  |  | 09/20/2022 | 11/13/2024 |  |  |  |  | 11/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1296873 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. after the pump was explanted, the aic experienced a shutdown. no further information was provided concerning the issue or resolution. additional information was provided that noted the patient expired with no allegation against the aic or impella related to the patient¿s death. | [5 empty] | The investigation for the reported aic - shutdown issue has been completed. the product was returned and the issue was confirmed. the root cause the issue was an eeprom programming issue. this report is being filed as part of a retrospective review of historical records.
14799 | 1220648-2024-22900 | 20662056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662056&pc=OZD | 02/02/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 02/02/2022 | 11/12/2024 |  |  | 04/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072513 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during efforts to deliver the impella, it was stated that the medical staff was unable to advance the pump due to heavy calcification of both groin access sites, stenosis of the right femoral and iliac arteries. the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14800 | 1220648-2024-22901 | 20662095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662095&pc=OZD | 02/03/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 02/03/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022108340 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during efforts to deliver the impella, the pump could not be advanced. the procedural outcome was the process was aborted. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14801 | 1220648-2024-22902 | 20662123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662123&pc=OZD | 02/03/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 02/03/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022131168 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support patient threw a clot to proximal left anterior descending artery causing st elevation. penumbra was used to retrieve the clot. the impella was weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14802 | 1220648-2024-22904 | 20662208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662208&pc=OZD | 02/21/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 02/21/2022 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032513 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the medical staff placed the impella pigtail catheter and the guidewire in the left ventricle. however, during insertion, severe calcification was observed. the decision was made to remove the impella pump and replace the device with an intra-aortic balloon pump, followed up with an emergent coronary artery bypass graft. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14803 | 1220648-2024-22903 | 20662209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662209&pc=OZD | 02/04/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 02/04/2022 | 11/12/2024 |  |  | 11/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022034401 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the cardiac physician attempted to cross the valve with a mp and j-wire: however, the cardiac surgeon experienced difficulty with advancing. after fluro, it was noted that the wire was tracking sub-intimal. the physician pulled back the wire and attempted to advance but was unable to advance past the dissection. the physician consulted with surgery who suggested pulling the wire and sheath out, and the case was aborted. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14804 | 1220648-2024-22905 | 20662218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662218&pc=OZD | 03/08/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 03/08/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022131224 |  | 005042 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, heparin was removed from the purge system due to the patient having heparin induced thrombocytopenia. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14805 | 1220648-2024-22899 | 20662502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662502&pc=OZD | 11/10/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/10/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022081733 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, there was suction at p5 with low lows, which resolved at p3. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14806 | 1220648-2024-22907 | 20662646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662646&pc=OZD | 10/16/2024 | 12/11/2024 |  | 11/12/2024 |  |  | 10/16/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547410 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 83-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella support, there was a lack of distal pulse in the impella leg. the sheath was repositioned and the patient was transferred to a hospital for a distal perfusion cannula. the patient expired after 9.31 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14807 | 1220648-2024-22906 | 20662769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662769&pc=PYX | 10/16/2024 | 12/05/2024 |  | 11/12/2024 |  |  | 10/16/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025484923 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella rp device for mechanical circulatory support. during impella support, the pump was explanted however after 9 days with concerns of hemolysis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. data logs were reviewed which showed pump ran without issue during the case. there were suction alarms triggered during support but resolved. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
14808 | 1220648-2024-22909 | 20662868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662868&pc=OZD | 03/13/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 03/13/2022 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022021627 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, heparin was removed from the purge system due to the patient¿s decreased platelets. weaning was successful, the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14809 | 1220648-2024-22910 | 20662873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662873&pc=OZD | 10/17/2024 | 08/22/2025 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329223 |  | 0042-0000-AU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella automated impella controller (aic) had a controller error alarm. the aic was exchanged. there was no harm to the patient reported. | [5 empty] | The automated impella controller was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised on manufacturer device report in accordance with updated procedures. h10 additional manufacture narrative revised on manufacturer device report in accordance with updated procedures.
14810 | 1220648-2024-22908 | 20662896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662896&pc=OZD | 10/16/2024 | 12/11/2024 |  | 11/12/2024 |  |  | 10/16/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541970 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant, had placed an impella cp to support a high risk coronary intervention 73-year-old female patient. hemostasis was complicated by bleeding and sign of perforation at the femoral. there was contrast blushing. this prompted, the team to treat with a covered stent. | [5 empty] | The impella device was discarded by the customer. and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the vascular damage has been completed. the product was not returned for investigation. a data log review was not required for this case run. according to the clinical team, a surgical approach was needed to resolve the access site bleeding complication. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was no provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14811 | 1220648-2024-22914 | 20662934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662934&pc=OZD | 10/17/2024 | 12/12/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524789 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the impella flows were lower than expected. the level was turned down to p5 with suction alarm, turned to p4 and resolved. tte obtained showing underfilled left ventricle (lv) and hyper dynamic right ventricle (rv). the physician decided to remove the impella. post case angiogram showed trace of aortic valve insufficiency (ai) and abnormal contrast to image in left annulus. | [5 empty] | The investigation for the heart valve damage has been completed. the pump was not returned for investigation. a data log review was not required for this case run. as per the clinician, the doctor was able to manipulate the catheter and successfully cross the lv. resistance met while crossing lv. the left annulus was diagnosed with abnormal contrast, and there was evidence of aortic insufficiency. the cause of the heart valve damage issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
14812 | 1220648-2024-22912 | 20662948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20662948&pc=OZD | 10/17/2024 | 12/12/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  | 03/31/2025 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024319112 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 40-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, there was a high purge pressure issue. tissue plasminogen activator was ran through the purge. the pump was succesfully weaned and the patient survived the explant. | [5 empty] | The investigation into high purge pressure has been completed. the pump was not returned for investigation. data logs showed a 15-minute gradual rise in purge pressure, with an active high purge pressure alarm for about five hours. the trend gradually returned to normal after eight hours. as per the clinician, tissue plasminogen activator was administered to resolve the high purge pressure. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and provided clinical details. b.7 information was inadvertently omitted from the initial manufacturer device report 1220648-2024-22912 and was added. g.1 revised reporting contact email since initial manufacturer device report 1220648-2024-22912 was submitted. | The impella device was not received from the customer and therefore. an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
14813 | 1220648-2024-22911 | 20663035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663035&pc=OZD | 06/24/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 06/24/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022131160 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. it was noted that impella 2.5 pump was unable to pass over 0.018" impella guidewire into descending aorta due to left iliac covered stent protruding into vessel. as a result, pump was explanted from right iliac artery. new impella 2.5 pump was opened, prepped and inserted via left femoral artery (lfa) without complications. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14814 | 1220648-2024-22916 | 20663056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663056&pc=OZD | 07/25/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 07/25/2022 | 11/12/2024 |  |  | 07/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 1488361 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. while on support, first puncture left fem failed due to missing puncture proglide was noted. bleeding to left fem was covered with stent. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14815 | 1220648-2024-22915 | 20663082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663082&pc=OZD | 06/28/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 06/28/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022152517 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the preclose procedure was completed; however, a small ooze was noted. a mynx closure device was deployed followed by manual compression and satisfactory hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14816 | 1220648-2024-22913 | 20663141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663141&pc=OZD | 06/27/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 06/27/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022124224 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was noted the impella pump could not be placed due to aortic valve stenosis. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14817 | 1220648-2024-22921 | 20663173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663173&pc=OZD | 08/15/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 08/15/2022 | 11/12/2024 |  |  | 11/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222822 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Other | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. while on support, it was noted that patient had 2- 8 second runs of ventricular tachycardia (vt). vt was resolved while mean arterial pressure (map) was maintained with impella. impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14818 | 1220648-2024-22918 | 20663221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663221&pc=OZD | 03/02/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 03/02/2022 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1591091 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomed reported for an automated impella controller (aic) that a battery failure was observed. a loaner console sent to initiate the return of aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery failure issue has been completed. the product was returned, and the issue was confirmed. data analysis: logs are consistent with complaint intake. subsequent boots following the complaint date trigger these alarms. log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 4 voltage of battery 1 had been declining for an extended period of time. device analysis: console arrived in danvers. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. per the results of the clinical review and investigation, the root cause was determined to be due to a defective battery. this report is being filed as part of a retrospective review of historical records.
14819 | 1220648-2024-22922 | 20663222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663222&pc=OZD | 08/28/2022 | 05/28/2025 |  | 11/12/2024 |  |  | 08/28/2022 | 11/12/2024 |  |  | 11/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222840 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, it was noted that patient's urine was blood tinged indicating hemolysis. activated clotted time (act) goals were discussed. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14820 | 1220648-2024-22917 | 20663227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663227&pc=OZD | 08/02/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 08/02/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022148665 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as pre-op placement was implanted with an impella 2.5 device for mechanical circulatory support. while on support, oozing and bleeding at groin site was noted. surgeon placed sutures at the groin site to stop the bleeding and blood products were required. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14821 | 1220648-2024-22920 | 20663236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663236&pc=OZD | 08/02/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 08/02/2022 | 11/12/2024 |  |  | 11/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023202889 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. impella was unable to be delivered due to patient's anatomy. impella placement procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14822 | 1220648-2024-22923 | 20663273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663273&pc=OZD | 08/31/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 08/31/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022114930 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was noted that the pump was inserted but could not pass through disease at aorta common femoral artery (cfa) bifurcation. procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14823 | 1220648-2024-22924 | 20663460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663460&pc=OZD | 01/11/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 01/11/2022 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022151098 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. during prep for the procedure the aic did not recognize the purge cassette. no patient involvement was noted. | [5 empty] | The investigation for the reported aic - purge disc/flag issues issue has been completed. the product was returned, and the issue was confirmed. the root cause of the aic inability to detect the purge pressure disc was due to a broken purge flag. this report is being filed as part of a retrospective review of historical records.
14824 | 1220648-2024-22925 | 20663468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663468&pc=OZD | 04/05/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/05/2021 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1440515 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was being placed on an impella 2.5 for hemodynamic support. during implantation attempt, the impella cp was opened and prepped. 14f x 23cm was initially attempted to be inserted but could not be advanced. the device was inserted under fluoroscopy but was unable to pass the motor housing into the iliac stent. the cp was removed and attempted plain old balloon angioplasty of the artery. there were several attempts without success to pass the cp into the artery. the surgeon then decided to insert an impella 2.5, however, could not pass the impella past the iliac stent. it was decided to abort the case and continue with medical management. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14825 | 1220648-2024-22928 | 20663535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663535&pc=OZD | 10/17/2024 | 12/13/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539127 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 being placed to support a surgical patient who had undergone coronary artery bypass grafting. it was reported that the case was aborted due to being unable to get atrioventricular opened to cross with catheter and 0.035 wire. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the av was unable to be opened. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
14826 | 1220648-2024-22932 | 20663620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663620&pc=OZD | 10/16/2024 | 12/20/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495719 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, there were frequent position in aorta alarms, but no position or suction alarms. the alarm was occurring and disappearing spontaneously repeatedly. since there were no hemodynamic problems, the catheter was weaned and removed. alarm did not occur at the time of removal. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation placement signal issue has been completed. there were no damages found and the pump. due to lack of clinical details and no data logs returned, the root cause of the optical signal issue cannot be determined the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22932.
14827 | 1220648-2024-22938 | 20663651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663651&pc=OZD | 04/09/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/09/2021 | 11/12/2024 |  |  | 04/30/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1431485 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that the 2.5 was weaned and removed. however, there was pulsatile bleeding noted. to resolve the issue, manual pressure was initiated and contralateral balloon tamponade performed and 20 mg protamine resulted in hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14828 | 1220648-2024-22931 | 20663654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663654&pc=PYX | 10/17/2024 | 12/26/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025495204 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella rp pump. it was stated that following insertion, the pump could not advance deep enough in the pulmonary artery. the catheter dislocated and the decision was made to abort the implant. there was no patient harm and an alternate option was utilized for support of the right side of the heart. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since it was not possible to get deep enough in the pulmonary artery. the following have been reported incorrectly or missing on manufacturer device report 1220648-2024-22931. g1 reporting contact email revised in accordance with updated procedures. g4 pma/510(k)number. h6 medical device problem code 2682 missing.
14829 | 1220648-2024-22951 | 20663743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663743&pc=OZD | 07/07/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 07/07/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060412 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, stenosis was observed in the left femoral artery when the impella sheath was removed. one balloon inflation was done via right femoral artery without incident. no sign bleeding or hematoma noted. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14830 | 1220648-2024-22952 | 20663840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663840&pc=OZD | 07/09/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 07/09/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022072412 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. during the implantation, multiple attempts were made to cross retrograde chronic total occlusion (cto) of circumflex artery with no success. moved to attempt cto of distal right coronary both retrograde and antegrade without success. the decision was made to abort the case due to fluoro time. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14831 | 1220648-2024-22926 | 20663903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663903&pc=OZD | 10/17/2024 | 12/24/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532235 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient supported on impella cp that was admitted in acute myocardial infarction and cardiogenic shock. the pump was placed and the team was giving the patient cardiopulmonary resuscitation chest compressions and the patient began to bleed. the team treated the bleeding with transfusions. patient remains on support therefore explant date and patient outcome are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the vascular damage has been completed. pump was not returned for investigation. data log review was not required for this case. as per the given clinical, chest compressions caused a sternum fracture and significant bleeding in the chest cavity, resulting in a right hemothorax. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between chest cavity bleeding and impella was unknown. corrections to the initial mfr report: g1 mdr reporting contact email.
14832 | 1220648-2024-22946 | 20663939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663939&pc=OZD | 10/17/2024 | 12/17/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562817 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a cp to support a patient admitted in a non-st elevated myocardial infarction. the pump was placed, and the percutaneous coronary intervention was accomplished, but the patient became agitated and sat upright on the table and caused hematoma. the bleed was treated by perclose that failed, then the staff tried contralateral ballooning. the patient was unable to recover and expired with the vascular damage. the angiogram had shown a tear in the left femoral artery. | [5 empty] | The investigation for the vascular damage has been completed. no product was returned and logs are not applicable. there was severe bleeding after the patient sat up during the procedure. physician noted that the patient should have been better sedated. the root cause of the vascular damage was most likely use-related. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including perforation) when, manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
14833 | 1220648-2024-22947 | 20663955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663955&pc=OZD | 06/24/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/24/2021 | 11/12/2024 |  |  | 03/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | GM |  | 2022068840 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, placement failed due to bleeding after the proglide was inserted. second approach with good placement and good hemodynamic support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14834 | 1220648-2024-22929 | 20663956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663956&pc=OZD | 10/17/2024 | 12/11/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553831 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the cp dove deep into the left ventricle (lv) and the inlet curled towards the mitral valve. the doctor pulled the cp out of the lv, and a kink was noted at the cannula right above the aortic valve. the doctor attempted to reposition to resolve the kink without success. the pump was explanted and replaced. the patient remained stable and there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. pump was not returned for investigation. data log review was not required for this case run. as per the clinical, pump was too deep into ventricle during advancement of sheath. pump was explanted. the cause of the positioning issue was most likely use related, based on given clinical details.
14835 | 1220648-2024-22933 | 20663983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663983&pc=OZD | 10/18/2024 | 12/27/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  | 05/31/2025 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351413 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the patient developed limb ischemia and hemolysis. fem-fem-bypass was performed. the impella leg was poorly supplied with blood. doppler hardly possible. due to hemolysis, there was a reduction of anticoagulation. the patient expired on support. | [5 empty] | The investigation of limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-22933 g1 reporting contact email revised on manufacturer device report 1220648-2024-22933 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
14836 | 1220648-2024-22941 | 20663990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20663990&pc=OZD | 10/18/2024 | 12/13/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545604 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp pump to support a patient with known hodgkin lymphoma, congestive heart failure and aortic stenosis. the team noted an access site bleed. the bleed was treated by the redressing of the site and re-sutures. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, . impella redressed and re-sutured to achieve hemostasis as per the clinical description provided. the following have been reported incorrectly on manufacturer device report 1220648-2024-22941. e4 should have been left blank. g1 reporting contact email.
14837 | 1220648-2024-22936 | 20664030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664030&pc=OZD | 09/01/2024 | 12/16/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [9 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded a publication from nyu columbia entitled clinical outcomes among cardiogenic shock patients supported with high-capacity impella axial flow pumps: a report from the cardiogenic shock working group in which cardiac patients were analyzed - 6,205 patients - 754 did have the impella 5.0/5.5 support. adverse events noted in the multicenter publication are limb ischemia, bleeding, hemolysis, and strokes. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. literature citation: fried, j., farr, m., kanwar, m., uriel, n., hernandez-montfort, j., blumer, v., li, s., sinha, s., garan, a., li, b., hall, s., hickey, g., mahr, c., nathan, s., schwartzman, a., kim, j., ton, v.-k., vishnevsky, o., vorovich, e., ¿ kapur, n. (2024). clinical outcomes among cardiogenic shock patients supported with high-capacity impella axial flow pumps: a report from the cardiogenic shock working group. the journal of heart and lung transplantation, 43(9), 1478¿1488. https://doi.org/10.1016/j.healun.2024.05.015 | The investigation for the reported access site bleed, hemolysis, limb ischemia, stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, hemolysis, limb ischemia, stroke could not be determined.
14838 | 1220648-2024-22942 | 20664043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664043&pc=OZD | 04/16/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/16/2021 | 11/12/2024 |  |  | 12/22/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054187 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support. the complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was reported that a pump had a start issue. the pump was restarted successfully. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14839 | 1220648-2024-22939 | 20664064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664064&pc=OZD | 10/18/2024 | 12/17/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support when they had hemolysis during support. the staff repositioned under echocardiogram and noticed a kink in the cannula. the pump was repositioned in the appropriate place, but the cannula would not come unkinked therefore the staff replaced the pump. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into product damage (cannula kink) and hemolysis has been completed. data logs confirmed the failure mode and clinical narrative. logs showed pump started and ran without issue for 6 hours and 15 minutes then suddenly flow dropped to 0.5 l/min that triggered auto suction response & suction alarms. this event remained to the rest of the case. logs also showed pump was in aortic area corresponded with the time hemolysis was reported. product was not returned for review. the root cause of low flow was most likely cannula kink since a kink was observed under echo during repositioning. the root cause of hemolysis was most likely due to pump was not in the proper position. the following have been reported incorrectly or missing from manufacturer device report.
14840 | 1220648-2024-22949 | 20664317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664317&pc=OZD | 09/01/2023 | 12/17/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded a publication from memorial hermann entitled prehabilitation maximizing functional mobility in patients with cardiogenic shock supported on axillary impella in which 5.0/5.5 patients had pt and noted the following adverse events/outcome information: the primary outcome of this study was in-hospital and 6-month mortality. the secondary outcomes were ambulation-related complications, including bleeding, hemolysis, stroke, pump malfunction, ventricular arrhythmia, and limb ischemia. all patients survived to hospital discharge and remained alive at 6 months. complications related to impella placement included five (13.5%) bleeding events, one (2.7%) case of hemolysis, one (2.7%) stroke, and two (5.4%) pump malfunctions. no ventricular arrhythmias or limb ischemia were reported. this report is to document the impella 5.5. | [5 empty] | D.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ literature citation: bhardwaj, a., salas de armas, i. a., bergeron, a., sauer, r. m., gilley, c., reeves, k., patarroyo-aponte, m., akay, m. h., patel, m., kumar, s., patel, j., marcano, j., nathan, s., gregoric, i. d., & kar, b. (2024). prehabilitation maximizing functional mobility in patients with cardiogenic shock supported on axillary impella. asaio journal, 70(8), 661¿666. https://doi.org/10.1097/mat.0000000000002170. | The investigation for the access site bleed, hemolysis, stroke, and pump stop has been completed. the pump was not returned for evaluation. clinical details were not sufficient to determine a root cause. therefore, the root cause of the access site bleed, hemolysis, stroke, and pump stop could not be determined. revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-22949 was submitted.
14841 | 1220648-2024-22927 | 20664357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664357&pc=OZD | 04/05/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/05/2021 | 11/12/2024 |  |  | 09/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021010266 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. the complainant reported that the patient was scheduled for a high-risk percutaneous coronary intervention by the cardiac physician. the pump was inserted and started. shortly after starting the pump, the pump stopped working. after running through the restart steps, the pump was removed and a new 2.5 was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14842 | 1220648-2024-22935 | 20664362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664362&pc=OZD | 10/18/2024 | 12/17/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534622 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was not flowing despite of fluids and repositioning. the cp was exchanged for a new cp. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. data logs were reviewed and showed low flow occurred when pump started and remained for 4 minutes, then flow was in normal range for about 19 minutes and dropped again and remained to the rest of the case. as the patient had mvd and poor lv function, the root cause of the low flow was most likely a patient condition related. corrections to the initial mfr report: g1 mdr reporting contact email.
14843 | 1220648-2024-22937 | 20664527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664527&pc=OZD | 10/18/2024 | 10/09/2025 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539050 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the aorta and left ventricle signals gradually decreased followed by an impella position unknown notice. pump placement was checked and adjusted without resolution of waveforms. motor current remained stable and unchanged, but flow has trended upward. there was no reported patient harm. the patient remains on support therefore the explant date is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issues has been completed. data logs were reviewed and placement signal shift or slight declining trend observed with impella position unknown alarms. slowly with ps declining trend, durability failure trend was found with placement signal drifting down until -70 mmhg and maxing out losing the placement monitoring. psnr alarms were recorded during this failure. no impact to optical 5s parameters were found during this time. psnr resolved with ps returning. the cause of the optical placement signal issue was most likely sensor silicone wear within design life. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. d6b updated. d10 concomitant medical products and therapy dates updated.
14844 | 1220648-2024-22934 | 20664653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664653&pc=OZD | 04/06/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/06/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022021728 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 for device for mechanical circulatory support. while on support there was a positioning issue. the pump position was showing unknown alarms. the field staff went over numbers to advise as the impella appeared to be in the incorrect position. no further information was provided and the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14845 | 1220648-2024-22954 | 20664759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664759&pc=OZD | 07/13/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 07/13/2021 | 11/12/2024 |  |  | 03/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022062116 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient was intubated and use of a wendl tube causes bleeding in nasopharynx. the impella migrated into the ventricle with inflow near to mitra valve annulus. consequences of possible reposition discussed with icu and cardiologist. because of hemolysis, bleeding in nasopharynx, the decision was made to wean impella. impella successfully weaned. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14846 | 1220648-2024-22950 | 20664796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664796&pc=OZD | 10/18/2024 | 12/16/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559865 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed without issues and started on auto. the peel away sheath was removed, and hematoma was noted. pressure held; forward tension sutures placed, and hematoma remained. pressure continued and figure 8 suture placed, and bleeding subsided. hematoma remained and continued. another dose of 10mg protamine given and hematoma mashed out and bleeding stopped. femstop placed prophylactically on pressure of 60. groin soft and not bleeding before leaving cath lab. additionally the patient had episode of sustained ventricular tachycardia requiring cardioversion. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22950.
14847 | 1220648-2024-22945 | 20664817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664817&pc=OZD | 06/21/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/21/2021 | 11/12/2024 |  |  | 07/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1432562 |  | 005048-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, after insertion, there was significant bleeding at the insertion site, so the patient was urgently weaned off. weaning was successful, the device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14848 | 1220648-2024-22943 | 20664874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664874&pc=OZD | 05/06/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/06/2021 | 11/12/2024 |  |  | 12/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022059027 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. during implantation there were delivery issues due to the patient's anatomy and inability to pass through the vasculature. there was a high degree of calcification, stenosis and kinking on both sides. it was reported a third proglide suture was needed to control bleeding from the access site. finally there was an uneventful implant, usage and explant of the impella 2.5. 2 stents were able to implanted with good results and stable hemodynamics. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14849 | 1220648-2024-22955 | 20664974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20664974&pc=OZD | 11/30/2022 | 03/13/2025 |  | 11/12/2024 |  |  | 11/30/2022 | 11/12/2024 |  |  | 01/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022049149 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella 2.5 device for mechanical circulatory support. during placement of the initial sheath, a perclose suture broke, and the decision was made to discontinue placement of the thirteen french sheath due to impella access issues. the procedure was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14850 | 1220648-2024-22944 | 20665132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665132&pc=OZD | 05/13/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/13/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022060418 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. it was reported that during implantation there were delivery issues. the procedure was aborted as the device was unable to pass through the patient¿s vasculature. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14851 | 1220648-2024-22940 | 20665314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665314&pc=PYX | 10/18/2024 | 12/18/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025536968 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, the rp was pulled back into the right ventricle (rv). attempts were made to readvance the pump, however the rp prolapsed deep into the rv and appeared to be kinked at the proximal cannula end near the inlet. the patient was taken to the cath lab to exchange the pump. the physician used tension release and the pump jumped forward to readvance itself across valve into main pulmonary artery. physicians confirmed good position and with improved flows and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into product damage (cannula kink) has been completed. the impella device was not returned for evaluation. product damage (cannula kink) changed to positioning issue as the issue occurred during repositioning attempts after the impella was pulled back into right ventricle (rv) in the unit when patient sat up in bed. the root cause of the positioning issue most likely was use related due to improper technique. the following have been reported incorrectly on manufacturer device report.
14852 | 1220648-2024-22953 | 20665365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665365&pc=OZD | 07/12/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 07/12/2021 | 11/12/2024 |  |  | 04/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022073565 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. the impella was implanted without any issues. upon successful explant of the impella, there was proglide failure, and the bleeding cannot stopped under manual compression. the decision was made to transfer the patient to or for vessel closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14853 | 1220648-2024-22930 | 20665368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665368&pc=OZD | 09/01/2024 | 12/16/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [9 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded a publication from nyu columbia entitled clinical outcomes among cardiogenic shock patients supported with high-capacity impella axial flow pumps: a report from the cardiogenic shock working group, in which cardiac patients were analyzed - 6,205 patients - 754 did have the impella 5.0/5.5 support. adverse events were noted in the multicenter publication. they were limb ischemia, bleeding, hemolysis, and strokes. | [5 empty] | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ literature citation: fried, j., farr, m., kanwar, m., uriel, n., hernandez-montfort, j., blumer, v., li, s., sinha, s. s., garan, a. r., li, b., hall, s., hickey, g. w., mahr, c., nathan, s., schwartzman, a., kim, j., ton, v.-k., vishnevsky, o. a., vorovich, e., ¿ kapur, n. k. (2024). clinical outcomes among cardiogenic shock patients supported with high-capacity impella axial flow pumps: a report from the cardiogenic shock working group. the journal of heart and lung transplantation, 43(9), 1478¿1488. https://doi.org/10.1016/j.healun.2024.05.015 | The investigation into access site bleeding, hemolysis, limb ischemia, and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22930.
14854 | 1220648-2024-22948 | 20665412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665412&pc=OZD | 06/30/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/30/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080538 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Heart Problem | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. during impella placement, the patient suffered a third episode of heart block. it was stated that to manage this complication, a temporary pacer was implanted for additional support. there was no report of patient harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14855 | 1220648-2024-22956 | 20665517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665517&pc=OZD | 10/18/2024 | 12/17/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541885 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that in an attempt to deliver the impella cp, non-flow limiting dissection of left subclavian occurred. motor would not cross from subclavian into aorta despite multiple attempts. sheath position was changed to attempt different angles of entry into aorta. at one point, the sheath advanced into aortic arch beyond subclavian stenosis but kink in sheath prohibited pump delivery. dissection of subclavian was noted. perforation requiring covered stent and/or balloon tamponade. there was persistent bleeding at the insertion site. covered stent was used to achieve hemostasis. decision was made to allow retrograde dissection flap to heal without intervention since natural direction of flow would keep flap against vessel wall. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) :¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the delivery issues and vascular damage has been completed. the impella was not returned for evaluation. clinical details were sufficient to determine the root cause of the delivery issues. the cause of the delivery issues was the use of an introducer not intended for axillary insertion with impella cp. the cause of the vascular damage was the use of an introducer not intended for axillary insertion with impella cp. clinical details state that the introducer was kinked during insertion. the cause of the introducer damage was the incorrect use of the 14fr x 25 cm introducer in an axillary insertion. corrections to the initial mfr report: g1 mdr reporting contact email.
14856 | 1220648-2024-22957 | 20665795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665795&pc=OZD | 10/18/2024 | 12/16/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support of 4 days for a patient admitted in with cardiomyopathy. the patient had persistent blood loss/ooze during the support and was treated by four units of red blood cells. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22957.
14857 | 1220648-2024-22964 | 20665802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20665802&pc=OZD | 04/16/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/16/2021 | 11/12/2024 |  |  | 07/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1452829 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient scheduled to undergo aortic valvuloplasty was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that after a successful insertion of the impella 2.5, there was a pump stop with a high motor current alarm which resolved after restarting and repositioning the device. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14858 | 1220648-2024-22965 | 20666091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20666091&pc=OZD | 04/19/2021 | 03/13/2025 |  | 11/12/2024 |  |  | 04/19/2021 | 11/12/2024 |  |  | 01/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054693 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced hemolysis. the pump was repositioning and the p level was lowered to p4. as a result of the ongoing hemolysis the patient was weaned and the impella 2.5 was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14859 | 1220648-2024-22961 | 20666307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20666307&pc=OZD | 04/16/2021 | 03/13/2025 |  | 11/12/2024 |  |  | 04/16/2021 | 11/12/2024 |  |  | 09/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021007211 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced a retroperitoneal bleed. it was reported that systemic heparin was withheld due to the bleed until the retroperitoneal bleed was stable. it was also reported that on (b)(6) 2021 the clinical staff was unable to locate doppler pulses on the patient's right foot, and as a result of the limb ischemia the impella 2.5 was weaned and explanted the following morning. the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14860 | 1220648-2024-22959 | 20666311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20666311&pc=OZD | 10/18/2024 | 11/12/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764189 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) showed a battery failure red alarm. the aic was replaced. no patient harm has been reported. | [5 empty] | The investigation has been completed. after reviewing the logs, the reported battery failure was confirmed. there were numerous instances of battery failure alarm (transport connection blown/missing) observed from the reported occurrence date. the console was returned and tested. the battery failure was able to be reproduced. results of running the batttest utility on telnet revealed that cell 3 in battery 2 had a voltage that was significantly less than the rest of the cells in the battery. the root cause of this issue was an internal cell imbalance in battery 2.